## Sheet1
This data recording was made with 5 temperatures sensors and a logging interval of 1 minutes over the course of 10 days.
Data recording was checked against a calibrated vaccine thermometer and showed a match to 1/2 degree ˚C (display resolution of the calibrated thermometer).
Recording began after a period of uncontrolled full-power cooling to show the temperatuere climb to the set point.					Cooler Bottom Shelf Temp. Over Time / Time Elapsed (min) / Temperature (˚C)
Cooler was then left open to warm to 9˚C (tallest spike) then cooler loaded with water bottles, thermostat was activated to see the cooling response.
Cooler was opened and closed several times at random (see slight temp. spikes).
Temperature was seen to spike at end of the recording when the ice was gone.

Year	Month	Day	Hour	Minute	Second	Ambient	Ice Chamber	Pipe	Bottom Shelf	Top Shelf
19	2	26	20	25	8	19.75	-38.06	-1.19	1.31	3.19
19	2	26	20	26	8	19.75	-37.94	-1.56	1.44	3.31
19	2	26	20	27	8	19.75	-38	-1.31	1.5	3.38
19	2	26	20	28	9	19.75	-37.94	-1.63	1.56	3.44
19	2	26	20	29	9	19.75	-37.94	-1.31	1.63	3.5
19	2	26	20	30	9	19.75	-37.94	-1.56	1.69	3.56
19	2	26	20	31	9	19.81	-37.88	-1.88	1.69	3.63
19	2	26	20	32	10	19.81	-37.94	-1.5	1.81	3.63
19	2	26	20	33	10	19.81	-37.94	-1.69	1.81	3.69
19	2	26	20	34	10	19.81	-37.94	-1.37	1.88	3.69
19	2	26	20	35	10	19.87	-37.88	-1.56	1.94	3.75
19	2	26	20	36	11	19.87	-37.81	-1.81	1.94	3.81
19	2	26	20	37	11	19.87	-37.81	-1.37	2	3.81
19	2	26	20	38	11	19.87	-37.81	-1.63	2.06	3.88
19	2	26	20	39	11	19.87	-37.88	-1.25	2.06	3.88
19	2	26	20	40	12	19.87	-37.81	-1.44	2.13	3.94
19	2	26	20	41	12	19.87	-37.81	-1.75	2.13	3.94
19	2	26	20	42	12	19.87	-37.88	-1.25	2.13	3.94
19	2	26	20	43	12	19.81	-37.81	-1.44	2.19	4
19	2	26	20	44	13	19.81	-37.88	-1.06	2.25	4
19	2	26	20	45	13	19.81	-37.81	-1.25	2.25	4.06
19	2	26	20	46	13	19.81	-37.81	-1.69	2.25	4.06
19	2	26	20	47	13	19.75	-37.81	-1.12	2.31	4.06
19	2	26	20	48	14	19.81	-37.69	-1.31	2.31	4.13
19	2	26	20	49	14	19.81	-37.81	-0.94	2.38	4.13
19	2	26	20	50	14	19.75	-37.75	-1.12	2.38	4.13
19	2	26	20	51	14	19.81	-37.75	-1.31	2.38	4.19
19	2	26	20	52	15	19.81	-37.75	-0.94	2.38	4.19
19	2	26	20	53	15	19.81	-37.69	-1.19	2.44	4.19
19	2	26	20	54	15	19.87	-37.81	-0.81	2.44	4.25
19	2	26	20	55	15	19.87	-37.75	-1	2.5	4.25
19	2	26	20	56	16	19.87	-37.81	-0.63	2.5	4.25
19	2	26	20	57	16	19.94	-37.81	-0.37	2.56	4.31
19	2	26	20	58	16	19.94	-37.75	-0.69	2.56	4.38
19	2	26	20	59	16	20	-37.81	-0.31	2.63	4.38
19	2	26	21	0	17	20	-37.75	-0.63	2.63	4.44
19	2	26	21	1	17	20	-37.81	-0.25	2.69	4.44
19	2	26	21	2	17	20	-37.81	-0.5	2.69	4.44
19	2	26	21	3	17	20.06	-37.75	-0.81	2.69	4.5
19	2	26	21	4	18	20	-37.75	-0.37	2.75	4.5
19	2	26	21	5	18	20	-37.69	-0.63	2.75	4.5
19	2	26	21	6	18	20	-37.75	-0.25	2.75	4.5
19	2	26	21	7	18	20	-37.69	-0.44	2.75	4.56
19	2	26	21	8	19	19.94	-37.69	-0.81	2.81	4.56
19	2	26	21	9	19	19.94	-37.75	-0.31	2.81	4.56
19	2	26	21	10	19	19.87	-37.63	-0.56	2.88	4.63
19	2	26	21	11	19	19.87	-37.69	-0.19	2.88	4.63
19	2	26	21	12	20	19.81	-37.69	-0.37	2.88	4.63
19	2	26	21	13	20	19.81	-37.63	-0.81	2.88	4.63
19	2	26	21	14	20	19.81	-37.69	-0.25	2.94	4.69
19	2	26	21	15	20	19.81	-37.63	-0.5	2.94	4.69
19	2	26	21	16	21	19.81	-37.69	-0.12	2.94	4.69
19	2	26	21	17	21	19.81	-37.63	-0.31	2.94	4.69
19	2	26	21	18	21	19.81	-37.63	-0.69	3	4.75
19	2	26	21	19	21	19.81	-37.63	-0.19	3	4.75
19	2	26	21	20	22	19.81	-37.56	-0.44	3	4.75
19	2	26	21	21	22	19.87	-37.63	0	3	4.75
19	2	26	21	22	22	19.87	-37.56	-0.25	3.06	4.81
19	2	26	21	23	22	19.87	-37.63	0.06	3.06	4.81
19	2	26	21	24	23	19.87	-37.63	-0.5	3.06	4.75
19	2	26	21	25	23	19.94	-37.56	-0.81	3.06	4.75
19	2	26	21	26	23	19.94	-37.63	-0.25	3.06	4.75
19	2	26	21	27	23	19.94	-37.56	-0.5	3.06	4.75
19	2	26	21	28	24	19.87	-37.63	-0.06	3.06	4.75
19	2	26	21	29	24	19.87	-37.63	-0.25	3.06	4.75
19	2	26	21	30	24	19.87	-37.56	-0.56	3.13	4.81
19	2	26	21	31	24	19.81	-37.63	-0.06	3.13	4.81
19	2	26	21	32	25	19.81	-37.56	-0.31	3.13	4.81
19	2	26	21	33	25	19.81	-37.63	0.06	3.13	4.81
19	2	26	21	34	25	19.75	-37.56	-0.12	3.13	4.81
19	2	26	21	35	25	19.75	-37.5	-0.5	3.19	4.88
19	2	26	21	36	26	19.75	-37.56	0	3.19	4.88
19	2	26	21	37	26	19.75	-37.5	-0.25	3.19	4.88
19	2	26	21	38	26	19.69	-37.5	0.12	3.19	4.88
19	2	26	21	39	26	19.75	-37.5	-0.06	3.19	4.88
19	2	26	21	40	27	19.69	-37.5	-0.5	3.19	4.94
19	2	26	21	41	27	19.75	-37.5	0.06	3.25	4.94
19	2	26	21	42	27	19.75	-37.5	-0.19	3.25	4.94
19	2	26	21	43	27	19.75	-37.56	0.19	3.25	4.94
19	2	26	21	44	28	19.75	-37.56	0	3.25	4.94
19	2	26	21	45	28	19.75	-37.56	-0.44	3.25	5
19	2	26	21	46	28	19.75	-37.56	0.12	3.25	5
19	2	26	21	47	28	19.81	-37.5	-0.19	3.31	5
19	2	26	21	48	29	19.81	-37.56	0.25	3.31	5
19	2	26	21	49	29	19.81	-37.56	0	3.31	5
19	2	26	21	50	29	19.81	-37.56	0.31	3.31	5.06
19	2	26	21	51	29	19.87	-37.56	-0.25	3.31	5
19	2	26	21	52	30	19.87	-37.56	-0.5	3.31	5
19	2	26	21	53	30	19.87	-37.56	0	3.31	5
19	2	26	21	54	30	19.94	-37.56	-0.25	3.31	5
19	2	26	21	55	30	19.94	-37.56	0.12	3.31	5
19	2	26	21	56	31	20	-37.56	-0.06	3.31	5
19	2	26	21	57	31	20	-37.5	-0.37	3.31	5
19	2	26	21	58	31	20	-37.56	0.12	3.38	5
19	2	26	21	59	31	19.94	-37.44	-0.12	3.38	5
19	2	26	22	0	32	19.94	-37.5	0.25	3.38	5.06
19	2	26	22	1	32	19.94	-37.5	0.06	3.38	5.06
19	2	26	22	2	32	19.87	-37.44	-0.31	3.38	5.06
19	2	26	22	3	32	19.87	-37.5	0.19	3.38	5.06
19	2	26	22	4	33	19.87	-37.44	-0.06	3.38	5.06
19	2	26	22	5	33	19.81	-37.44	0.31	3.44	5.06
19	2	26	22	6	33	19.81	-37.44	0.06	3.44	5.13
19	2	26	22	7	33	19.81	-37.44	-0.31	3.44	5.06
19	2	26	22	8	34	19.81	-37.44	0.25	3.44	5.13
19	2	26	22	9	34	19.81	-37.38	0	3.44	5.13
19	2	26	22	10	34	19.81	-37.44	0.37	3.44	5.13
19	2	26	22	11	34	19.81	-37.44	0.19	3.44	5.13
19	2	26	22	12	35	19.81	-37.44	-0.19	3.5	5.13
19	2	26	22	13	35	19.87	-37.38	0.37	3.5	5.13
19	2	26	22	14	35	19.87	-37.38	0.06	3.5	5.19
19	2	26	22	15	35	19.87	-37.44	0.5	3.5	5.19
19	2	26	22	16	36	19.87	-37.38	0.19	3.5	5.19
19	2	26	22	17	36	19.87	-37.44	0.5	3.5	5.19
19	2	26	22	18	36	19.87	-37.5	-0.06	3.5	5.19
19	2	26	22	19	36	19.87	-37.44	-0.37	3.5	5.19
19	2	26	22	20	37	19.87	-37.5	0.19	3.5	5.19
19	2	26	22	21	37	19.94	-37.44	-0.06	3.5	5.19
19	2	26	22	22	37	19.94	-37.44	0.37	3.5	5.19
19	2	26	22	23	37	19.94	-37.44	0.12	3.5	5.19
19	2	26	22	24	38	19.87	-37.38	-0.19	3.56	5.19
19	2	26	22	25	38	19.87	-37.38	0.31	3.5	5.19
19	2	26	22	26	38	19.87	-37.31	0.06	3.56	5.19
19	2	26	22	27	38	19.81	-37.38	0.5	3.56	5.19
19	2	26	22	28	39	19.81	-37.31	0.25	3.56	5.19
19	2	26	22	29	39	19.75	-37.31	-0.12	3.56	5.25
19	2	26	22	30	39	19.75	-37.31	0.37	3.56	5.25
19	2	26	22	31	39	19.69	-37.25	0.12	3.56	5.25
19	2	26	22	32	40	19.69	-37.31	0.5	3.63	5.25
19	2	26	22	33	40	19.69	-37.31	0.31	3.63	5.25
19	2	26	22	34	40	19.69	-37.31	-0.12	3.63	5.25
19	2	26	22	35	41	19.69	-37.31	0.44	3.63	5.31
19	2	26	22	36	41	19.69	-37.25	0.19	3.63	5.31
19	2	26	22	37	41	19.69	-37.31	0.63	3.63	5.31
19	2	26	22	38	41	19.69	-37.25	0.37	3.69	5.31
19	2	26	22	39	42	19.69	-37.31	-0.06	3.63	5.31
19	2	26	22	40	42	19.69	-37.31	0.5	3.69	5.31
19	2	26	22	41	42	19.69	-37.25	0.25	3.69	5.38
19	2	26	22	42	42	19.69	-37.31	0.63	3.69	5.38
19	2	26	22	43	43	19.69	-37.25	0.44	3.69	5.38
19	2	26	22	44	43	19.75	-37.25	0.75	3.69	5.38
19	2	26	22	45	43	19.75	-37.31	0.19	3.69	5.38
19	2	26	22	46	43	19.75	-37.25	-0.12	3.69	5.38
19	2	26	22	47	44	19.75	-37.31	0.37	3.69	5.38
19	2	26	22	48	44	19.75	-37.25	0.12	3.69	5.38
19	2	26	22	49	44	19.75	-37.31	0.56	3.69	5.38
19	2	26	22	50	44	19.75	-37.31	0.31	3.69	5.38
19	2	26	22	51	45	19.75	-37.25	0	3.69	5.38
19	2	26	22	52	45	19.81	-37.25	0.5	3.69	5.38
19	2	26	22	53	45	19.75	-37.19	0.25	3.75	5.38
19	2	26	22	54	45	19.81	-37.25	0.63	3.69	5.38
19	2	26	22	55	46	19.81	-37.19	0.44	3.69	5.38
19	2	26	22	56	46	19.81	-37.19	0.06	3.75	5.44
19	2	26	22	57	46	19.81	-37.19	0.56	3.75	5.44
19	2	26	22	58	46	19.87	-37.19	0.31	3.75	5.44
19	2	26	22	59	47	19.87	-37.19	0.75	3.75	5.44
19	2	26	23	0	47	19.87	-37.19	0.5	3.75	5.44
19	2	26	23	1	47	19.87	-37.13	0.06	3.75	5.44
19	2	26	23	2	47	19.87	-37.19	0.63	3.75	5.44
19	2	26	23	3	48	19.87	-37.13	0.37	3.75	5.5
19	2	26	23	4	48	19.94	-37.19	0.75	3.81	5.5
19	2	26	23	5	48	19.94	-37.19	0.56	3.81	5.5
19	2	26	23	6	48	19.94	-37.19	0.12	3.81	5.5
19	2	26	23	7	49	19.94	-37.19	0.69	3.81	5.5
19	2	26	23	8	49	19.94	-37.06	0.44	3.81	5.5
19	2	26	23	9	49	19.87	-37.13	0.81	3.81	5.5
19	2	26	23	10	49	19.87	-37.06	0.56	3.81	5.5
19	2	26	23	11	50	19.87	-37.13	0.88	3.88	5.5
19	2	26	23	12	50	19.81	-37.19	0.31	3.88	5.5
19	2	26	23	13	50	19.81	-37.13	0	3.81	5.5
19	2	26	23	14	50	19.75	-37.19	0.5	3.81	5.5
19	2	26	23	15	51	19.75	-37.13	0.25	3.81	5.5
19	2	26	23	16	51	19.69	-37.19	0.69	3.81	5.5
19	2	26	23	17	51	19.69	-37.13	0.5	3.81	5.5
19	2	26	23	18	51	19.69	-37.06	0.19	3.81	5.5
19	2	26	23	19	52	19.69	-37.13	0.69	3.81	5.5
19	2	26	23	20	52	19.69	-37.06	0.44	3.88	5.5
19	2	26	23	21	52	19.69	-37.13	0.81	3.88	5.5
19	2	26	23	22	52	19.69	-37.13	0.56	3.88	5.5
19	2	26	23	23	53	19.69	-37.06	0.19	3.88	5.56
19	2	26	23	24	53	19.69	-37.13	0.75	3.88	5.56
19	2	26	23	25	53	19.69	-37	0.44	3.88	5.56
19	2	26	23	26	53	19.75	-37.06	0.88	3.88	5.56
19	2	26	23	27	54	19.75	-37.06	0.63	3.88	5.56
19	2	26	23	28	54	19.75	-37.06	0.19	3.88	5.56
19	2	26	23	29	54	19.75	-37.06	0.75	3.94	5.56
19	2	26	23	30	54	19.75	-37	0.5	3.94	5.56
19	2	26	23	31	55	19.75	-37.06	0.94	3.94	5.63
19	2	26	23	32	55	19.75	-37.06	0.63	3.94	5.63
19	2	26	23	33	55	19.75	-37.06	0.19	3.94	5.63
19	2	26	23	34	55	19.75	-37.06	0.81	3.94	5.63
19	2	26	23	35	56	19.75	-37.06	0.44	3.94	5.63
19	2	26	23	36	56	19.75	-37.06	0.88	3.94	5.63
19	2	26	23	37	56	19.75	-37.06	0.69	3.94	5.63
19	2	26	23	38	56	19.75	-37.06	1	3.94	5.63
19	2	26	23	39	57	19.69	-37.13	0.44	3.94	5.63
19	2	26	23	40	57	19.69	-37.06	0.12	3.94	5.63
19	2	26	23	41	57	19.62	-37.13	0.63	3.94	5.63
19	2	26	23	42	57	19.62	-37.06	0.37	3.94	5.56
19	2	26	23	43	58	19.56	-37.13	0.81	3.94	5.63
19	2	26	23	44	58	19.56	-37.13	0.56	3.94	5.56
19	2	26	23	45	58	19.56	-37.06	0.25	3.94	5.63
19	2	26	23	46	58	19.56	-37.13	0.75	3.94	5.63
19	2	26	23	47	59	19.56	-37.06	0.5	4	5.63
19	2	26	23	48	59	19.56	-37.13	0.88	3.94	5.63
19	2	26	23	49	59	19.56	-37.06	0.69	3.94	5.63
19	2	26	23	50	59	19.56	-37.06	0.31	3.94	5.63
19	2	26	23	52	0	19.56	-37.06	0.81	3.94	5.63
19	2	26	23	53	0	19.56	-37	0.56	4	5.63
19	2	26	23	54	0	19.56	-37.06	0.94	4	5.63
19	2	26	23	55	0	19.56	-37	0.69	4	5.69
19	2	26	23	56	1	19.62	-37	0.25	4	5.69
19	2	26	23	57	1	19.62	-37.06	0.81	4	5.63
19	2	26	23	58	1	19.62	-37	0.56	4	5.69
19	2	26	23	59	1	19.62	-37.06	0.94	4	5.69
19	2	27	0	0	2	19.62	-37.06	0.69	4	5.69
19	2	27	0	1	2	19.69	-37.06	0.25	4	5.69
19	2	27	0	2	2	19.69	-37.06	0.88	4	5.69
19	2	27	0	3	2	19.69	-37	0.56	4.06	5.69
19	2	27	0	4	3	19.75	-37.06	1	4.06	5.69
19	2	27	0	5	3	19.75	-37	0.75	4.06	5.69
19	2	27	0	6	3	19.75	-37	1.06	4.06	5.69
19	2	27	0	7	3	19.81	-37	0.5	4.06	5.69
19	2	27	0	8	4	19.81	-37	0.19	4.06	5.69
19	2	27	0	9	4	19.81	-37	0.69	4	5.63
19	2	27	0	10	4	19.87	-37	0.44	4	5.63
19	2	27	0	11	4	19.87	-37	0.88	4	5.63
19	2	27	0	12	5	19.87	-37	0.63	4	5.63
19	2	27	0	13	5	19.87	-36.94	0.31	4	5.63
19	2	27	0	14	5	19.87	-37	0.81	4	5.63
19	2	27	0	15	5	19.94	-36.94	0.56	4	5.63
19	2	27	0	16	6	19.94	-37	0.94	4	5.63
19	2	27	0	17	6	19.94	-37	0.75	4	5.69
19	2	27	0	18	6	19.94	-36.94	0.37	4	5.69
19	2	27	0	19	6	19.94	-37	0.88	4	5.69
19	2	27	0	20	7	19.94	-36.94	0.63	4	5.69
19	2	27	0	21	7	19.94	-36.94	1	4	5.69
19	2	27	0	22	7	19.87	-36.88	0.75	4	5.69
19	2	27	0	23	7	19.87	-36.94	0.31	4.06	5.69
19	2	27	0	24	8	19.81	-36.94	0.88	4.06	5.69
19	2	27	0	25	8	19.81	-36.88	0.63	4.06	5.69
19	2	27	0	26	8	19.75	-36.94	1	4.06	5.69
19	2	27	0	27	8	19.75	-36.94	0.75	4.06	5.69
19	2	27	0	28	9	19.69	-36.94	0.31	4.06	5.75
19	2	27	0	29	9	19.69	-36.94	0.88	4.06	5.75
19	2	27	0	30	9	19.62	-36.88	0.63	4.06	5.75
19	2	27	0	31	9	19.62	-36.94	1	4.06	5.75
19	2	27	0	32	10	19.56	-36.94	0.75	4.06	5.75
19	2	27	0	33	10	19.56	-36.94	1.06	4.06	5.75
19	2	27	0	34	10	19.56	-37	0.5	4.06	5.75
19	2	27	0	35	10	19.56	-37	0.19	4.06	5.75
19	2	27	0	36	11	19.56	-37.06	0.69	4.06	5.69
19	2	27	0	37	11	19.56	-36.94	0.5	4.06	5.69
19	2	27	0	38	11	19.56	-36.94	0.88	4.06	5.69
19	2	27	0	39	11	19.56	-36.94	0.69	4.06	5.69
19	2	27	0	40	12	19.56	-36.88	0.31	4.06	5.69
19	2	27	0	41	12	19.56	-36.94	0.81	4.06	5.69
19	2	27	0	42	12	19.56	-36.88	0.56	4.06	5.69
19	2	27	0	43	12	19.56	-36.94	0.94	4.06	5.69
19	2	27	0	44	13	19.62	-36.88	0.75	4.06	5.69
19	2	27	0	45	13	19.62	-36.81	0.37	4.06	5.69
19	2	27	0	46	13	19.62	-36.88	0.88	4.06	5.69
19	2	27	0	47	13	19.62	-36.88	0.63	4.06	5.75
19	2	27	0	48	14	19.62	-36.88	1	4.06	5.75
19	2	27	0	49	14	19.62	-36.88	0.81	4.06	5.75
19	2	27	0	50	14	19.62	-36.88	0.31	4.06	5.75
19	2	27	0	51	14	19.69	-36.88	0.88	4.13	5.75
19	2	27	0	52	15	19.69	-36.88	0.63	4.06	5.75
19	2	27	0	53	15	19.69	-36.88	1.06	4.06	5.75
19	2	27	0	54	15	19.69	-36.88	0.81	4.06	5.75
19	2	27	0	55	15	19.69	-36.88	0.37	4.13	5.75
19	2	27	0	56	16	19.69	-36.88	0.94	4.06	5.75
19	2	27	0	57	16	19.75	-36.81	0.63	4.13	5.75
19	2	27	0	58	16	19.69	-36.88	1.06	4.13	5.75
19	2	27	0	59	16	19.75	-36.88	0.81	4.13	5.75
19	2	27	1	0	17	19.75	-36.88	1.12	4.13	5.75
19	2	27	1	1	17	19.75	-36.94	0.56	4.13	5.75
19	2	27	1	2	17	19.75	-36.94	0.25	4.13	5.75
19	2	27	1	3	17	19.75	-36.94	0.81	4.13	5.75
19	2	27	1	4	18	19.75	-36.88	0.56	4.13	5.75
19	2	27	1	5	18	19.75	-36.88	0.94	4.06	5.75
19	2	27	1	6	18	19.81	-36.88	0.75	4.06	5.69
19	2	27	1	7	18	19.81	-36.88	0.37	4.06	5.75
19	2	27	1	8	19	19.81	-36.88	0.88	4.06	5.69
19	2	27	1	9	19	19.87	-36.88	0.63	4.13	5.75
19	2	27	1	10	19	19.81	-36.88	1	4.13	5.75
19	2	27	1	11	19	19.87	-36.88	0.81	4.13	5.75
19	2	27	1	12	20	19.87	-36.81	0.37	4.13	5.75
19	2	27	1	13	20	19.81	-36.88	0.94	4.13	5.75
19	2	27	1	14	20	19.81	-36.81	0.69	4.13	5.75
19	2	27	1	15	20	19.81	-36.88	1.06	4.13	5.75
19	2	27	1	16	21	19.81	-36.88	0.81	4.13	5.75
19	2	27	1	17	21	19.75	-36.88	0.37	4.13	5.75
19	2	27	1	18	21	19.75	-36.88	0.94	4.13	5.75
19	2	27	1	19	21	19.69	-36.81	0.69	4.13	5.75
19	2	27	1	20	22	19.62	-36.88	1.12	4.13	5.75
19	2	27	1	21	22	19.62	-36.81	0.88	4.13	5.81
19	2	27	1	22	22	19.62	-36.88	0.44	4.13	5.81
19	2	27	1	23	22	19.62	-36.81	1	4.13	5.81
19	2	27	1	24	23	19.62	-36.75	0.69	4.19	5.81
19	2	27	1	25	23	19.62	-36.81	1.12	4.13	5.81
19	2	27	1	26	23	19.62	-36.81	0.88	4.13	5.81
19	2	27	1	27	23	19.62	-36.81	1.19	4.19	5.81
19	2	27	1	28	24	19.62	-36.88	0.63	4.19	5.81
19	2	27	1	29	24	19.62	-36.88	0.31	4.19	5.81
19	2	27	1	30	24	19.62	-36.94	0.81	4.19	5.75
19	2	27	1	31	25	19.62	-36.88	0.63	4.13	5.75
19	2	27	1	32	25	19.62	-36.88	1	4.13	5.75
19	2	27	1	33	25	19.62	-36.81	0.75	4.13	5.75
19	2	27	1	34	25	19.62	-36.81	0.44	4.13	5.75
19	2	27	1	35	26	19.62	-36.88	0.94	4.13	5.75
19	2	27	1	36	26	19.69	-36.75	0.75	4.13	5.75
19	2	27	1	37	26	19.69	-36.81	1.12	4.19	5.81
19	2	27	1	38	26	19.69	-36.75	0.88	4.13	5.81
19	2	27	1	39	27	19.69	-36.75	0.44	4.19	5.81
19	2	27	1	40	27	19.69	-36.81	1	4.13	5.81
19	2	27	1	41	27	19.75	-36.81	0.75	4.19	5.81
19	2	27	1	42	27	19.69	-36.81	1.12	4.19	5.81
19	2	27	1	43	28	19.75	-36.81	0.94	4.19	5.81
19	2	27	1	44	28	19.69	-36.81	0.44	4.19	5.81
19	2	27	1	45	28	19.69	-36.81	1.06	4.19	5.81
19	2	27	1	46	28	19.69	-36.75	0.75	4.19	5.81
19	2	27	1	47	29	19.62	-36.81	1.19	4.19	5.81
19	2	27	1	48	29	19.62	-36.81	0.94	4.19	5.81
19	2	27	1	49	29	19.56	-36.81	0.5	4.19	5.88
19	2	27	1	50	29	19.56	-36.81	1.06	4.19	5.88
19	2	27	1	51	30	19.56	-36.75	0.75	4.19	5.88
19	2	27	1	52	30	19.5	-36.81	1.19	4.19	5.88
19	2	27	1	53	30	19.5	-36.81	1	4.25	5.88
19	2	27	1	54	30	19.5	-36.81	1.25	4.25	5.88
19	2	27	1	55	31	19.5	-36.88	0.69	4.25	5.88
19	2	27	1	56	31	19.5	-36.81	0.44	4.19	5.81
19	2	27	1	57	31	19.5	-36.81	0.88	4.19	5.81
19	2	27	1	58	31	19.5	-36.75	0.69	4.19	5.81
19	2	27	1	59	32	19.5	-36.81	1.06	4.19	5.81
19	2	27	2	0	32	19.5	-36.75	0.88	4.19	5.81
19	2	27	2	1	32	19.5	-36.69	0.5	4.19	5.81
19	2	27	2	2	32	19.5	-36.75	1	4.19	5.81
19	2	27	2	3	33	19.56	-36.69	0.81	4.19	5.81
19	2	27	2	4	33	19.56	-36.75	1.19	4.19	5.81
19	2	27	2	5	33	19.56	-36.69	0.94	4.19	5.81
19	2	27	2	6	33	19.56	-36.69	0.56	4.19	5.88
19	2	27	2	7	34	19.56	-36.69	1.06	4.19	5.88
19	2	27	2	8	34	19.56	-36.63	0.81	4.19	5.88
19	2	27	2	9	34	19.56	-36.69	1.19	4.19	5.88
19	2	27	2	10	34	19.62	-36.69	1	4.19	5.88
19	2	27	2	11	35	19.62	-36.69	0.56	4.19	5.88
19	2	27	2	12	35	19.62	-36.75	1.12	4.19	5.88
19	2	27	2	13	35	19.62	-36.69	0.81	4.25	5.88
19	2	27	2	14	35	19.62	-36.69	1.25	4.25	5.88
19	2	27	2	15	36	19.69	-36.69	1	4.25	5.88
19	2	27	2	16	36	19.69	-36.75	0.56	4.25	5.88
19	2	27	2	17	36	19.75	-36.75	1.12	4.25	5.88
19	2	27	2	18	36	19.75	-36.75	0.81	4.25	5.94
19	2	27	2	19	37	19.75	-36.81	1.25	4.25	5.88
19	2	27	2	20	37	19.75	-36.75	1	4.25	5.94
19	2	27	2	21	37	19.75	-36.75	1.31	4.25	5.94
19	2	27	2	22	37	19.81	-36.81	0.69	4.25	5.88
19	2	27	2	23	38	19.75	-36.81	0.44	4.25	5.88
19	2	27	2	24	38	19.75	-36.81	0.94	4.25	5.88
19	2	27	2	25	38	19.75	-36.81	0.69	4.25	5.88
19	2	27	2	26	38	19.75	-36.81	1.12	4.25	5.88
19	2	27	2	27	39	19.69	-36.81	0.94	4.25	5.88
19	2	27	2	28	39	19.69	-36.75	0.56	4.25	5.88
19	2	27	2	29	39	19.62	-36.75	1.06	4.25	5.88
19	2	27	2	30	39	19.62	-36.75	0.81	4.25	5.88
19	2	27	2	31	40	19.56	-36.75	1.19	4.25	5.88
19	2	27	2	32	40	19.56	-36.75	1	4.25	5.88
19	2	27	2	33	40	19.56	-36.69	0.63	4.25	5.88
19	2	27	2	34	40	19.56	-36.75	1.12	4.25	5.88
19	2	27	2	35	41	19.56	-36.69	0.88	4.25	5.88
19	2	27	2	36	41	19.56	-36.75	1.25	4.25	5.88
19	2	27	2	37	41	19.56	-36.75	1	4.25	5.88
19	2	27	2	38	41	19.62	-36.69	0.56	4.25	5.94
19	2	27	2	39	42	19.62	-36.69	1.19	4.25	5.88
19	2	27	2	40	42	19.62	-36.63	0.88	4.25	5.94
19	2	27	2	41	42	19.62	-36.69	1.25	4.25	5.94
19	2	27	2	42	42	19.62	-36.63	1.06	4.25	5.94
19	2	27	2	43	43	19.62	-36.63	0.63	4.25	5.94
19	2	27	2	44	43	19.62	-36.69	1.19	4.31	5.94
19	2	27	2	45	43	19.69	-36.63	0.88	4.31	5.94
19	2	27	2	46	43	19.62	-36.69	1.31	4.31	5.94
19	2	27	2	47	44	19.69	-36.69	1.06	4.31	5.94
19	2	27	2	48	44	19.69	-36.69	1.37	4.31	5.94
19	2	27	2	49	44	19.69	-36.75	0.75	4.31	5.94
19	2	27	2	50	44	19.69	-36.75	0.5	4.31	5.94
19	2	27	2	51	45	19.69	-36.75	1	4.31	5.88
19	2	27	2	52	45	19.69	-36.69	0.75	4.31	5.88
19	2	27	2	53	45	19.69	-36.75	1.19	4.25	5.88
19	2	27	2	54	45	19.69	-36.75	0.94	4.25	5.88
19	2	27	2	55	46	19.69	-36.69	0.63	4.25	5.88
19	2	27	2	56	46	19.69	-36.75	1.12	4.31	5.88
19	2	27	2	57	46	19.69	-36.69	0.88	4.25	5.88
19	2	27	2	58	46	19.62	-36.75	1.25	4.31	5.88
19	2	27	2	59	47	19.62	-36.75	1.06	4.25	5.94
19	2	27	3	0	47	19.56	-36.69	0.69	4.31	5.94
19	2	27	3	1	47	19.56	-36.75	1.19	4.25	5.94
19	2	27	3	2	47	19.5	-36.69	0.88	4.31	5.94
19	2	27	3	3	48	19.5	-36.69	1.31	4.31	5.94
19	2	27	3	4	48	19.5	-36.63	1.06	4.31	5.94
19	2	27	3	5	48	19.5	-36.63	0.63	4.31	5.94
19	2	27	3	6	48	19.44	-36.69	1.19	4.31	5.94
19	2	27	3	7	49	19.5	-36.63	0.94	4.31	5.94
19	2	27	3	8	49	19.5	-36.69	1.31	4.31	5.94
19	2	27	3	9	49	19.5	-36.63	1.06	4.31	5.94
19	2	27	3	10	49	19.5	-36.63	0.69	4.31	6
19	2	27	3	11	50	19.5	-36.63	1.25	4.31	5.94
19	2	27	3	12	50	19.5	-36.63	0.94	4.31	6
19	2	27	3	13	50	19.5	-36.63	1.37	4.31	6
19	2	27	3	14	50	19.56	-36.63	1.12	4.31	6
19	2	27	3	15	51	19.56	-36.63	1.44	4.38	6
19	2	27	3	16	51	19.56	-36.69	0.88	4.31	6
19	2	27	3	17	51	19.62	-36.69	0.56	4.31	6
19	2	27	3	18	51	19.62	-36.69	1.06	4.31	5.94
19	2	27	3	19	52	19.62	-36.69	0.81	4.31	5.94
19	2	27	3	20	52	19.69	-36.69	1.19	4.31	5.94
19	2	27	3	21	52	19.69	-36.69	1	4.31	5.94
19	2	27	3	22	52	19.69	-36.63	0.69	4.31	5.94
19	2	27	3	23	53	19.62	-36.69	1.19	4.31	5.94
19	2	27	3	24	53	19.62	-36.63	0.94	4.31	5.94
19	2	27	3	25	53	19.62	-36.69	1.31	4.31	5.94
19	2	27	3	26	53	19.56	-36.69	1.12	4.31	5.94
19	2	27	3	27	54	19.56	-36.63	0.75	4.31	5.94
19	2	27	3	28	54	19.5	-36.69	1.25	4.31	5.94
19	2	27	3	29	54	19.5	-36.63	1	4.31	5.94
19	2	27	3	30	54	19.5	-36.69	1.37	4.31	5.94
19	2	27	3	31	55	19.5	-36.63	1.12	4.31	6
19	2	27	3	32	55	19.5	-36.63	0.69	4.31	6
19	2	27	3	33	55	19.5	-36.63	1.25	4.31	6
19	2	27	3	34	55	19.44	-36.56	1	4.31	6
19	2	27	3	35	56	19.5	-36.63	1.37	4.38	6
19	2	27	3	36	56	19.5	-36.56	1.12	4.38	6
19	2	27	3	37	56	19.5	-36.56	0.75	4.38	6
19	2	27	3	38	56	19.5	-36.56	1.31	4.38	6
19	2	27	3	39	57	19.5	-36.5	1	4.38	6
19	2	27	3	40	57	19.5	-36.56	1.37	4.38	6
19	2	27	3	41	57	19.5	-36.56	1.19	4.38	6
19	2	27	3	42	57	19.5	-36.63	1.5	4.38	6.06
19	2	27	3	43	58	19.56	-36.69	0.88	4.38	6
19	2	27	3	44	58	19.56	-36.63	0.56	4.38	6
19	2	27	3	45	58	19.56	-36.69	1.06	4.38	6
19	2	27	3	46	58	19.56	-36.63	0.81	4.38	6
19	2	27	3	47	59	19.56	-36.69	1.25	4.38	6
19	2	27	3	48	59	19.56	-36.63	1.06	4.31	6
19	2	27	3	49	59	19.56	-36.56	0.75	4.38	6
19	2	27	3	50	59	19.56	-36.63	1.19	4.31	5.94
19	2	27	3	52	0	19.62	-36.56	0.94	4.31	5.94
19	2	27	3	53	0	19.62	-36.56	1.31	4.31	6
19	2	27	3	54	0	19.62	-36.56	1.12	4.31	6
19	2	27	3	55	0	19.56	-36.56	0.75	4.38	6
19	2	27	3	56	1	19.56	-36.56	1.25	4.38	6
19	2	27	3	57	1	19.5	-36.5	1	4.38	6
19	2	27	3	58	1	19.5	-36.56	1.37	4.31	6
19	2	27	3	59	1	19.44	-36.56	1.12	4.38	6
19	2	27	4	0	2	19.44	-36.56	0.75	4.38	6
19	2	27	4	1	2	19.37	-36.56	1.31	4.38	6
19	2	27	4	2	2	19.37	-36.56	1.06	4.38	6
19	2	27	4	3	2	19.37	-36.63	1.44	4.38	6
19	2	27	4	4	3	19.31	-36.56	1.19	4.38	6
19	2	27	4	5	3	19.31	-36.56	0.75	4.38	6
19	2	27	4	6	3	19.31	-36.56	1.31	4.38	6
19	2	27	4	7	3	19.31	-36.5	1	4.38	6
19	2	27	4	8	4	19.37	-36.56	1.44	4.38	6
19	2	27	4	9	4	19.37	-36.56	1.19	4.38	6
19	2	27	4	10	4	19.37	-36.56	1.5	4.38	6
19	2	27	4	11	4	19.37	-36.69	0.88	4.38	6
19	2	27	4	12	5	19.44	-36.63	0.63	4.38	6
19	2	27	4	13	5	19.44	-36.69	1.12	4.38	5.94
19	2	27	4	14	5	19.44	-36.63	0.88	4.38	5.94
19	2	27	4	15	5	19.5	-36.63	1.25	4.31	5.94
19	2	27	4	16	6	19.5	-36.63	1.06	4.31	5.94
19	2	27	4	17	6	19.5	-36.63	0.69	4.31	5.94
19	2	27	4	18	6	19.56	-36.63	1.25	4.31	5.94
19	2	27	4	19	6	19.56	-36.56	0.94	4.31	5.94
19	2	27	4	20	7	19.56	-36.63	1.31	4.31	5.94
19	2	27	4	21	7	19.62	-36.56	1.12	4.31	5.94
19	2	27	4	22	7	19.62	-36.5	0.69	4.31	5.94
19	2	27	4	23	7	19.62	-36.56	1.25	4.31	5.94
19	2	27	4	24	8	19.62	-36.5	0.94	4.31	5.94
19	2	27	4	25	8	19.62	-36.56	1.37	4.31	5.94
19	2	27	4	26	8	19.62	-36.56	1.12	4.38	5.94
19	2	27	4	27	8	19.69	-36.56	0.69	4.31	5.94
19	2	27	4	28	9	19.69	-36.5	1.25	4.31	6
19	2	27	4	29	9	19.69	-36.5	1	4.38	6
19	2	27	4	30	9	19.69	-36.5	1.37	4.38	6
19	2	27	4	31	9	19.69	-36.56	1.12	4.38	6
19	2	27	4	32	10	19.62	-36.56	0.75	4.38	6
19	2	27	4	33	10	19.62	-36.56	1.31	4.38	6
19	2	27	4	34	10	19.62	-36.5	1	4.38	6
19	2	27	4	35	11	19.56	-36.5	1.37	4.38	6
19	2	27	4	36	11	19.56	-36.5	1.19	4.38	6
19	2	27	4	37	11	19.5	-36.5	1.5	4.38	6
19	2	27	4	38	11	19.5	-36.56	0.88	4.38	6
19	2	27	4	39	12	19.5	-36.56	0.63	4.38	6
19	2	27	4	40	12	19.5	-36.63	1.12	4.31	6
19	2	27	4	41	12	19.44	-36.63	0.88	4.31	5.94
19	2	27	4	42	12	19.44	-36.63	1.25	4.31	5.94
19	2	27	4	43	13	19.5	-36.63	1.06	4.31	5.94
19	2	27	4	44	13	19.5	-36.56	0.75	4.31	5.94
19	2	27	4	45	13	19.5	-36.63	1.19	4.31	5.94
19	2	27	4	46	13	19.5	-36.56	1	4.31	6
19	2	27	4	47	14	19.5	-36.56	1.37	4.31	6
19	2	27	4	48	14	19.5	-36.56	1.12	4.38	6
19	2	27	4	49	14	19.5	-36.56	0.75	4.31	6
19	2	27	4	50	14	19.5	-36.56	1.25	4.38	6
19	2	27	4	51	15	19.5	-36.56	1	4.38	6
19	2	27	4	52	15	19.56	-36.63	1.37	4.38	6
19	2	27	4	53	15	19.56	-36.56	1.12	4.38	6
19	2	27	4	54	15	19.56	-36.56	0.69	4.38	6
19	2	27	4	55	16	19.56	-36.56	1.25	4.38	6
19	2	27	4	56	16	19.56	-36.5	1	4.38	6
19	2	27	4	57	16	19.56	-36.56	1.37	4.38	6
19	2	27	4	58	16	19.56	-36.5	1.19	4.38	6
19	2	27	4	59	17	19.56	-36.56	0.75	4.38	6
19	2	27	5	0	17	19.56	-36.56	1.31	4.38	6
19	2	27	5	1	17	19.5	-36.5	1	4.38	6
19	2	27	5	2	17	19.5	-36.56	1.37	4.38	6
19	2	27	5	3	18	19.44	-36.5	1.12	4.38	6
19	2	27	5	4	18	19.44	-36.56	1.44	4.38	6
19	2	27	5	5	18	19.37	-36.56	0.88	4.38	6
19	2	27	5	6	18	19.37	-36.56	0.63	4.38	6
19	2	27	5	7	19	19.37	-36.56	1.12	4.38	6
19	2	27	5	8	19	19.37	-36.56	0.88	4.38	5.94
19	2	27	5	9	19	19.37	-36.56	1.25	4.31	5.94
19	2	27	5	10	19	19.37	-36.5	1.06	4.38	5.94
19	2	27	5	11	20	19.37	-36.5	0.69	4.38	6
19	2	27	5	12	20	19.44	-36.5	1.19	4.38	5.94
19	2	27	5	13	20	19.44	-36.44	0.94	4.38	6
19	2	27	5	14	20	19.44	-36.5	1.31	4.38	5.94
19	2	27	5	15	21	19.5	-36.5	1.12	4.38	6
19	2	27	5	16	21	19.5	-36.5	0.75	4.38	6
19	2	27	5	17	21	19.5	-36.56	1.25	4.38	6
19	2	27	5	18	21	19.5	-36.5	1	4.38	6
19	2	27	5	19	22	19.56	-36.5	1.44	4.38	6
19	2	27	5	20	22	19.56	-36.5	1.12	4.38	6
19	2	27	5	21	22	19.56	-36.5	0.75	4.38	6
19	2	27	5	22	22	19.56	-36.5	1.31	4.38	6
19	2	27	5	23	23	19.56	-36.44	1.06	4.38	6
19	2	27	5	24	23	19.56	-36.5	1.44	4.38	6
19	2	27	5	25	23	19.56	-36.44	1.19	4.38	6
19	2	27	5	26	23	19.62	-36.44	0.81	4.38	6
19	2	27	5	27	24	19.62	-36.44	1.31	4.38	6
19	2	27	5	28	24	19.62	-36.38	1.06	4.38	6.06
19	2	27	5	29	24	19.62	-36.44	1.44	4.38	6.06
19	2	27	5	30	24	19.56	-36.44	1.19	4.38	6.06
19	2	27	5	31	25	19.56	-36.44	1.5	4.38	6.06
19	2	27	5	32	25	19.5	-36.5	0.94	4.38	6.06
19	2	27	5	33	25	19.5	-36.5	0.69	4.38	6
19	2	27	5	34	25	19.5	-36.5	1.12	4.38	6
19	2	27	5	35	26	19.44	-36.5	0.88	4.38	6
19	2	27	5	36	26	19.44	-36.5	1.31	4.38	6
19	2	27	5	37	26	19.44	-36.44	1.06	4.38	6
19	2	27	5	38	26	19.37	-36.38	0.75	4.38	6
19	2	27	5	39	27	19.37	-36.44	1.25	4.38	6
19	2	27	5	40	27	19.37	-36.44	1	4.38	6
19	2	27	5	41	27	19.37	-36.44	1.37	4.38	6
19	2	27	5	42	27	19.37	-36.5	1.19	4.38	6
19	2	27	5	43	28	19.37	-36.44	0.81	4.38	6
19	2	27	5	44	28	19.37	-36.5	1.31	4.38	6
19	2	27	5	45	28	19.37	-36.44	1.06	4.38	6
19	2	27	5	46	28	19.37	-36.5	1.44	4.38	6
19	2	27	5	47	29	19.44	-36.5	1.19	4.38	6
19	2	27	5	48	29	19.44	-36.5	0.75	4.38	6
19	2	27	5	49	29	19.44	-36.5	1.31	4.38	6
19	2	27	5	50	29	19.44	-36.44	1.06	4.38	6
19	2	27	5	51	30	19.44	-36.5	1.44	4.38	6
19	2	27	5	52	30	19.44	-36.44	1.19	4.38	6
19	2	27	5	53	30	19.5	-36.44	0.81	4.38	6.06
19	2	27	5	54	30	19.44	-36.44	1.37	4.38	6
19	2	27	5	55	31	19.5	-36.38	1	4.44	6.06
19	2	27	5	56	31	19.5	-36.44	1.5	4.44	6.06
19	2	27	5	57	31	19.5	-36.44	1.25	4.44	6.06
19	2	27	5	58	31	19.5	-36.44	1.56	4.44	6.06
19	2	27	5	59	32	19.5	-36.44	1	4.44	6.06
19	2	27	6	0	32	19.5	-36.44	0.75	4.38	6
19	2	27	6	1	32	19.5	-36.5	1.25	4.38	6
19	2	27	6	2	32	19.5	-36.44	1	4.44	6
19	2	27	6	3	33	19.56	-36.44	1.37	4.44	6
19	2	27	6	4	33	19.56	-36.44	1.19	4.38	6
19	2	27	6	5	33	19.56	-36.38	0.81	4.44	6.06
19	2	27	6	6	33	19.56	-36.44	1.31	4.44	6
19	2	27	6	7	34	19.56	-36.31	1.06	4.44	6
19	2	27	6	8	34	19.62	-36.38	1.44	4.44	6
19	2	27	6	9	34	19.62	-36.38	1.19	4.44	6
19	2	27	6	10	34	19.62	-36.38	0.81	4.44	6.06
19	2	27	6	11	35	19.56	-36.44	1.37	4.44	6
19	2	27	6	12	35	19.62	-36.38	1.06	4.44	6.06
19	2	27	6	13	35	19.62	-36.44	1.5	4.44	6.06
19	2	27	6	14	35	19.56	-36.38	1.25	4.44	6.06
19	2	27	6	15	36	19.56	-36.38	0.81	4.44	6.06
19	2	27	6	16	36	19.56	-36.44	1.37	4.44	6.06
19	2	27	6	17	36	19.56	-36.38	1.12	4.44	6.06
19	2	27	6	18	36	19.5	-36.38	1.5	4.44	6.06
19	2	27	6	19	37	19.56	-36.38	1.25	4.44	6.06
19	2	27	6	20	37	19.56	-36.38	0.88	4.44	6.06
19	2	27	6	21	37	19.56	-36.38	1.37	4.44	6.06
19	2	27	6	22	37	19.56	-36.31	1.06	4.44	6.06
19	2	27	6	23	38	19.56	-36.38	1.5	4.44	6.06
19	2	27	6	24	38	19.56	-36.38	1.31	4.44	6.06
19	2	27	6	25	38	19.56	-36.38	1.56	4.44	6.13
19	2	27	6	26	38	19.62	-36.44	1	4.44	6.06
19	2	27	6	27	39	19.62	-36.44	0.69	4.44	6.06
19	2	27	6	28	39	19.62	-36.5	1.19	4.44	6.06
19	2	27	6	29	39	19.62	-36.38	0.94	4.44	6.06
19	2	27	6	30	39	19.62	-36.38	1.37	4.44	6.06
19	2	27	6	31	40	19.62	-36.38	1.12	4.44	6
19	2	27	6	32	40	19.62	-36.38	0.81	4.44	6
19	2	27	6	33	40	19.62	-36.38	1.31	4.44	6
19	2	27	6	34	40	19.62	-36.31	1.06	4.44	6
19	2	27	6	35	41	19.56	-36.31	1.44	4.44	6
19	2	27	6	36	41	19.56	-36.31	1.19	4.44	6
19	2	27	6	37	41	19.56	-36.31	0.81	4.44	6.06
19	2	27	6	38	42	19.5	-36.31	1.37	4.44	6.06
19	2	27	6	39	42	19.44	-36.31	1.06	4.44	6.06
19	2	27	6	40	42	19.44	-36.31	1.44	4.44	6.06
19	2	27	6	41	42	19.44	-36.31	1.25	4.44	6.06
19	2	27	6	42	43	19.44	-36.31	0.81	4.44	6.06
19	2	27	6	43	43	19.44	-36.38	1.37	4.44	6.06
19	2	27	6	44	43	19.44	-36.31	1.12	4.44	6.06
19	2	27	6	45	43	19.44	-36.31	1.5	4.44	6.06
19	2	27	6	46	44	19.44	-36.31	1.25	4.44	6.06
19	2	27	6	47	44	19.44	-36.31	0.88	4.44	6.06
19	2	27	6	48	44	19.44	-36.38	1.44	4.44	6.06
19	2	27	6	49	44	19.44	-36.31	1.06	4.44	6.13
19	2	27	6	50	45	19.5	-36.38	1.5	4.44	6.06
19	2	27	6	51	45	19.5	-36.38	1.25	4.44	6.13
19	2	27	6	52	45	19.56	-36.38	1.56	4.44	6.13
19	2	27	6	53	45	19.62	-36.44	1	4.44	6.13
19	2	27	6	54	46	19.62	-36.44	0.69	4.44	6.06
19	2	27	6	55	46	19.69	-36.44	1.19	4.44	6.06
19	2	27	6	56	46	19.69	-36.38	1	4.44	6.06
19	2	27	6	57	46	19.69	-36.44	1.37	4.44	6.06
19	2	27	6	58	47	19.75	-36.38	1.12	4.44	6.06
19	2	27	6	59	47	19.75	-36.38	0.81	4.44	6.06
19	2	27	7	0	47	19.75	-36.38	1.31	4.44	6.06
19	2	27	7	1	47	19.75	-36.31	1.06	4.44	6.06
19	2	27	7	2	48	19.75	-36.38	1.44	4.44	6.06
19	2	27	7	3	48	19.69	-36.38	1.25	4.44	6.06
19	2	27	7	4	48	19.69	-36.38	0.88	4.44	6.06
19	2	27	7	5	48	19.69	-36.38	1.44	4.44	6.06
19	2	27	7	6	49	19.62	-36.31	1.12	4.44	6.06
19	2	27	7	7	49	19.62	-36.38	1.5	4.44	6.06
19	2	27	7	8	49	19.56	-36.31	1.25	4.44	6.06
19	2	27	7	9	49	19.5	-36.31	0.81	4.44	6.06
19	2	27	7	10	50	19.5	-36.31	1.37	4.44	6.06
19	2	27	7	11	50	19.5	-36.25	1.12	4.44	6.06
19	2	27	7	12	50	19.44	-36.31	1.5	4.44	6.06
19	2	27	7	13	50	19.44	-36.31	1.25	4.5	6.06
19	2	27	7	14	51	19.44	-36.31	0.81	4.44	6.06
19	2	27	7	15	51	19.44	-36.31	1.37	4.44	6.06
19	2	27	7	16	51	19.5	-36.31	1.12	4.44	6.13
19	2	27	7	17	51	19.5	-36.38	1.5	4.5	6.13
19	2	27	7	18	52	19.5	-36.31	1.31	4.5	6.13
19	2	27	7	19	52	19.56	-36.31	1.56	4.5	6.13
19	2	27	7	20	52	19.56	-36.38	1.06	4.5	6.13
19	2	27	7	21	52	19.56	-36.38	0.75	4.5	6.06
19	2	27	7	22	53	19.56	-36.38	1.25	4.5	6.06
19	2	27	7	23	53	19.5	-36.38	1	4.44	6.06
19	2	27	7	24	53	19.5	-36.38	1.44	4.44	6.06
19	2	27	7	25	53	19.5	-36.38	1.19	4.44	6.06
19	2	27	7	26	54	19.5	-36.31	0.88	4.44	6.06
19	2	27	7	27	54	19.5	-36.38	1.37	4.44	6.06
19	2	27	7	28	54	19.5	-36.31	1.12	4.44	6.06
19	2	27	7	29	54	19.5	-36.31	1.5	4.44	6.06
19	2	27	7	30	55	19.5	-36.31	1.31	4.5	6.06
19	2	27	7	31	55	19.5	-36.31	0.94	4.5	6.06
19	2	27	7	32	55	19.5	-36.31	1.44	4.44	6.06
19	2	27	7	33	55	19.5	-36.25	1.19	4.5	6.06
19	2	27	7	34	56	19.56	-36.25	1.56	4.5	6.13
19	2	27	7	35	56	19.56	-36.25	1.37	4.5	6.13
19	2	27	7	36	56	19.56	-36.25	0.94	4.5	6.13
19	2	27	7	37	56	19.62	-36.25	1.5	4.5	6.13
19	2	27	7	38	57	19.62	-36.25	1.25	4.5	6.13
19	2	27	7	39	57	19.62	-36.31	1.63	4.5	6.13
19	2	27	7	40	57	19.69	-36.31	1.37	4.5	6.13
19	2	27	7	41	57	19.69	-36.31	0.94	4.5	6.13
19	2	27	7	42	58	19.69	-36.31	1.5	4.5	6.13
19	2	27	7	43	58	19.75	-36.25	1.19	4.5	6.13
19	2	27	7	44	58	19.69	-36.31	1.63	4.5	6.13
19	2	27	7	45	58	19.75	-36.31	1.37	4.5	6.19
19	2	27	7	46	59	19.75	-36.31	1.69	4.5	6.19
19	2	27	7	47	59	19.75	-36.38	1.06	4.5	6.13
19	2	27	7	48	59	19.81	-36.38	0.81	4.5	6.13
19	2	27	7	50	0	19.81	-36.38	1.31	4.5	6.13
19	2	27	7	51	0	19.81	-36.31	1.06	4.5	6.13
19	2	27	7	52	0	19.87	-36.31	1.44	4.5	6.13
19	2	27	7	53	0	19.87	-36.31	1.25	4.5	6.13
19	2	27	7	54	1	19.87	-36.25	0.94	4.5	6.13
19	2	27	7	55	1	19.81	-36.31	1.44	4.5	6.13
19	2	27	7	56	1	19.81	-36.19	1.19	4.5	6.13
19	2	27	7	57	1	19.81	-36.25	1.56	4.5	6.13
19	2	27	7	58	2	19.81	-36.25	1.37	4.5	6.13
19	2	27	7	59	2	19.75	-36.19	1	4.5	6.13
19	2	27	8	0	2	19.69	-36.25	1.5	4.5	6.13
19	2	27	8	1	2	19.69	-36.19	1.25	4.5	6.13
19	2	27	8	2	3	19.69	-36.25	1.63	4.5	6.13
19	2	27	8	3	3	19.69	-36.25	1.37	4.5	6.13
19	2	27	8	4	3	19.69	-36.25	0.94	4.5	6.13
19	2	27	8	5	3	19.69	-36.25	1.5	4.5	6.13
19	2	27	8	6	4	19.69	-36.25	1.25	4.56	6.19
19	2	27	8	7	4	19.69	-36.25	1.63	4.56	6.19
19	2	27	8	8	4	19.75	-36.25	1.44	4.56	6.19
19	2	27	8	9	4	19.75	-36.25	1	4.56	6.19
19	2	27	8	10	5	19.75	-36.19	1.56	4.56	6.19
19	2	27	8	11	5	19.81	-36.19	1.25	4.56	6.19
19	2	27	8	12	5	19.81	-36.25	1.63	4.56	6.19
19	2	27	8	13	5	19.81	-36.19	1.44	4.56	6.19
19	2	27	8	14	6	19.87	-36.19	1.75	4.56	6.19
19	2	27	8	15	6	19.87	-36.25	1.19	4.56	6.19
19	2	27	8	16	6	19.94	-36.19	0.88	4.56	6.19
19	2	27	8	17	6	19.94	-36.25	1.37	4.56	6.13
19	2	27	8	18	7	20	-36.19	1.19	4.56	6.13
19	2	27	8	19	7	20	-36.25	1.56	4.56	6.19
19	2	27	8	20	7	20	-36.25	1.37	4.56	6.13
19	2	27	8	21	7	19.94	-36.19	1.06	4.56	6.13
19	2	27	8	22	8	19.94	-36.25	1.56	4.56	6.19
19	2	27	8	23	8	19.87	-36.19	1.31	4.56	6.19
19	2	27	8	24	8	19.87	-36.25	1.69	4.56	6.19
19	2	27	8	25	8	19.81	-36.19	1.44	4.56	6.19
19	2	27	8	26	9	19.81	-36.19	1.12	4.56	6.19
19	2	27	8	27	9	19.75	-36.19	1.63	4.56	6.19
19	2	27	8	28	9	19.75	-36.19	1.31	4.56	6.19
19	2	27	8	29	9	19.75	-36.25	1.69	4.56	6.19
19	2	27	8	30	10	19.75	-36.19	1.5	4.56	6.19
19	2	27	8	31	10	19.75	-36.19	1.06	4.63	6.19
19	2	27	8	32	10	19.75	-36.25	1.63	4.56	6.19
19	2	27	8	33	10	19.75	-36.19	1.31	4.63	6.25
19	2	27	8	34	11	19.75	-36.19	1.69	4.63	6.25
19	2	27	8	35	11	19.75	-36.19	1.5	4.63	6.25
19	2	27	8	36	11	19.75	-36.25	1.06	4.63	6.25
19	2	27	8	37	11	19.81	-36.25	1.63	4.63	6.25
19	2	27	8	38	12	19.81	-36.19	1.31	4.63	6.25
19	2	27	8	39	12	19.81	-36.25	1.75	4.63	6.25
19	2	27	8	40	12	19.81	-36.19	1.44	4.63	6.25
19	2	27	8	41	12	19.81	-36.19	1.81	4.63	6.25
19	2	27	8	42	13	19.81	-36.31	1.19	4.63	6.25
19	2	27	8	43	13	19.81	-36.25	0.88	4.63	6.25
19	2	27	8	44	13	19.81	-36.31	1.37	4.63	6.19
19	2	27	8	45	13	19.87	-36.25	1.19	4.63	6.19
19	2	27	8	46	14	19.87	-36.25	1.56	4.63	6.19
19	2	27	8	47	14	19.87	-36.25	1.31	4.63	6.19
19	2	27	8	48	14	19.87	-36.19	1	4.63	6.19
19	2	27	8	49	14	19.87	-36.25	1.5	4.56	6.19
19	2	27	8	50	15	19.81	-36.19	1.25	4.56	6.19
19	2	27	8	51	15	19.81	-36.19	1.63	4.56	6.19
19	2	27	8	52	15	19.75	-36.25	1.44	4.63	6.19
19	2	27	8	53	15	19.75	-36.19	1.06	4.56	6.25
19	2	27	8	54	16	19.69	-36.19	1.56	4.63	6.19
19	2	27	8	55	16	19.69	-36.19	1.25	4.63	6.25
19	2	27	8	56	16	19.62	-36.19	1.69	4.63	6.25
19	2	27	8	57	17	19.62	-36.19	1.44	4.63	6.25
19	2	27	8	58	17	19.62	-36.19	1	4.63	6.25
19	2	27	8	59	17	19.62	-36.19	1.56	4.63	6.25
19	2	27	9	0	17	19.62	-36.13	1.31	4.63	6.25
19	2	27	9	1	18	19.62	-36.19	1.69	4.63	6.25
19	2	27	9	2	18	19.62	-36.13	1.44	4.63	6.25
19	2	27	9	3	18	19.62	-36.19	1.06	4.63	6.25
19	2	27	9	4	18	19.62	-36.13	1.63	4.63	6.25
19	2	27	9	5	19	19.62	-36.13	1.31	4.63	6.25
19	2	27	9	6	19	19.62	-36.13	1.69	4.63	6.25
19	2	27	9	7	19	19.62	-36.13	1.5	4.63	6.25
19	2	27	9	8	19	19.69	-36.13	1.75	4.63	6.31
19	2	27	9	9	20	19.69	-36.19	1.19	4.63	6.25
19	2	27	9	10	20	19.69	-36.19	0.88	4.63	6.25
19	2	27	9	11	20	19.69	-36.25	1.37	4.63	6.19
19	2	27	9	12	20	19.75	-36.19	1.12	4.63	6.19
19	2	27	9	13	21	19.75	-36.19	1.56	4.56	6.19
19	2	27	9	14	21	19.75	-36.19	1.31	4.63	6.19
19	2	27	9	15	21	19.81	-36.19	1	4.63	6.19
19	2	27	9	16	21	19.81	-36.19	1.5	4.56	6.19
19	2	27	9	17	22	19.87	-36.19	1.25	4.63	6.19
19	2	27	9	18	22	19.87	-36.19	1.63	4.56	6.19
19	2	27	9	19	22	19.87	-36.19	1.44	4.63	6.19
19	2	27	9	20	22	19.87	-36.19	1.06	4.56	6.19
19	2	27	9	21	23	19.94	-36.19	1.56	4.63	6.19
19	2	27	9	22	23	19.94	-36.13	1.31	4.63	6.19
19	2	27	9	23	23	19.94	-36.19	1.69	4.56	6.25
19	2	27	9	24	23	20	-36.13	1.5	4.63	6.25
19	2	27	9	25	24	19.94	-36.13	1.06	4.63	6.25
19	2	27	9	26	24	20	-36.13	1.63	4.63	6.25
19	2	27	9	27	24	20	-36.06	1.37	4.63	6.25
19	2	27	9	28	24	20	-36.13	1.75	4.63	6.25
19	2	27	9	29	25	20	-36.13	1.5	4.63	6.25
19	2	27	9	30	25	20	-36.13	1.12	4.63	6.25
19	2	27	9	31	25	20	-36.13	1.69	4.63	6.25
19	2	27	9	32	25	20	-36.06	1.37	4.63	6.25
19	2	27	9	33	26	20	-36.13	1.75	4.63	6.25
19	2	27	9	34	26	20.06	-36.06	1.56	4.69	6.31
19	2	27	9	35	26	20.06	-36.06	1.81	4.63	6.31
19	2	27	9	36	26	20.06	-36.19	1.25	4.63	6.25
19	2	27	9	37	27	20.06	-36.13	1	4.63	6.25
19	2	27	9	38	27	20.06	-36.19	1.5	4.63	6.25
19	2	27	9	39	27	20.06	-36.13	1.25	4.63	6.25
19	2	27	9	40	27	20.06	-36.13	1.63	4.63	6.25
19	2	27	9	41	28	20.06	-36.13	1.44	4.63	6.25
19	2	27	9	42	28	20	-36.13	1.12	4.63	6.25
19	2	27	9	43	28	20	-36.13	1.56	4.63	6.25
19	2	27	9	44	28	19.94	-36.13	1.31	4.63	6.25
19	2	27	9	45	29	19.94	-36.13	1.69	4.63	6.25
19	2	27	9	46	29	19.87	-36.13	1.5	4.63	6.25
19	2	27	9	47	29	19.87	-36.13	1.12	4.63	6.25
19	2	27	9	48	29	19.81	-36.13	1.63	4.63	6.25
19	2	27	9	49	30	19.75	-36.06	1.37	4.63	6.25
19	2	27	9	50	30	19.75	-36.13	1.75	4.69	6.25
19	2	27	9	51	30	19.69	-36.13	1.5	4.63	6.25
19	2	27	9	52	30	19.69	-36.13	1.12	4.63	6.25
19	2	27	9	53	31	19.62	-36.13	1.69	4.63	6.25
19	2	27	9	54	31	19.62	-36.06	1.44	4.69	6.31
19	2	27	9	55	31	19.62	-36.13	1.81	4.69	6.31
19	2	27	9	56	31	19.62	-36.06	1.56	4.69	6.31
19	2	27	9	57	32	19.62	-36.13	1.19	4.69	6.31
19	2	27	9	58	32	19.62	-36.13	1.69	4.69	6.31
19	2	27	9	59	32	19.62	-36.06	1.44	4.69	6.31
19	2	27	10	0	32	19.62	-36.13	1.81	4.69	6.31
19	2	27	10	1	33	19.62	-36.06	1.63	4.69	6.31
19	2	27	10	2	33	19.62	-36.13	1.94	4.69	6.31
19	2	27	10	3	33	19.69	-36.19	1.31	4.69	6.31
19	2	27	10	4	34	19.62	-36.19	1.06	4.69	6.31
19	2	27	10	5	34	19.69	-36.19	1.5	4.69	6.31
19	2	27	10	6	34	19.69	-36.13	1.31	4.69	6.25
19	2	27	10	7	34	19.69	-36.19	1.69	4.69	6.25
19	2	27	10	8	35	19.69	-36.13	1.5	4.69	6.25
19	2	27	10	9	35	19.75	-36.06	1.19	4.69	6.25
19	2	27	10	10	35	19.75	-36.13	1.63	4.69	6.25
19	2	27	10	11	35	19.81	-36.06	1.44	4.69	6.25
19	2	27	10	12	36	19.81	-36.13	1.81	4.69	6.25
19	2	27	10	13	36	19.81	-36.06	1.63	4.69	6.25
19	2	27	10	14	36	19.81	-36.06	1.19	4.69	6.31
19	2	27	10	15	36	19.81	-36.06	1.69	4.63	6.31
19	2	27	10	16	37	19.87	-36.06	1.5	4.69	6.31
19	2	27	10	17	37	19.87	-36.06	1.81	4.69	6.31
19	2	27	10	18	37	19.87	-36.06	1.63	4.69	6.31
19	2	27	10	19	37	19.81	-36.06	1.19	4.69	6.31
19	2	27	10	20	38	19.81	-36.06	1.75	4.69	6.31
19	2	27	10	21	38	19.75	-36	1.44	4.69	6.31
19	2	27	10	22	38	19.75	-36.06	1.81	4.69	6.31
19	2	27	10	23	38	19.69	-36.06	1.56	4.69	6.31
19	2	27	10	24	39	19.69	-36.06	1.19	4.69	6.31
19	2	27	10	25	39	19.69	-36.06	1.75	4.69	6.31
19	2	27	10	26	39	19.62	-36	1.44	4.69	6.31
19	2	27	10	27	39	19.62	-36.06	1.81	4.69	6.31
19	2	27	10	28	40	19.62	-36	1.56	4.69	6.38
19	2	27	10	29	40	19.69	-36	1.88	4.69	6.38
19	2	27	10	30	40	19.69	-36.06	1.31	4.69	6.31
19	2	27	10	31	40	19.69	-36.06	1	4.69	6.31
19	2	27	10	32	41	19.69	-36.13	1.5	4.69	6.31
19	2	27	10	33	41	19.69	-36.06	1.25	4.69	6.31
19	2	27	10	34	41	19.75	-36.06	1.69	4.69	6.31
19	2	27	10	35	41	19.75	-36.06	1.5	4.69	6.31
19	2	27	10	36	42	19.75	-36	1.12	4.69	6.31
19	2	27	10	37	42	19.75	-36.06	1.63	4.69	6.31
19	2	27	10	38	42	19.75	-36	1.37	4.69	6.25
19	2	27	10	39	42	19.81	-36.06	1.75	4.69	6.31
19	2	27	10	40	43	19.81	-36.06	1.56	4.69	6.31
19	2	27	10	41	43	19.81	-36.06	1.19	4.63	6.31
19	2	27	10	42	43	19.81	-36.06	1.69	4.69	6.25
19	2	27	10	43	43	19.81	-36	1.44	4.69	6.31
19	2	27	10	44	44	19.81	-36.06	1.81	4.69	6.31
19	2	27	10	45	44	19.81	-36.06	1.63	4.69	6.31
19	2	27	10	46	44	19.87	-36.06	1.19	4.69	6.31
19	2	27	10	47	44	19.87	-36.06	1.75	4.69	6.31
19	2	27	10	48	45	19.87	-36	1.44	4.69	6.31
19	2	27	10	49	45	19.87	-36.06	1.88	4.69	6.31
19	2	27	10	50	45	19.87	-36.06	1.63	4.69	6.31
19	2	27	10	51	45	19.87	-36.06	1.19	4.69	6.31
19	2	27	10	52	46	19.87	-36	1.75	4.69	6.31
19	2	27	10	53	46	19.87	-36	1.44	4.69	6.31
19	2	27	10	54	46	19.87	-36	1.88	4.69	6.31
19	2	27	10	55	46	19.87	-36	1.56	4.69	6.31
19	2	27	10	56	47	19.87	-36	1.88	4.69	6.38
19	2	27	10	57	47	19.87	-36.06	1.31	4.69	6.31
19	2	27	10	58	47	19.87	-36.06	1.06	4.69	6.31
19	2	27	10	59	47	19.87	-36.06	1.5	4.69	6.31
19	2	27	11	0	48	19.87	-36.06	1.31	4.69	6.31
19	2	27	11	1	48	19.87	-36.06	1.69	4.69	6.31
19	2	27	11	2	48	19.81	-36.06	1.5	4.69	6.31
19	2	27	11	3	48	19.81	-36	1.19	4.69	6.31
19	2	27	11	4	49	19.81	-36.06	1.69	4.69	6.31
19	2	27	11	5	49	19.75	-36	1.44	4.69	6.31
19	2	27	11	6	49	19.75	-36.06	1.81	4.69	6.31
19	2	27	11	7	49	19.75	-36.06	1.63	4.69	6.31
19	2	27	11	8	50	19.75	-36	1.19	4.69	6.31
19	2	27	11	9	50	19.75	-36	1.75	4.69	6.31
19	2	27	11	10	50	19.75	-35.94	1.44	4.69	6.31
19	2	27	11	11	51	19.75	-36	1.81	4.69	6.31
19	2	27	11	12	51	19.75	-35.94	1.63	4.69	6.31
19	2	27	11	13	51	19.75	-35.94	1.19	4.69	6.31
19	2	27	11	14	51	19.81	-36	1.75	4.69	6.31
19	2	27	11	15	52	19.81	-35.94	1.5	4.69	6.31
19	2	27	11	16	52	19.81	-36	1.88	4.69	6.31
19	2	27	11	17	52	19.81	-36	1.63	4.69	6.38
19	2	27	11	18	52	19.87	-36	1.25	4.75	6.38
19	2	27	11	19	53	19.87	-36	1.75	4.75	6.38
19	2	27	11	20	53	19.87	-36	1.44	4.75	6.38
19	2	27	11	21	53	19.87	-36	1.88	4.75	6.38
19	2	27	11	22	53	19.87	-36	1.63	4.75	6.38
19	2	27	11	23	54	19.87	-36	1.94	4.75	6.38
19	2	27	11	24	54	19.87	-36	1.37	4.75	6.31
19	2	27	11	25	54	19.94	-36	1.06	4.69	6.31
19	2	27	11	26	54	19.87	-36	1.56	4.75	6.31
19	2	27	11	27	55	19.87	-36	1.31	4.69	6.31
19	2	27	11	28	55	19.94	-36	1.75	4.69	6.31
19	2	27	11	29	55	19.94	-36	1.56	4.69	6.31
19	2	27	11	30	55	19.94	-35.94	1.19	4.69	6.31
19	2	27	11	31	56	19.94	-36	1.69	4.69	6.31
19	2	27	11	32	56	19.94	-36	1.44	4.69	6.31
19	2	27	11	33	56	19.94	-36	1.81	4.69	6.31
19	2	27	11	34	56	19.94	-35.94	1.56	4.69	6.31
19	2	27	11	35	57	19.94	-36	1.19	4.69	6.31
19	2	27	11	36	57	19.87	-36	1.69	4.69	6.31
19	2	27	11	37	57	19.87	-35.94	1.44	4.69	6.31
19	2	27	11	38	57	19.87	-36	1.81	4.69	6.31
19	2	27	11	39	58	19.81	-36	1.63	4.69	6.31
19	2	27	11	40	58	19.81	-36	1.19	4.69	6.31
19	2	27	11	41	58	19.75	-36	1.75	4.69	6.31
19	2	27	11	42	58	19.69	-36	1.44	4.69	6.38
19	2	27	11	43	59	19.69	-36	1.88	4.69	6.31
19	2	27	11	44	59	19.62	-36	1.63	4.75	6.31
19	2	27	11	45	59	19.62	-36	1.19	4.75	6.38
19	2	27	11	46	59	19.62	-36	1.75	4.69	6.38
19	2	27	11	48	0	19.62	-35.94	1.44	4.75	6.38
19	2	27	11	49	0	19.62	-36	1.88	4.75	6.38
19	2	27	11	50	0	19.62	-35.94	1.63	4.75	6.38
19	2	27	11	51	0	19.62	-35.94	1.94	4.75	6.38
19	2	27	11	52	1	19.62	-36	1.31	4.75	6.31
19	2	27	11	53	1	19.69	-35.94	1.06	4.69	6.31
19	2	27	11	54	1	19.69	-36	1.56	4.69	6.31
19	2	27	11	55	1	19.69	-35.94	1.31	4.69	6.31
19	2	27	11	56	2	19.69	-35.94	1.75	4.69	6.31
19	2	27	11	57	2	19.69	-35.94	1.5	4.69	6.31
19	2	27	11	58	2	19.69	-35.94	1.19	4.69	6.31
19	2	27	11	59	2	19.69	-35.94	1.69	4.69	6.31
19	2	27	12	0	3	19.69	-35.94	1.44	4.69	6.31
19	2	27	12	1	3	19.69	-35.94	1.81	4.69	6.31
19	2	27	12	2	3	19.69	-35.94	1.56	4.69	6.31
19	2	27	12	3	3	19.75	-35.94	1.19	4.69	6.31
19	2	27	12	4	4	19.75	-35.94	1.69	4.69	6.31
19	2	27	12	5	4	19.75	-35.88	1.44	4.69	6.31
19	2	27	12	6	4	19.75	-35.94	1.81	4.69	6.31
19	2	27	12	7	4	19.75	-35.94	1.63	4.69	6.31
19	2	27	12	8	5	19.75	-35.94	1.19	4.69	6.31
19	2	27	12	9	5	19.75	-35.94	1.75	4.69	6.31
19	2	27	12	10	5	19.75	-35.88	1.44	4.75	6.31
19	2	27	12	11	5	19.75	-35.94	1.81	4.69	6.31
19	2	27	12	12	6	19.69	-35.88	1.63	4.75	6.38
19	2	27	12	13	6	19.69	-35.88	1.19	4.75	6.38
19	2	27	12	14	6	19.62	-35.88	1.75	4.75	6.31
19	2	27	12	15	6	19.62	-35.88	1.44	4.75	6.38
19	2	27	12	16	7	19.62	-35.94	1.88	4.75	6.38
19	2	27	12	17	7	19.62	-35.88	1.63	4.75	6.38
19	2	27	12	18	7	19.56	-35.88	1.94	4.75	6.38
19	2	27	12	19	8	19.56	-35.94	1.37	4.75	6.31
19	2	27	12	20	8	19.56	-35.94	1.06	4.75	6.31
19	2	27	12	21	8	19.62	-35.94	1.56	4.75	6.31
19	2	27	12	22	8	19.56	-35.94	1.37	4.69	6.31
19	2	27	12	23	9	19.62	-35.94	1.75	4.69	6.31
19	2	27	12	24	9	19.62	-35.88	1.56	4.69	6.31
19	2	27	12	25	9	19.62	-35.88	1.25	4.69	6.31
19	2	27	12	26	9	19.62	-35.88	1.69	4.69	6.31
19	2	27	12	27	10	19.62	-35.88	1.44	4.69	6.31
19	2	27	12	28	10	19.62	-35.94	1.81	4.69	6.31
19	2	27	12	29	10	19.62	-35.88	1.56	4.69	6.31
19	2	27	12	30	10	19.62	-35.88	1.25	4.75	6.31
19	2	27	12	31	11	19.69	-35.88	1.75	4.69	6.31
19	2	27	12	32	11	19.69	-35.81	1.5	4.69	6.31
19	2	27	12	33	11	19.69	-35.88	1.88	4.75	6.31
19	2	27	12	34	11	19.69	-35.88	1.63	4.75	6.31
19	2	27	12	35	12	19.69	-35.88	1.19	4.75	6.38
19	2	27	12	36	12	19.69	-35.94	1.75	4.75	6.31
19	2	27	12	37	12	19.69	-35.88	1.5	4.75	6.38
19	2	27	12	38	12	19.69	-35.88	1.88	4.75	6.31
19	2	27	12	39	13	19.69	-35.88	1.69	4.75	6.38
19	2	27	12	40	13	19.75	-35.88	1.25	4.75	6.38
19	2	27	12	41	13	19.75	-35.88	1.81	4.75	6.38
19	2	27	12	42	13	19.75	-35.81	1.5	4.75	6.38
19	2	27	12	43	14	19.75	-35.88	1.94	4.75	6.38
19	2	27	12	44	14	19.75	-35.88	1.69	4.75	6.38
19	2	27	12	45	14	19.75	-35.88	2	4.75	6.38
19	2	27	12	46	14	19.75	-35.88	1.37	4.75	6.38
19	2	27	12	47	15	19.75	-35.88	1.12	4.75	6.38
19	2	27	12	48	15	19.75	-35.94	1.63	4.75	6.31
19	2	27	12	49	15	19.75	-35.94	1.37	4.75	6.31
19	2	27	12	50	15	19.69	-35.94	1.75	4.75	6.31
19	2	27	12	51	16	19.69	-35.94	1.56	4.75	6.31
19	2	27	12	52	16	19.62	-35.88	1.25	4.75	6.31
19	2	27	12	53	16	19.62	-35.94	1.75	4.75	6.31
19	2	27	12	54	16	19.56	-35.88	1.5	4.75	6.31
19	2	27	12	55	17	19.56	-35.94	1.81	4.75	6.31
19	2	27	12	56	17	19.5	-35.88	1.69	4.75	6.31
19	2	27	12	57	17	19.5	-35.88	1.25	4.75	6.38
19	2	27	12	58	17	19.5	-35.94	1.75	4.75	6.31
19	2	27	12	59	18	19.5	-35.88	1.5	4.75	6.38
19	2	27	13	0	18	19.5	-35.88	1.88	4.75	6.31
19	2	27	13	1	18	19.5	-35.88	1.69	4.75	6.31
19	2	27	13	2	18	19.56	-35.88	1.25	4.75	6.38
19	2	27	13	3	19	19.56	-35.88	1.81	4.75	6.38
19	2	27	13	4	19	19.56	-35.88	1.56	4.75	6.38
19	2	27	13	5	19	19.62	-35.88	1.94	4.75	6.38
19	2	27	13	6	19	19.62	-35.88	1.69	4.75	6.38
19	2	27	13	7	20	19.62	-35.88	1.25	4.75	6.38
19	2	27	13	8	20	19.69	-35.88	1.75	4.75	6.38
19	2	27	13	9	20	19.69	-35.81	1.5	4.75	6.38
19	2	27	13	10	20	19.69	-35.88	1.94	4.75	6.38
19	2	27	13	11	21	19.69	-35.81	1.69	4.75	6.38
19	2	27	13	12	21	19.69	-35.81	2	4.75	6.38
19	2	27	13	13	21	19.69	-35.94	1.44	4.75	6.38
19	2	27	13	14	21	19.69	-35.88	1.12	4.75	6.38
19	2	27	13	15	22	19.62	-35.94	1.63	4.75	6.31
19	2	27	13	16	22	19.62	-35.88	1.44	4.75	6.31
19	2	27	13	17	22	19.56	-35.94	1.75	4.75	6.31
19	2	27	13	18	22	19.56	-35.94	1.63	4.69	6.31
19	2	27	13	19	23	19.56	-35.88	1.25	4.69	6.31
19	2	27	13	20	23	19.5	-35.94	1.75	4.75	6.31
19	2	27	13	21	23	19.5	-35.88	1.5	4.75	6.31
19	2	27	13	22	23	19.5	-35.88	1.88	4.69	6.31
19	2	27	13	23	24	19.56	-35.81	1.69	4.75	6.31
19	2	27	13	24	24	19.56	-35.88	1.25	4.75	6.31
19	2	27	13	25	24	19.56	-35.88	1.81	4.75	6.31
19	2	27	13	26	24	19.56	-35.88	1.5	4.75	6.38
19	2	27	13	27	25	19.56	-35.88	1.88	4.75	6.38
19	2	27	13	28	25	19.56	-35.88	1.69	4.75	6.38
19	2	27	13	29	25	19.56	-35.88	1.25	4.75	6.38
19	2	27	13	30	26	19.62	-35.88	1.81	4.75	6.38
19	2	27	13	31	26	19.62	-35.81	1.56	4.75	6.38
19	2	27	13	32	26	19.62	-35.81	1.94	4.75	6.38
19	2	27	13	33	26	19.62	-35.88	1.69	4.75	6.38
19	2	27	13	34	27	19.62	-35.88	1.31	4.75	6.38
19	2	27	13	35	27	19.62	-35.88	1.81	4.75	6.38
19	2	27	13	36	27	19.62	-35.81	1.56	4.75	6.38
19	2	27	13	37	27	19.69	-35.81	1.94	4.75	6.38
19	2	27	13	38	28	19.69	-35.88	1.69	4.75	6.38
19	2	27	13	39	28	19.69	-35.81	2	4.75	6.44
19	2	27	13	40	28	19.69	-35.94	1.44	4.75	6.38
19	2	27	13	41	28	19.69	-35.88	1.12	4.75	6.38
19	2	27	13	42	29	19.69	-35.88	1.63	4.75	6.38
19	2	27	13	43	29	19.69	-35.88	1.37	4.75	6.31
19	2	27	13	44	29	19.69	-35.88	1.75	4.75	6.31
19	2	27	13	45	29	19.75	-35.88	1.56	4.75	6.31
19	2	27	13	46	30	19.75	-35.88	1.25	4.75	6.31
19	2	27	13	47	30	19.75	-35.88	1.75	4.75	6.31
19	2	27	13	48	30	19.75	-35.81	1.5	4.75	6.31
19	2	27	13	49	30	19.75	-35.81	1.88	4.69	6.31
19	2	27	13	50	31	19.75	-35.88	1.69	4.75	6.31
19	2	27	13	51	31	19.75	-35.81	1.31	4.75	6.31
19	2	27	13	52	31	19.81	-35.88	1.75	4.75	6.31
19	2	27	13	53	31	19.81	-35.81	1.5	4.75	6.38
19	2	27	13	54	32	19.81	-35.88	1.88	4.75	6.38
19	2	27	13	55	32	19.81	-35.88	1.63	4.75	6.31
19	2	27	13	56	32	19.87	-35.88	1.19	4.75	6.38
19	2	27	13	57	32	19.87	-35.81	1.75	4.75	6.31
19	2	27	13	58	33	19.87	-35.81	1.5	4.75	6.38
19	2	27	13	59	33	19.87	-35.81	1.94	4.75	6.38
19	2	27	14	0	33	19.87	-35.88	1.69	4.75	6.38
19	2	27	14	1	33	19.87	-35.88	1.25	4.75	6.38
19	2	27	14	2	34	19.87	-35.81	1.81	4.75	6.38
19	2	27	14	3	34	19.87	-35.75	1.5	4.75	6.38
19	2	27	14	4	34	19.81	-35.81	1.94	4.75	6.38
19	2	27	14	5	34	19.81	-35.81	1.69	4.75	6.38
19	2	27	14	6	35	19.75	-35.81	2	4.75	6.38
19	2	27	14	7	35	19.75	-35.88	1.37	4.75	6.38
19	2	27	14	8	35	19.75	-35.88	1.12	4.75	6.31
19	2	27	14	9	35	19.75	-35.88	1.63	4.75	6.31
19	2	27	14	10	36	19.69	-35.88	1.37	4.69	6.31
19	2	27	14	11	36	19.75	-35.88	1.75	4.75	6.31
19	2	27	14	12	36	19.69	-35.88	1.56	4.69	6.31
19	2	27	14	13	36	19.75	-35.81	1.25	4.69	6.31
19	2	27	14	14	37	19.75	-35.88	1.69	4.69	6.31
19	2	27	14	15	37	19.75	-35.81	1.44	4.69	6.31
19	2	27	14	16	37	19.75	-35.88	1.81	4.69	6.31
19	2	27	14	17	37	19.75	-35.88	1.56	4.69	6.31
19	2	27	14	18	38	19.75	-35.81	1.25	4.69	6.31
19	2	27	14	19	38	19.75	-35.81	1.75	4.69	6.31
19	2	27	14	20	38	19.75	-35.75	1.5	4.69	6.31
19	2	27	14	21	38	19.75	-35.81	1.88	4.69	6.31
19	2	27	14	22	39	19.75	-35.81	1.63	4.69	6.31
19	2	27	14	23	39	19.75	-35.81	1.25	4.75	6.31
19	2	27	14	24	39	19.81	-35.88	1.75	4.75	6.31
19	2	27	14	25	39	19.81	-35.81	1.5	4.75	6.31
19	2	27	14	26	40	19.75	-35.88	1.88	4.75	6.31
19	2	27	14	27	40	19.75	-35.81	1.63	4.75	6.38
19	2	27	14	28	40	19.69	-35.81	1.25	4.75	6.38
19	2	27	14	29	40	19.69	-35.75	1.75	4.75	6.38
19	2	27	14	30	41	19.62	-35.75	1.5	4.75	6.38
19	2	27	14	31	41	19.62	-35.81	1.88	4.75	6.38
19	2	27	14	32	41	19.56	-35.81	1.63	4.75	6.38
19	2	27	14	33	41	19.56	-35.81	1.94	4.75	6.38
19	2	27	14	34	42	19.56	-35.88	1.37	4.75	6.38
19	2	27	14	35	42	19.56	-35.88	1.06	4.75	6.31
19	2	27	14	36	42	19.56	-35.88	1.56	4.75	6.31
19	2	27	14	37	43	19.5	-35.88	1.31	4.75	6.31
19	2	27	14	38	43	19.5	-35.88	1.75	4.69	6.31
19	2	27	14	39	43	19.5	-35.88	1.56	4.69	6.31
19	2	27	14	40	43	19.5	-35.81	1.19	4.69	6.31
19	2	27	14	41	44	19.56	-35.88	1.69	4.69	6.31
19	2	27	14	42	44	19.56	-35.81	1.44	4.75	6.31
19	2	27	14	43	44	19.56	-35.81	1.81	4.69	6.31
19	2	27	14	44	44	19.56	-35.81	1.56	4.69	6.31
19	2	27	14	45	45	19.56	-35.81	1.25	4.69	6.31
19	2	27	14	46	45	19.56	-35.81	1.75	4.69	6.31
19	2	27	14	47	45	19.62	-35.81	1.5	4.69	6.31
19	2	27	14	48	45	19.62	-35.88	1.88	4.69	6.31
19	2	27	14	49	46	19.62	-35.81	1.69	4.75	6.31
19	2	27	14	50	46	19.69	-35.81	1.25	4.69	6.31
19	2	27	14	51	46	19.69	-35.88	1.81	4.75	6.31
19	2	27	14	52	46	19.69	-35.81	1.5	4.75	6.38
19	2	27	14	53	47	19.75	-35.81	1.88	4.75	6.38
19	2	27	14	54	47	19.75	-35.75	1.69	4.75	6.38
19	2	27	14	55	47	19.75	-35.81	1.25	4.75	6.38
19	2	27	14	56	47	19.75	-35.81	1.81	4.75	6.38
19	2	27	14	57	48	19.75	-35.75	1.56	4.75	6.38
19	2	27	14	58	48	19.75	-35.81	1.94	4.75	6.38
19	2	27	14	59	48	19.75	-35.75	1.69	4.75	6.38
19	2	27	15	0	48	19.75	-35.75	2	4.75	6.38
19	2	27	15	1	49	19.75	-35.88	1.37	4.75	6.38
19	2	27	15	2	49	19.75	-35.81	1.06	4.75	6.31
19	2	27	15	3	49	19.69	-35.88	1.56	4.75	6.31
19	2	27	15	4	49	19.62	-35.81	1.37	4.75	6.31
19	2	27	15	5	50	19.62	-35.81	1.75	4.69	6.31
19	2	27	15	6	50	19.56	-35.81	1.56	4.69	6.31
19	2	27	15	7	50	19.56	-35.75	1.25	4.75	6.31
19	2	27	15	8	50	19.56	-35.81	1.75	4.69	6.31
19	2	27	15	9	51	19.56	-35.75	1.5	4.69	6.31
19	2	27	15	10	51	19.5	-35.75	1.88	4.69	6.31
19	2	27	15	11	51	19.5	-35.75	1.69	4.75	6.31
19	2	27	15	12	51	19.5	-35.75	1.31	4.75	6.31
19	2	27	15	13	52	19.5	-35.75	1.81	4.69	6.31
19	2	27	15	14	52	19.5	-35.69	1.56	4.75	6.31
19	2	27	15	15	52	19.56	-35.75	1.94	4.75	6.38
19	2	27	15	16	52	19.56	-35.75	1.69	4.75	6.38
19	2	27	15	17	53	19.56	-35.75	1.25	4.75	6.38
19	2	27	15	18	53	19.56	-35.75	1.81	4.75	6.38
19	2	27	15	19	53	19.56	-35.75	1.5	4.75	6.38
19	2	27	15	20	53	19.56	-35.75	1.94	4.75	6.38
19	2	27	15	21	54	19.56	-35.75	1.69	4.75	6.38
19	2	27	15	22	54	19.56	-35.75	1.25	4.75	6.38
19	2	27	15	23	54	19.62	-35.75	1.81	4.75	6.38
19	2	27	15	24	54	19.62	-35.75	1.5	4.75	6.38
19	2	27	15	25	55	19.62	-35.81	1.94	4.75	6.38
19	2	27	15	26	55	19.62	-35.81	1.69	4.75	6.38
19	2	27	15	27	55	19.62	-35.75	2	4.75	6.38
19	2	27	15	28	55	19.62	-35.81	1.44	4.75	6.38
19	2	27	15	29	56	19.62	-35.81	1.12	4.75	6.38
19	2	27	15	30	56	19.62	-35.81	1.63	4.75	6.31
19	2	27	15	31	56	19.62	-35.81	1.44	4.75	6.31
19	2	27	15	32	56	19.62	-35.81	1.81	4.75	6.31
19	2	27	15	33	57	19.69	-35.75	1.63	4.75	6.31
19	2	27	15	34	57	19.62	-35.75	1.25	4.75	6.31
19	2	27	15	35	57	19.62	-35.81	1.75	4.75	6.31
19	2	27	15	36	57	19.69	-35.75	1.5	4.75	6.31
19	2	27	15	37	58	19.62	-35.81	1.88	4.75	6.31
19	2	27	15	38	58	19.69	-35.81	1.69	4.75	6.31
19	2	27	15	39	58	19.69	-35.75	1.31	4.75	6.31
19	2	27	15	40	58	19.69	-35.81	1.81	4.75	6.31
19	2	27	15	41	59	19.69	-35.75	1.5	4.75	6.38
19	2	27	15	42	59	19.75	-35.81	1.88	4.75	6.38
19	2	27	15	43	59	19.75	-35.75	1.69	4.75	6.38
19	2	27	15	44	59	19.75	-35.75	1.25	4.75	6.38
19	2	27	15	46	0	19.69	-35.75	1.81	4.75	6.38
19	2	27	15	47	0	19.75	-35.69	1.5	4.75	6.38
19	2	27	15	48	0	19.69	-35.69	1.94	4.75	6.38
19	2	27	15	49	1	19.69	-35.69	1.69	4.75	6.38
19	2	27	15	50	1	19.69	-35.69	1.31	4.75	6.38
19	2	27	15	51	1	19.69	-35.75	1.81	4.75	6.38
19	2	27	15	52	1	19.62	-35.69	1.5	4.75	6.38
19	2	27	15	53	2	19.62	-35.75	1.94	4.75	6.38
19	2	27	15	54	2	19.62	-35.75	1.69	4.75	6.38
19	2	27	15	55	2	19.62	-35.75	2	4.75	6.38
19	2	27	15	56	2	19.62	-35.75	1.5	4.75	6.38
19	2	27	15	57	3	19.62	-35.75	1.19	4.75	6.38
19	2	27	15	58	3	19.62	-35.75	1.69	4.75	6.38
19	2	27	15	59	3	19.62	-35.75	1.44	4.75	6.38
19	2	27	16	0	3	19.62	-35.75	1.81	4.75	6.38
19	2	27	16	1	4	19.62	-35.75	1.63	4.75	6.31
19	2	27	16	2	4	19.62	-35.69	1.31	4.75	6.38
19	2	27	16	3	4	19.62	-35.69	1.75	4.75	6.31
19	2	27	16	4	4	19.69	-35.75	1.5	4.75	6.38
19	2	27	16	5	5	19.62	-35.75	1.88	4.75	6.31
19	2	27	16	6	5	19.69	-35.75	1.69	4.75	6.38
19	2	27	16	7	5	19.69	-35.75	1.37	4.75	6.38
19	2	27	16	8	5	19.69	-35.75	1.81	4.75	6.38
19	2	27	16	9	6	19.69	-35.69	1.56	4.75	6.38
19	2	27	16	10	6	19.69	-35.75	1.94	4.75	6.38
19	2	27	16	11	6	19.69	-35.69	1.75	4.75	6.38
19	2	27	16	12	6	19.69	-35.69	1.31	4.75	6.38
19	2	27	16	13	7	19.69	-35.75	1.88	4.75	6.38
19	2	27	16	14	7	19.69	-35.69	1.63	4.75	6.38
19	2	27	16	15	7	19.69	-35.69	2	4.75	6.38
19	2	27	16	16	7	19.69	-35.69	1.75	4.75	6.38
19	2	27	16	17	8	19.69	-35.69	1.37	4.75	6.38
19	2	27	16	18	8	19.75	-35.75	1.88	4.81	6.38
19	2	27	16	19	8	19.69	-35.69	1.63	4.75	6.44
19	2	27	16	20	8	19.69	-35.69	2	4.81	6.44
19	2	27	16	21	9	19.69	-35.63	1.81	4.81	6.44
19	2	27	16	22	9	19.62	-35.63	2.06	4.81	6.44
19	2	27	16	23	9	19.62	-35.69	1.5	4.81	6.44
19	2	27	16	24	9	19.56	-35.69	1.25	4.81	6.38
19	2	27	16	25	10	19.5	-35.69	1.69	4.75	6.38
19	2	27	16	26	10	19.5	-35.69	1.5	4.75	6.38
19	2	27	16	27	10	19.5	-35.69	1.88	4.75	6.38
19	2	27	16	28	11	19.44	-35.69	1.69	4.75	6.38
19	2	27	16	29	11	19.44	-35.63	1.37	4.75	6.38
19	2	27	16	30	11	19.44	-35.69	1.81	4.75	6.38
19	2	27	16	31	11	19.44	-35.69	1.56	4.75	6.38
19	2	27	16	32	12	19.44	-35.69	1.94	4.75	6.38
19	2	27	16	33	12	19.44	-35.69	1.75	4.75	6.38
19	2	27	16	34	12	19.44	-35.69	1.37	4.75	6.38
19	2	27	16	35	12	19.5	-35.69	1.88	4.75	6.38
19	2	27	16	36	13	19.44	-35.63	1.63	4.75	6.38
19	2	27	16	37	13	19.5	-35.63	2	4.75	6.38
19	2	27	16	38	13	19.5	-35.63	1.75	4.75	6.38
19	2	27	16	39	13	19.5	-35.63	1.37	4.75	6.44
19	2	27	16	40	14	19.5	-35.69	1.88	4.81	6.38
19	2	27	16	41	14	19.56	-35.63	1.63	4.81	6.38
19	2	27	16	42	14	19.56	-35.69	2	4.81	6.44
19	2	27	16	43	14	19.62	-35.63	1.81	4.81	6.44
19	2	27	16	44	15	19.62	-35.63	1.37	4.81	6.44
19	2	27	16	45	15	19.62	-35.63	1.88	4.81	6.44
19	2	27	16	46	15	19.69	-35.56	1.56	4.81	6.44
19	2	27	16	47	15	19.69	-35.56	2	4.81	6.44
19	2	27	16	48	16	19.69	-35.56	1.75	4.81	6.44
19	2	27	16	49	16	19.69	-35.63	2.06	4.81	6.44
19	2	27	16	50	16	19.75	-35.69	1.5	4.81	6.38
19	2	27	16	51	16	19.75	-35.69	1.25	4.81	6.38
19	2	27	16	52	17	19.75	-35.69	1.69	4.81	6.38
19	2	27	16	53	17	19.69	-35.63	1.5	4.81	6.38
19	2	27	16	54	17	19.69	-35.69	1.88	4.75	6.38
19	2	27	16	55	17	19.62	-35.69	1.63	4.75	6.38
19	2	27	16	56	18	19.62	-35.63	1.31	4.81	6.38
19	2	27	16	57	18	19.56	-35.69	1.81	4.75	6.38
19	2	27	16	58	18	19.56	-35.69	1.56	4.81	6.38
19	2	27	16	59	18	19.5	-35.69	1.94	4.81	6.38
19	2	27	17	0	19	19.44	-35.63	1.75	4.81	6.38
19	2	27	17	1	19	19.44	-35.63	1.37	4.75	6.38
19	2	27	17	2	19	19.44	-35.63	1.88	4.75	6.38
19	2	27	17	3	19	19.44	-35.63	1.56	4.81	6.38
19	2	27	17	4	20	19.44	-35.63	1.94	4.75	6.38
19	2	27	17	5	20	19.44	-35.63	1.75	4.75	6.38
19	2	27	17	6	20	19.44	-35.63	1.31	4.75	6.38
19	2	27	17	7	20	19.44	-35.63	1.88	4.75	6.38
19	2	27	17	8	21	19.44	-35.56	1.63	4.81	6.44
19	2	27	17	9	21	19.44	-35.63	2	4.81	6.44
19	2	27	17	10	21	19.44	-35.63	1.75	4.81	6.44
19	2	27	17	11	21	19.44	-35.63	1.31	4.81	6.44
19	2	27	17	12	22	19.44	-35.63	1.88	4.81	6.44
19	2	27	17	13	22	19.44	-35.56	1.56	4.81	6.44
19	2	27	17	14	22	19.5	-35.63	2	4.81	6.44
19	2	27	17	15	22	19.5	-35.63	1.75	4.81	6.44
19	2	27	17	16	23	19.5	-35.63	2.06	4.81	6.44
19	2	27	17	17	23	19.5	-35.63	1.5	4.81	6.44
19	2	27	17	18	23	19.5	-35.63	1.25	4.81	6.44
19	2	27	17	19	23	19.5	-35.63	1.75	4.81	6.38
19	2	27	17	20	24	19.5	-35.63	1.5	4.81	6.38
19	2	27	17	21	24	19.5	-35.63	1.88	4.81	6.38
19	2	27	17	22	24	19.5	-35.63	1.69	4.75	6.38
19	2	27	17	23	24	19.5	-35.63	1.31	4.81	6.38
19	2	27	17	24	25	19.5	-35.63	1.81	4.81	6.38
19	2	27	17	25	25	19.44	-35.56	1.56	4.81	6.38
19	2	27	17	26	25	19.44	-35.56	1.94	4.81	6.38
19	2	27	17	27	26	19.37	-35.56	1.69	4.75	6.38
19	2	27	17	28	26	19.37	-35.56	1.31	4.81	6.38
19	2	27	17	29	26	19.31	-35.56	1.88	4.75	6.38
19	2	27	17	30	26	19.31	-35.56	1.63	4.75	6.38
19	2	27	17	31	27	19.31	-35.56	2	4.75	6.38
19	2	27	17	32	27	19.31	-35.56	1.75	4.75	6.44
19	2	27	17	33	27	19.31	-35.56	1.31	4.81	6.38
19	2	27	17	34	27	19.31	-35.56	1.88	4.81	6.38
19	2	27	17	35	28	19.37	-35.56	1.63	4.81	6.38
19	2	27	17	36	28	19.37	-35.56	2	4.81	6.38
19	2	27	17	37	28	19.37	-35.56	1.75	4.81	6.38
19	2	27	17	38	28	19.44	-35.63	1.31	4.81	6.44
19	2	27	17	39	29	19.44	-35.56	1.88	4.81	6.38
19	2	27	17	40	29	19.44	-35.5	1.56	4.81	6.44
19	2	27	17	41	29	19.5	-35.56	2	4.81	6.44
19	2	27	17	42	29	19.5	-35.56	1.75	4.81	6.44
19	2	27	17	43	30	19.5	-35.56	2.06	4.81	6.44
19	2	27	17	44	30	19.56	-35.69	1.5	4.81	6.38
19	2	27	17	45	30	19.56	-35.63	1.19	4.81	6.38
19	2	27	17	46	30	19.56	-35.69	1.69	4.81	6.38
19	2	27	17	47	31	19.56	-35.69	1.44	4.81	6.38
19	2	27	17	48	31	19.62	-35.69	1.88	4.75	6.38
19	2	27	17	49	31	19.56	-35.63	1.69	4.75	6.38
19	2	27	17	50	31	19.56	-35.63	1.31	4.75	6.38
19	2	27	17	51	32	19.56	-35.63	1.81	4.75	6.38
19	2	27	17	52	32	19.56	-35.56	1.56	4.75	6.38
19	2	27	17	53	32	19.5	-35.63	1.94	4.75	6.38
19	2	27	17	54	32	19.5	-35.56	1.75	4.75	6.38
19	2	27	17	55	33	19.44	-35.5	1.37	4.75	6.38
19	2	27	17	56	33	19.44	-35.56	1.88	4.75	6.38
19	2	27	17	57	33	19.37	-35.56	1.63	4.75	6.38
19	2	27	17	58	33	19.37	-35.56	2	4.75	6.38
19	2	27	17	59	34	19.37	-35.56	1.75	4.81	6.38
19	2	27	18	0	34	19.37	-35.56	1.37	4.81	6.44
19	2	27	18	1	34	19.37	-35.56	1.88	4.81	6.44
19	2	27	18	2	34	19.37	-35.56	1.63	4.81	6.38
19	2	27	18	3	35	19.37	-35.56	2	4.81	6.44
19	2	27	18	4	35	19.37	-35.56	1.75	4.81	6.38
19	2	27	18	5	35	19.37	-35.56	1.37	4.81	6.44
19	2	27	18	6	35	24.56	-35.19	1.94	4.63	5.06
19	2	27	18	7	36	19.37	-35.5	1.56	4.81	6.44
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	2	27	18	9	3	19.37	-35.44	2.13	4.81	6.44
19	2	27	18	10	3	19.44	-35.44	2	4.81	6.44
19	2	27	18	11	4	19.44	-35.44	2.31	4.88	6.5
19	2	27	18	12	4	19.44	-35.44	2.06	4.88	6.44
19	2	27	18	13	4	19.44	-35.44	1.75	4.88	6.5
19	2	27	18	14	4	19.37	-35.5	2.13	4.88	6.5
19	2	27	18	15	5	19.37	-35.5	1.88	4.88	6.5
19	2	27	18	16	5	19.37	-35.56	2.19	4.88	6.5
19	2	27	18	17	5	19.37	-35.56	2	4.88	6.5
19	2	27	18	18	5	19.31	-35.56	1.63	4.88	6.5
19	2	27	18	19	6	19.31	-35.56	2.06	4.88	6.5
19	2	27	18	20	6	19.31	-35.56	1.81	4.88	6.5
19	2	27	18	21	6	19.31	-35.56	2.19	4.88	6.5
19	2	27	18	22	6	19.31	-35.56	2	4.94	6.5
19	2	27	18	23	7	19.31	-35.56	1.56	4.88	6.5
19	2	27	18	24	7	19.31	-35.56	2.06	4.88	6.5
19	2	27	18	25	7	19.37	-35.5	1.81	4.88	6.5
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	2	27	18	27	23	19.37	-35.5	2.5	4.94	6.5
19	2	27	18	28	23	19.44	-35.44	2.56	4.94	6.56
19	2	27	18	29	23	19.44	-35.5	2.81	5	85
19	2	27	18	30	23	19.5	-35.56	3	5.19	85
19	2	27	18	31	24	19.5	-35.5	3.19	5.44	7.38
19	2	27	18	32	24	19.56	-35.56	3.44	5.69	7.69
19	2	27	18	33	24	19.56	-35.5	3.63	5.88	7.88
19	2	27	18	34	24	19.56	-35.56	3.88	6.06	8
19	2	27	18	35	25	19.56	-35.5	4	6.25	8.13
19	2	27	18	36	25	19.5	-35.56	4.25	6.5	8.25
19	2	27	18	37	25	19.5	-35.63	4.44	6.63	8.38
19	2	27	18	38	25	19.5	-35.56	4.5	6.81	8.44
19	2	27	18	39	26	19.5	-35.63	4.69	7	8.5
19	2	27	18	40	26	19.56	-35.5	4.75	7.13	8.63
19	2	27	18	41	26	19.5	-35.5	4.81	7.25	8.69
19	2	27	18	42	26	19.56	-35.5	5	7.38	8.75
19	2	27	18	43	27	19.56	-35.44	5.06	7.5	8.81
19	2	27	18	44	27	19.56	-35.5	5.25	7.69	8.88
19	2	27	18	45	27	19.56	-35.44	5.31	7.75	8.94
19	2	27	18	46	28	19.56	-35.44	5.38	7.88	9
19	2	27	18	47	28	19.62	-35.5	5.56	8	9.06
19	2	27	18	48	28	19.62	-35.5	5.63	8.06	9.13
19	2	27	18	49	28	19.62	-35.56	5.81	8.13	9.19
19	2	27	18	50	29	19.62	-35.5	5.88	8.19	9.25
19	2	27	18	51	29	19.69	-35.44	5.94	8.31	9.25
19	2	27	18	52	29	19.69	-35.44	6.06	8.31	9.31
19	2	27	18	53	29	19.69	-35.44	6.13	8.38	9.38
19	2	27	18	54	30	19.69	-35.44	6.31	8.44	9.38
19	2	27	18	55	30	19.69	-35.44	6.38	8.5	9.44
19	2	27	18	56	30	19.62	-35.38	6.38	8.56	9.5
19	2	27	18	57	30	19.62	-35.44	6.56	8.63	9.5
19	2	27	18	58	31	19.62	-35.31	6.25	8.63	9.5
19	2	27	18	59	31	19.62	-35.38	6.44	8.69	9.56
19	2	27	19	0	31	19.56	-35.31	6.25	8.69	9.56
19	2	27	19	1	31	19.56	-35.31	6.31	8.75	9.56
19	2	27	19	2	32	19.56	-35.31	6.56	8.81	9.56
19	2	27	19	3	32	19.56	-35.31	6.25	8.81	9.56
19	2	27	19	4	32	19.56	-35.38	6.5	8.88	9.56
19	2	27	19	5	32	19.56	-35.31	6.25	8.88	9.56
19	2	27	19	6	33	19.56	-35.25	6.38	8.88	9.56
19	2	27	19	7	33	19.56	-35.25	6.63	8.94	9.56
19	2	27	19	8	33	19.56	-35.25	6.31	8.94	9.56
19	2	27	19	9	33	19.56	-35.31	6.56	9	9.56
19	2	27	19	10	34	19.62	-35.25	6.38	9	9.56
19	2	27	19	11	34	19.62	-35.25	6	9	9.56
19	2	27	19	12	34	19.62	-35.19	6.38	9	9.56
19	2	27	19	13	34	19.69	-35.13	6.13	9.06	9.56
19	2	27	19	14	35	19.69	-35.19	6.38	9.06	9.56
19	2	27	19	15	35	19.69	-35.19	6.19	9.06	9.56
19	2	27	19	16	35	19.69	-35.19	5.81	9.06	9.56
19	2	27	19	17	36	19.75	-35.19	6.25	9.06	9.56
19	2	27	19	18	36	19.75	-35.13	5.75	9.06	9.56
19	2	27	19	19	36	19.81	-35.25	6.13	9.06	9.56
19	2	27	19	20	36	19.81	-35.25	5.69	9.06	9.56
19	2	27	19	21	37	19.81	-35.25	5.31	9.13	9.56
19	2	27	19	22	37	19.81	-35.19	5.88	9.13	9.56
19	2	27	19	23	37	19.81	-35.19	5.38	9.13	9.56
19	2	27	19	24	37	19.81	-35.25	5.81	9.13	9.56
19	2	27	19	25	38	19.81	-35.25	5.38	9.13	9.56
19	2	27	19	26	38	19.75	-35.31	5	9.13	9.5
19	2	27	19	27	38	19.81	-35.38	5.44	9.13	9.5
19	2	27	19	28	38	19.81	-35.31	4.69	9.13	9.5
19	2	27	19	29	39	19.75	-35.38	5.25	9.13	9.5
19	2	27	19	30	39	19.81	-35.44	4.63	9.13	9.5
19	2	27	19	31	39	19.81	-35.44	5.13	9.13	9.5
19	2	27	19	32	39	19.81	-35.38	5.06	9.13	9.5
19	2	27	19	33	40	19.81	-35.31	4.63	9.13	9.5
19	2	27	19	34	40	19.87	-35.31	5.19	9.13	9.5
19	2	27	19	35	40	19.87	-35.31	4.5	9.13	9.5
19	2	27	19	36	40	19.87	-35.38	5.06	9.13	9.5
19	2	27	19	37	41	19.94	-35.38	4.31	9.06	9.5
19	2	27	19	38	41	19.94	-35.44	3.63	9.06	9.44
19	2	27	19	39	41	19.94	-35.38	4.38	9.06	9.44
19	2	27	19	40	41	19.94	-35.38	3.63	9.06	9.44
19	2	27	19	41	42	20	-35.44	4.31	9.06	9.38
19	2	27	19	42	42	20	-35.44	3.63	9.06	9.38
19	2	27	19	43	42	20	-35.5	2.94	9	9.38
19	2	27	19	44	43	20.06	-35.38	3.81	9	9.38
19	2	27	19	45	43	20.06	-35.44	3.06	9	9.38
19	2	27	19	46	43	20.06	-35.38	3.81	9	9.38
19	2	27	19	47	43	20	-35.38	3	9	9.31
19	2	27	19	48	44	20	-35.38	3.69	9	9.31
19	2	27	19	49	44	20	-35.38	3.56	9	9.31
19	2	27	19	50	44	19.94	-35.38	2.69	8.94	9.31
19	2	27	19	51	44	19.94	-35.38	3.5	8.94	9.31
19	2	27	19	52	45	19.87	-35.31	2.69	8.94	9.25
19	2	27	19	53	45	19.87	-35.31	3.38	8.94	9.25
19	2	27	19	54	45	19.81	-35.38	2.63	8.94	9.25
19	2	27	19	55	45	19.81	-35.38	1.81	8.88	9.25
19	2	27	19	56	46	19.81	-35.31	2.88	8.88	9.19
19	2	27	19	57	46	19.81	-35.38	1.81	8.88	9.19
19	2	27	19	58	46	19.75	-35.31	2.69	8.88	9.19
19	2	27	19	59	46	19.81	-35.31	1.81	8.88	9.19
19	2	27	20	0	47	19.81	-35.31	2.5	8.81	9.13
19	2	27	20	1	47	19.81	-35.31	1.94	8.81	9.13
19	2	27	20	2	47	19.81	-35.38	1.12	8.81	9.13
19	2	27	20	3	47	19.81	-35.31	2.13	8.81	9.06
19	2	27	20	4	48	19.87	-35.31	1.25	8.75	9.06
19	2	27	20	5	48	19.87	-35.25	2.13	8.75	9.06
19	2	27	20	6	48	19.81	-35.25	1.37	8.75	9.06
19	2	27	20	7	49	19.81	-35.38	0.25	8.75	9
19	2	27	20	8	49	19.81	-35.25	1.5	8.75	9
19	2	27	20	9	49	19.75	-35.25	0.37	8.69	9
19	2	27	20	10	49	19.75	-35.19	1.44	8.69	8.94
19	2	27	20	11	50	19.69	-35.19	0.56	8.69	8.94
19	2	27	20	12	50	19.69	-35.13	1.5	8.69	8.94
19	2	27	20	13	50	19.69	-35.13	1	8.69	8.88
19	2	27	20	14	50	19.69	-35.19	0	8.69	8.88
19	2	27	20	15	51	19.69	-35.13	1.12	8.63	8.88
19	2	27	20	16	51	19.69	-35.13	-0.12	8.63	8.88
19	2	27	20	17	51	19.69	-35.13	0.94	8.63	8.81
19	2	27	20	18	51	19.69	-35.13	-0.06	8.63	8.81
19	2	27	20	19	52	19.75	-35.19	-0.25	8.56	8.81
19	2	27	20	20	52	19.75	-35.06	-0.25	8.56	8.75
19	2	27	20	21	52	19.75	-35.13	-1.44	8.56	8.69
19	2	27	20	22	52	19.75	-35.06	-0.12	8.5	8.69
19	2	27	20	23	53	19.81	-35.06	-1.06	8.5	8.63
19	2	27	20	24	53	19.81	-35.06	0	8.5	8.63
19	2	27	20	25	53	19.81	-35.06	-0.81	8.44	8.56
19	2	27	20	26	53	19.87	-35.25	-2.13	8.5	8.56
19	2	27	20	27	54	19.81	-35	-0.56	8.44	8.56
19	2	27	20	28	54	19.87	-35.06	-1.75	8.44	8.5
19	2	27	20	29	54	19.81	-35	-0.5	8.44	8.5
19	2	27	20	30	54	19.81	-35	-1.44	8.44	8.5
19	2	27	20	31	55	19.81	-34.94	-0.31	8.38	8.5
19	2	27	20	32	55	19.81	-34.94	-1.25	8.38	8.44
19	2	27	20	33	55	19.81	-35.06	-2.44	8.38	8.44
19	2	27	20	34	55	19.81	-34.94	-0.88	8.38	8.44
19	2	27	20	35	56	19.81	-34.94	-2	8.31	8.38
19	2	27	20	36	56	19.81	-34.88	-0.75	8.31	8.38
19	2	27	20	37	56	19.81	-34.88	-1.81	8.31	8.38
19	2	27	20	38	57	19.81	-34.81	-0.69	8.31	8.38
19	2	27	20	39	57	19.81	-34.81	-1.63	8.31	8.31
19	2	27	20	40	57	19.81	-34.94	-2.88	8.25	8.31
19	2	27	20	41	57	19.81	-34.81	-1.31	8.25	8.31
19	2	27	20	42	58	19.87	-34.81	-2.63	8.25	8.31
19	2	27	20	43	58	19.87	-34.81	-1.31	8.25	8.25
19	2	27	20	44	58	19.87	-34.81	-2.25	8.25	8.25
19	2	27	20	45	58	19.87	-34.75	-1.12	8.19	8.25
19	2	27	20	46	59	19.87	-34.75	-2.31	8.19	8.19
19	2	27	20	47	59	19.87	-34.94	-3.75	8.19	8.19
19	2	27	20	48	59	19.81	-34.75	-2.06	8.19	8.19
19	2	27	20	49	59	19.81	-34.81	-3.38	8.13	8.13
19	2	27	20	51	0	19.81	-34.69	-1.94	8.13	8.13
19	2	27	20	52	0	19.75	-34.75	-3.06	8.13	8.13
19	2	27	20	53	0	19.75	-34.69	-1.81	8.13	8.06
19	2	27	20	54	0	19.69	-34.75	-2.81	8.06	8.06
19	2	27	20	55	1	19.69	-34.88	-4.19	8.06	8.06
19	2	27	20	56	1	19.69	-34.69	-2.5	8.06	8
19	2	27	20	57	1	19.69	-34.81	-4	8.06	8
19	2	27	20	58	1	19.69	-34.69	-2.5	8.06	7.94
19	2	27	20	59	2	19.69	-34.69	-3.75	8	7.94
19	2	27	21	0	2	19.69	-34.69	-2.44	8	7.88
19	2	27	21	1	2	19.69	-34.69	-3.56	8	7.88
19	2	27	21	2	2	19.75	-34.75	-3.63	8	7.88
19	2	27	21	3	3	19.75	-34.63	-3.19	7.94	7.81
19	2	27	21	4	3	19.75	-34.69	-4.38	7.94	7.81
19	2	27	21	5	3	19.75	-34.56	-2.81	7.94	7.81
19	2	27	21	6	3	19.81	-34.56	-4	7.94	7.81
19	2	27	21	7	4	19.81	-34.56	-2.63	7.88	7.81
19	2	27	21	8	4	19.81	-34.56	-3.81	7.88	7.81
19	2	27	21	9	4	19.81	-34.56	-2.56	7.88	7.81
19	2	27	21	10	5	19.87	-34.56	-3.31	7.88	7.75
19	2	27	21	11	5	19.87	-34.75	-4.94	7.88	7.81
19	2	27	21	12	5	19.87	-34.56	-3.19	7.88	7.75
19	2	27	21	13	5	19.81	-34.63	-4.69	7.81	7.75
19	2	27	21	14	6	19.81	-34.56	-3.19	7.81	7.69
19	2	27	21	15	6	19.81	-34.63	-4.5	7.81	7.69
19	2	27	21	16	10	19.75	-34.56	-3.06	7.81	7.69
19	2	27	21	17	10	19.75	-34.5	-4.25	7.75	7.63
19	2	27	21	18	10	19.69	-34.69	-5.88	7.75	7.63
19	2	27	21	19	10	19.69	-34.5	-3.94	7.75	7.63
19	2	27	21	20	11	19.69	-34.56	-5.44	7.75	7.56
19	2	27	21	21	11	19.69	-34.44	-3.81	7.69	7.56
19	2	27	21	22	11	19.69	-34.5	-5.19	7.69	7.5
19	2	27	21	23	11	19.69	-34.44	-3.69	7.69	7.5
19	2	27	21	24	12	19.69	-34.44	-4.94	7.69	7.44
19	2	27	21	25	12	19.69	-34.44	-4.06	7.69	7.44
19	2	27	21	26	12	19.69	-34.44	-4.63	7.63	7.44
19	2	27	21	27	12	19.69	-34.5	-6.25	7.63	7.44
19	2	27	21	28	13	19.69	-34.38	-4.31	7.63	7.38
19	2	27	21	29	13	19.75	-34.44	-5.88	7.63	7.38
19	2	27	21	30	13	19.75	-34.38	-4.25	7.56	7.38
19	2	27	21	31	13	19.75	-34.44	-5.63	7.56	7.31
19	2	27	21	32	14	19.75	-34.38	-4.13	7.56	7.31
19	2	27	21	33	14	19.75	-34.38	-5.5	7.56	7.25
19	2	27	21	34	14	19.81	-34.5	-5.75	7.5	7.25
19	2	27	21	35	14	19.81	-34.38	-5.13	7.5	7.19
19	2	27	21	36	15	19.81	-34.44	-6.31	7.5	7.25
19	2	27	21	37	15	19.87	-34.19	-4.5	7.5	7.25
19	2	27	21	38	15	19.87	-34.25	-6.19	7.5	7.25
19	2	27	21	39	16	19.87	-34.19	-4.56	7.44	7.19
19	2	27	21	40	16	19.87	-34.25	-6.13	7.44	7.19
19	2	27	21	41	16	19.87	-34.19	-4.56	7.44	7.19
19	2	27	21	42	16	19.94	-34.31	-5.88	7.38	7.13
19	2	27	21	43	17	19.94	-34.44	-6.5	7.38	7.13
19	2	27	21	44	17	20	-34.31	-5.56	7.38	7.06
19	2	27	21	45	17	20	-34.44	-7.13	7.38	7.06
19	2	27	21	46	17	20.06	-34.31	-5.19	7.31	7.06
19	2	27	21	47	18	20.06	-34.38	-6.88	7.31	7
19	2	27	21	48	18	20.06	-34.25	-5.19	7.31	7
19	2	27	21	49	18	20.06	-34.31	-6.75	7.31	7
19	2	27	21	50	18	20.06	-34.31	-5.13	7.25	6.94
19	2	27	21	51	19	20	-34.38	-6.5	7.25	6.88
19	2	27	21	52	19	20	-34.44	-6.31	7.25	6.94
19	2	27	21	53	19	20	-34.19	-5.94	7.25	6.88
19	2	27	21	54	19	19.94	-34.31	-7.88	7.19	6.94
19	2	27	21	55	20	19.94	-34.19	-5.81	7.19	6.88
19	2	27	21	56	20	19.87	-34.25	-7.63	7.19	6.88
19	2	27	21	57	20	19.87	-34.19	-5.81	7.19	6.81
19	2	27	21	58	20	19.87	-34.25	-7.31	7.13	6.81
19	2	27	21	59	21	19.87	-34.19	-5.63	7.13	6.75
19	2	27	22	0	21	19.87	-34.25	-7.06	7.13	6.75
19	2	27	22	1	21	19.87	-34.19	-5.63	7.13	6.75
19	2	27	22	2	22	19.87	-34.25	-6.81	7.06	6.69
19	2	27	22	3	22	19.87	-34.44	-8.75	7.06	6.69
19	2	27	22	4	22	19.87	-34.25	-6.5	7.06	6.69
19	2	27	22	5	22	19.87	-34.31	-8.19	7.06	6.69
19	2	27	22	6	23	19.87	-34.19	-6.25	7	6.63
19	2	27	22	7	23	19.94	-34.31	-8	7	6.63
19	2	27	22	8	23	19.94	-34.19	-6.25	7	6.56
19	2	27	22	9	23	19.94	-34.25	-7.81	7	6.56
19	2	27	22	10	24	19.94	-34.19	-6.19	7	6.56
19	2	27	22	11	24	19.94	-34.19	-7.5	6.94	6.5
19	2	27	22	12	24	19.94	-34.31	-7.75	6.94	6.5
19	2	27	22	13	24	19.94	-34.19	-7.25	6.94	6.44
19	2	27	22	14	25	19.94	-34.38	-9.25	6.94	6.5
19	2	27	22	15	25	19.94	-34.19	-7	6.88	6.44
19	2	27	22	16	25	19.94	-34.31	-9.38	6.88	6.44
19	2	27	22	17	25	20	-34.13	-7.25	6.88	6.38
19	2	27	22	18	26	19.94	-34.25	-8.94	6.81	6.31
19	2	27	22	19	26	19.94	-34.13	-7	6.81	6.31
19	2	27	22	20	26	19.94	-34.19	-8.63	6.81	6.25
19	2	27	22	21	26	19.87	-34.13	-6.88	6.75	6.25
19	2	27	22	22	27	19.87	-34.13	-8.19	6.75	6.25
19	2	27	22	23	27	19.81	-34.25	-8	6.75	6.25
19	2	27	22	24	27	19.81	-34.13	-7.94	6.75	6.19
19	2	27	22	25	27	19.75	-34.31	-9.69	6.75	6.19
19	2	27	22	26	28	19.75	-34	-7.44	6.69	6.19
19	2	27	22	27	28	19.75	-34.19	-9.5	6.69	6.19
19	2	27	22	28	28	19.75	-34.06	-7.38	6.69	6.19
19	2	27	22	29	29	19.75	-34.19	-9.31	6.63	6.13
19	2	27	22	30	29	19.75	-34.06	-7.38	6.63	6.13
19	2	27	22	31	29	19.75	-34.13	-9	6.63	6.06
19	2	27	22	32	29	19.81	-34.06	-7.25	6.63	6.06
19	2	27	22	33	30	19.81	-34.13	-9	6.56	6
19	2	27	22	34	30	19.81	-34.13	-7.94	6.56	6
19	2	27	22	35	30	19.81	-34.13	-8.75	6.56	5.94
19	2	27	22	36	30	19.81	-34.31	-10.81	6.56	6
19	2	27	22	37	31	19.87	-34.13	-8.5	6.5	5.88
19	2	27	22	38	31	19.87	-34.25	-10.75	6.5	5.94
19	2	27	22	39	31	19.87	-34.06	-8.38	6.5	5.81
19	2	27	22	40	31	19.87	-34.19	-10.38	6.5	5.81
19	2	27	22	41	32	19.94	-34.06	-8.25	6.44	5.81
19	2	27	22	42	32	19.94	-34.19	-10.13	6.44	5.75
19	2	27	22	43	32	19.94	-34.06	-8.13	6.44	5.75
19	2	27	22	44	32	19.87	-34.13	-9.75	6.38	5.69
19	2	27	22	45	33	19.87	-34.06	-7.94	6.38	5.69
19	2	27	22	46	33	19.81	-34.06	-9.44	6.38	5.63
19	2	27	22	47	33	19.81	-34.19	-9.63	6.38	5.69
19	2	27	22	48	33	19.81	-34.06	-9.38	6.31	5.56
19	2	27	22	49	34	19.81	-34.31	-11.69	6.31	5.63
19	2	27	22	50	34	19.75	-34.06	-9.06	6.31	5.56
19	2	27	22	51	34	19.75	-34.25	-11.13	6.31	5.56
19	2	27	22	52	34	19.81	-34.06	-8.88	6.25	5.5
19	2	27	22	53	35	19.81	-34.19	-10.88	6.25	5.5
19	2	27	22	54	35	19.81	-34.06	-8.75	6.25	5.5
19	2	27	22	55	35	19.81	-34.06	-10.44	6.19	5.5
19	2	27	22	56	36	19.87	-34	-8.5	6.19	5.44
19	2	27	22	57	36	19.87	-34.06	-10.25	6.19	5.44
19	2	27	22	58	36	19.87	-34	-8.44	6.19	5.44
19	2	27	22	59	36	19.87	-34.06	-10.19	6.13	5.38
19	2	27	23	0	37	19.94	-34.13	-9.75	6.13	5.38
19	2	27	23	1	37	19.94	-34.06	-10	6.13	5.31
19	2	27	23	2	37	19.94	-34.25	-12.19	6.06	5.38
19	2	27	23	3	37	20	-34.06	-9.75	6.06	5.25
19	2	27	23	4	38	20	-34.19	-11.88	6.06	5.31
19	2	27	23	5	38	20	-34	-9.44	6	5.25
19	2	27	23	6	38	20	-34.19	-11.25	6	5.31
19	2	27	23	7	38	20	-34.06	-9.06	6	5.25
19	2	27	23	8	39	19.94	-34.13	-11.25	6	5.25
19	2	27	23	9	39	19.94	-34.06	-9.13	6	5.25
19	2	27	23	10	39	19.87	-34.13	-11.06	5.94	5.19
19	2	27	23	11	39	19.87	-34.06	-9.06	5.94	5.19
19	2	27	23	12	40	19.81	-34.06	-10.81	5.94	5.13
19	2	27	23	13	40	19.75	-34	-8.94	5.94	5.19
19	2	27	23	14	40	19.75	-34.06	-10.38	5.88	5.13
19	2	27	23	15	40	19.75	-34.25	-10.44	5.88	5.13
19	2	27	23	16	41	19.69	-34.06	-10.06	5.88	5.06
19	2	27	23	17	41	19.69	-34.25	-12.38	5.88	5.13
19	2	27	23	18	41	19.69	-34	-10	5.81	5.06
19	2	27	23	19	41	19.69	-34.19	-12.63	5.81	5.06
19	2	27	23	20	42	19.69	-34	-10.06	5.81	5
19	2	27	23	21	42	19.69	-34.13	-12.19	5.75	5
19	2	27	23	22	42	19.69	-34	-9.88	5.75	4.94
19	2	27	23	23	42	19.75	-34.06	-11.94	5.75	4.94
19	2	27	23	24	43	19.69	-34	-9.75	5.75	4.94
19	2	27	23	25	43	19.75	-34.06	-11.75	5.69	4.88
19	2	27	23	26	43	19.75	-34	-9.69	5.69	4.88
19	2	27	23	27	44	19.75	-34	-11.5	5.69	4.88
19	2	27	23	28	44	19.75	-33.94	-9.5	5.69	4.88
19	2	27	23	29	44	19.75	-34	-11.38	5.63	4.81
19	2	27	23	30	44	19.75	-33.94	-9.75	5.63	4.81
19	2	27	23	31	45	19.75	-33.94	-11.06	5.63	4.75
19	2	27	23	32	45	19.75	-34.13	-11.69	5.56	4.81
19	2	27	23	33	45	19.75	-33.88	-10.88	5.56	4.75
19	2	27	23	34	45	19.81	-34.13	-13.31	5.56	4.81
19	2	27	23	35	46	19.81	-33.94	-10.75	5.5	4.75
19	2	27	23	36	46	19.81	-34.13	-13.38	5.5	4.75
19	2	27	23	37	46	19.81	-33.94	-10.75	5.5	4.69
19	2	27	23	38	46	19.75	-34.06	-13	5.5	4.69
19	2	27	23	39	47	19.75	-33.94	-10.56	5.44	4.63
19	2	27	23	40	47	19.69	-34.06	-12.88	5.44	4.63
19	2	27	23	41	47	19.69	-33.94	-10.56	5.44	4.56
19	2	27	23	42	47	19.62	-34	-12.5	5.44	4.56
19	2	27	23	43	48	19.62	-33.94	-10.38	5.38	4.56
19	2	27	23	44	48	19.56	-34	-12.38	5.38	4.56
19	2	27	23	45	48	19.56	-33.94	-10.31	5.38	4.56
19	2	27	23	46	48	19.56	-34	-12.19	5.31	4.5
19	2	27	23	47	49	19.5	-33.94	-10.19	5.31	4.5
19	2	27	23	48	49	19.5	-33.94	-11.94	5.31	4.44
19	2	27	23	49	49	19.56	-34	-10.75	5.31	4.5
19	2	27	23	50	49	19.56	-33.94	-11.75	5.25	4.38
19	2	27	23	51	50	19.56	-34.13	-12.25	5.25	4.44
19	2	27	23	52	50	19.56	-33.81	-11.19	5.25	4.44
19	2	27	23	53	50	19.56	-34.06	-13.81	5.25	4.5
19	2	27	23	54	50	19.56	-33.81	-11.25	5.19	4.44
19	2	27	23	55	51	19.62	-34	-13.81	5.19	4.44
19	2	27	23	56	51	19.62	-33.88	-11.19	5.19	4.38
19	2	27	23	57	51	19.62	-34	-13.38	5.19	4.38
19	2	27	23	58	52	19.69	-33.88	-10.94	5.13	4.31
19	2	27	23	59	52	19.69	-34	-13.19	5.13	4.31
19	2	28	0	0	52	19.69	-33.88	-10.88	5.13	4.31
19	2	28	0	1	52	19.69	-34	-12.88	5.13	4.25
19	2	28	0	2	53	19.69	-33.94	-10.69	5.06	4.25
19	2	28	0	3	53	19.69	-33.94	-12.56	5.06	4.25
19	2	28	0	4	53	19.62	-33.94	-10.5	5.06	4.25
19	2	28	0	5	53	19.62	-33.94	-12.38	5.06	4.19
19	2	28	0	6	54	19.56	-33.88	-10.44	5	4.19
19	2	28	0	7	54	19.56	-33.88	-12.06	5	4.19
19	2	28	0	8	54	19.5	-33.94	-10.94	5	4.19
19	2	28	0	9	54	19.5	-33.94	-11.94	4.94	4.13
19	2	28	0	10	55	19.44	-34.06	-12.69	4.94	4.19
19	2	28	0	11	55	19.44	-33.88	-11.56	4.94	4.13
19	2	28	0	12	55	19.37	-34.13	-13.94	4.94	4.19
19	2	28	0	13	55	19.37	-33.94	-11.38	4.94	4.13
19	2	28	0	14	56	19.37	-34.13	-13.75	4.94	4.13
19	2	28	0	15	56	19.37	-33.94	-11.19	4.88	4.13
19	2	28	0	16	56	19.37	-34.13	-13.56	4.88	4.13
19	2	28	0	17	56	19.37	-34	-11.13	4.88	4.06
19	2	28	0	18	57	19.37	-34.06	-13.25	4.88	4.06
19	2	28	0	19	57	19.37	-33.94	-11	4.81	4
19	2	28	0	20	57	19.37	-34	-12.88	4.81	4
19	2	28	0	21	57	19.37	-33.94	-10.75	4.81	4
19	2	28	0	22	58	19.37	-34	-12.75	4.75	4
19	2	28	0	23	58	19.37	-34	-10.75	4.75	4
19	2	28	0	24	58	19.37	-33.94	-12.38	4.75	4
19	2	28	0	25	58	19.31	-33.94	-10.5	4.75	4
19	2	28	0	26	59	19.31	-34	-12.31	4.69	3.94
19	2	28	0	27	59	19.31	-34	-11.19	4.75	3.94
19	2	28	0	28	59	19.25	-34	-12	4.69	3.88
19	2	28	0	30	0	19.25	-34.19	-12.75	4.69	3.94
19	2	28	0	31	0	19.25	-34	-11.75	4.69	3.88
19	2	28	0	32	0	19.19	-34.19	-14.19	4.69	3.94
19	2	28	0	33	0	19.25	-34	-11.63	4.63	3.88
19	2	28	0	34	1	19.19	-34.13	-13.88	4.63	3.88
19	2	28	0	35	1	19.25	-33.94	-11.38	4.63	3.88
19	2	28	0	36	1	19.25	-34.13	-13.75	4.63	3.88
19	2	28	0	37	1	19.25	-34	-11.31	4.63	3.81
19	2	28	0	38	2	19.25	-34.06	-13.38	4.56	3.81
19	2	28	0	39	2	19.25	-34	-11.13	4.56	3.81
19	2	28	0	40	2	19.25	-34.06	-13.13	4.56	3.81
19	2	28	0	41	2	19.25	-34	-11.06	4.56	3.81
19	2	28	0	42	3	19.25	-34.06	-12.81	4.5	3.81
19	2	28	0	43	3	19.25	-34	-10.81	4.5	3.81
19	2	28	0	44	3	19.31	-34	-12.63	4.5	3.75
19	2	28	0	45	3	19.31	-34	-10.69	4.5	3.81
19	2	28	0	46	4	19.31	-34.06	-12.5	4.5	3.69
19	2	28	0	47	4	19.31	-34.06	-11.38	4.5	3.75
19	2	28	0	48	4	19.31	-33.94	-12.13	4.44	3.69
19	2	28	0	49	5	19.37	-34.19	-13	4.44	3.75
19	2	28	0	50	5	19.37	-34	-11.94	4.44	3.69
19	2	28	0	51	5	19.37	-34.19	-14.13	4.44	3.75
19	2	28	0	52	5	19.37	-34	-11.69	4.38	3.63
19	2	28	0	53	6	19.44	-34.19	-14.06	4.38	3.69
19	2	28	0	54	6	19.44	-34	-11.5	4.38	3.63
19	2	28	0	55	6	19.44	-34.13	-13.75	4.38	3.69
19	2	28	0	56	6	19.5	-34	-11.38	4.38	3.63
19	2	28	0	57	7	19.5	-34.13	-13.5	4.38	3.63
19	2	28	0	58	7	19.5	-34.06	-11.31	4.31	3.63
19	2	28	0	59	7	19.56	-34	-13.06	4.31	3.63
19	2	28	1	0	7	19.56	-34	-11	4.31	3.63
19	2	28	1	1	8	19.56	-34.06	-13	4.31	3.63
19	2	28	1	2	8	19.56	-34	-10.94	4.25	3.56
19	2	28	1	3	8	19.56	-34	-12.56	4.25	3.56
19	2	28	1	4	8	19.56	-34	-10.69	4.25	3.63
19	2	28	1	5	9	19.56	-34	-12.5	4.25	3.56
19	2	28	1	6	11	19.5	-34.13	-12.06	4.19	3.56
19	2	28	1	7	11	19.5	-34	-12.06	4.19	3.5
19	2	28	1	8	11	19.44	-34.19	-13.5	4.19	3.56
19	2	28	1	9	12	19.44	-34	-11.81	4.19	3.5
19	2	28	1	10	12	19.37	-34.25	-14.13	4.19	3.56
19	2	28	1	11	12	19.37	-34.06	-11.56	4.19	3.5
19	2	28	1	12	12	19.31	-34.13	-13.56	4.19	3.5
19	2	28	1	13	13	19.31	-34	-11.25	4.13	3.5
19	2	28	1	14	13	19.25	-34.13	-13.44	4.13	3.5
19	2	28	1	15	13	19.25	-34	-11.19	4.13	3.44
19	2	28	1	16	14	19.25	-34.06	-13	4.13	3.5
19	2	28	1	17	14	19.25	-34	-11	4.13	3.44
19	2	28	1	18	14	19.25	-34.13	-12.81	4.06	3.44
19	2	28	1	19	14	19.25	-34.06	-10.81	4.06	3.44
19	2	28	1	20	15	19.25	-34.06	-12.31	4.06	3.44
19	2	28	1	21	15	19.25	-34.06	-10.5	4.06	3.44
19	2	28	1	22	15	19.25	-34.06	-12	4.06	3.44
19	2	28	1	23	15	19.25	-34.19	-11.75	4.06	3.5
19	2	28	1	24	16	19.25	-34.13	-11.75	4.06	3.44
19	2	28	1	25	16	19.31	-34.25	-13.31	4.06	3.44
19	2	28	1	26	16	19.31	-34	-11.63	4	3.44
19	2	28	1	27	16	19.31	-34.25	-13.75	4.06	3.44
19	2	28	1	28	17	19.31	-34.06	-11.31	4	3.38
19	2	28	1	29	17	19.31	-34.13	-13.13	4	3.44
19	2	28	1	30	17	19.31	-34	-10.94	4	3.44
19	2	28	1	31	17	19.31	-34.13	-12.88	4	3.44
19	2	28	1	32	18	19.37	-34.06	-10.81	3.94	3.44
19	2	28	1	33	18	19.37	-34.06	-12.38	3.94	3.44
19	2	28	1	34	18	19.37	-34	-10.5	3.94	3.44
19	2	28	1	35	18	19.37	-34.06	-12.19	3.94	3.44
19	2	28	1	36	19	19.31	-34.06	-10.38	3.94	3.44
19	2	28	1	37	19	19.31	-34.06	-11.69	3.94	3.44
19	2	28	1	38	19	19.31	-34.13	-10.94	3.94	3.5
19	2	28	1	39	19	19.25	-34.06	-11.31	3.88	3.44
19	2	28	1	40	20	19.25	-34.31	-13.06	3.94	3.5
19	2	28	1	41	20	19.19	-34.13	-11.31	3.88	3.44
19	2	28	1	42	20	19.19	-34.25	-13.13	3.88	3.44
19	2	28	1	43	20	19.19	-34.06	-10.81	3.88	3.44
19	2	28	1	44	21	19.19	-34.25	-12.81	3.88	3.44
19	2	28	1	45	21	19.19	-34.13	-10.69	3.88	3.44
19	2	28	1	46	21	19.19	-34.06	-12.25	3.81	3.44
19	2	28	1	47	21	19.25	-34	-10.31	3.81	3.44
19	2	28	1	48	22	19.25	-34.13	-11.88	3.81	3.44
19	2	28	1	49	22	19.25	-34.13	-10.13	3.81	3.44
19	2	28	1	50	22	19.31	-34.06	-11.44	3.81	3.44
19	2	28	1	51	23	19.31	-34.06	-9.88	3.81	3.44
19	2	28	1	52	23	19.31	-34.13	-11.25	3.81	3.38
19	2	28	1	53	23	19.37	-34.31	-12.19	3.81	3.44
19	2	28	1	54	23	19.37	-34	-10.75	3.81	3.38
19	2	28	1	55	24	19.37	-34.25	-12.81	3.75	3.44
19	2	28	1	56	24	19.37	-34.13	-10.56	3.75	3.44
19	2	28	1	57	24	19.37	-34.19	-12.19	3.75	3.44
19	2	28	1	58	24	19.44	-34.13	-10.19	3.75	3.44
19	2	28	1	59	25	19.44	-34.13	-11.81	3.75	3.44
19	2	28	2	0	25	19.44	-34.13	-9.94	3.75	3.44
19	2	28	2	1	25	19.44	-34.19	-11.5	3.75	3.38
19	2	28	2	2	25	19.44	-34.19	-9.81	3.75	3.38
19	2	28	2	3	26	19.44	-34.13	-11	3.69	3.38
19	2	28	2	4	26	19.44	-34.13	-9.88	3.69	3.44
19	2	28	2	5	26	19.44	-34.19	-10.81	3.69	3.38
19	2	28	2	6	26	19.44	-34.38	-12.31	3.69	3.44
19	2	28	2	7	27	19.44	-34.13	-10.25	3.69	3.38
19	2	28	2	8	27	19.44	-34.31	-12.19	3.69	3.44
19	2	28	2	9	27	19.37	-34.19	-10.06	3.69	3.38
19	2	28	2	10	27	19.37	-34.19	-11.63	3.69	3.38
19	2	28	2	11	28	19.31	-34.13	-9.75	3.69	3.38
19	2	28	2	12	28	19.25	-34.13	-11.13	3.69	3.44
19	2	28	2	13	28	19.25	-34.06	-9.44	3.63	3.44
19	2	28	2	14	28	19.25	-34.13	-10.81	3.63	3.44
19	2	28	2	15	29	19.19	-34.13	-9.25	3.69	3.44
19	2	28	2	16	29	19.19	-34.19	-10.38	3.63	3.44
19	2	28	2	17	29	19.19	-34.44	-11.31	3.63	3.44
19	2	28	2	18	29	19.19	-34.19	-9.94	3.63	3.44
19	2	28	2	19	30	19.19	-34.38	-11.69	3.63	3.44
19	2	28	2	20	30	19.19	-34.19	-9.63	3.63	3.44
19	2	28	2	21	30	19.19	-34.25	-11.19	3.63	3.5
19	2	28	2	22	30	19.19	-34.19	-9.38	3.63	3.44
19	2	28	2	23	31	19.25	-34.25	-10.81	3.63	3.44
19	2	28	2	24	31	19.25	-34.25	-9.19	3.56	3.44
19	2	28	2	25	31	19.25	-34.25	-10.5	3.63	3.44
19	2	28	2	26	32	19.25	-34.25	-9	3.63	3.44
19	2	28	2	27	32	19.25	-34.25	-10	3.56	3.44
19	2	28	2	28	32	19.25	-34.44	-11.19	3.56	3.5
19	2	28	2	29	32	19.25	-34.25	-9.88	3.56	3.44
19	2	28	2	30	33	19.25	-34.44	-11.5	3.56	3.44
19	2	28	2	31	33	19.25	-34.25	-9.44	3.56	3.44
19	2	28	2	32	33	19.25	-34.38	-10.81	3.56	3.44
19	2	28	2	33	33	19.31	-34.25	-9.06	3.56	3.44
19	2	28	2	34	34	19.25	-34.31	-10.56	3.56	3.44
19	2	28	2	35	34	19.25	-34.31	-8.94	3.56	3.44
19	2	28	2	36	34	19.25	-34.25	-10.06	3.56	3.44
19	2	28	2	37	34	19.19	-34.25	-8.63	3.5	3.5
19	2	28	2	38	35	19.19	-34.31	-9.75	3.5	3.44
19	2	28	2	39	35	19.19	-34.56	-11.25	3.5	3.5
19	2	28	2	40	35	19.12	-34.31	-9.19	3.5	3.5
19	2	28	2	41	35	19.12	-34.44	-10.81	3.5	3.5
19	2	28	2	42	36	19.12	-34.31	-9	3.5	3.5
19	2	28	2	43	36	19.12	-34.38	-10.25	3.5	3.5
19	2	28	2	44	36	19.12	-34.31	-8.63	3.5	3.5
19	2	28	2	45	36	19.12	-34.38	-10	3.5	3.5
19	2	28	2	46	37	19.12	-34.31	-8.5	3.5	3.5
19	2	28	2	47	37	19.12	-34.31	-9.44	3.5	3.5
19	2	28	2	48	37	19.12	-34.5	-10.13	3.5	3.56
19	2	28	2	49	37	19.19	-34.38	-9.13	3.5	3.5
19	2	28	2	50	38	19.19	-34.56	-10.5	3.5	3.5
19	2	28	2	51	38	19.19	-34.31	-8.63	3.5	3.5
19	2	28	2	52	38	19.19	-34.44	-10.13	3.44	3.5
19	2	28	2	53	38	19.19	-34.38	-8.5	3.44	3.5
19	2	28	2	54	39	19.19	-34.38	-9.63	3.5	3.5
19	2	28	2	55	39	19.19	-34.38	-8.19	3.44	3.5
19	2	28	2	56	39	19.19	-34.44	-9.31	3.44	3.5
19	2	28	2	57	40	19.25	-34.56	-9.13	3.44	3.5
19	2	28	2	58	40	19.25	-34.38	-8.81	3.44	3.5
19	2	28	2	59	40	19.25	-34.56	-10.31	3.44	3.56
19	2	28	3	0	40	19.25	-34.44	-8.44	3.44	3.5
19	2	28	3	1	41	19.19	-34.44	-9.63	3.44	3.56
19	2	28	3	2	41	19.19	-34.38	-8.06	3.44	3.56
19	2	28	3	3	41	19.19	-34.44	-9.25	3.44	3.56
19	2	28	3	4	41	19.12	-34.38	-7.88	3.44	3.56
19	2	28	3	5	42	19.12	-34.44	-9.06	3.44	3.56
19	2	28	3	6	42	19.06	-34.56	-8.88	3.44	3.56
19	2	28	3	7	42	19.06	-34.44	-8.44	3.44	3.56
19	2	28	3	8	42	19.06	-34.63	-9.75	3.44	3.56
19	2	28	3	9	43	19	-34.44	-8.06	3.44	3.56
19	2	28	3	10	43	19	-34.56	-9.31	3.44	3.56
19	2	28	3	11	43	19	-34.5	-7.81	3.44	3.56
19	2	28	3	12	43	19	-34.5	-8.88	3.44	3.56
19	2	28	3	13	44	19	-34.44	-7.56	3.44	3.56
19	2	28	3	14	44	19	-34.44	-8.56	3.44	3.56
19	2	28	3	15	44	19	-34.63	-9.31	3.44	3.56
19	2	28	3	16	44	19.06	-34.44	-8.31	3.38	3.56
19	2	28	3	17	45	19.06	-34.63	-9.56	3.38	3.56
19	2	28	3	18	45	19.06	-34.5	-7.88	3.38	3.56
19	2	28	3	19	45	19.06	-34.5	-8.88	3.38	3.56
19	2	28	3	20	45	19.06	-34.5	-7.44	3.38	3.56
19	2	28	3	21	46	19.06	-34.5	-8.5	3.38	3.63
19	2	28	3	22	46	19.06	-34.5	-7.25	3.38	3.63
19	2	28	3	23	46	19.12	-34.5	-8.13	3.38	3.56
19	2	28	3	24	46	19.12	-34.69	-9.56	3.38	3.63
19	2	28	3	25	47	19.12	-34.56	-7.75	3.38	3.63
19	2	28	3	26	47	19.12	-34.63	-8.75	3.38	3.63
19	2	28	3	27	47	19.19	-34.56	-7.31	3.38	3.63
19	2	28	3	28	48	19.19	-34.63	-8.38	3.38	3.63
19	2	28	3	29	48	19.19	-34.63	-7.06	3.38	3.63
19	2	28	3	30	48	19.19	-34.63	-7.94	3.38	3.63
19	2	28	3	31	48	19.25	-34.69	-7.81	3.38	3.63
19	2	28	3	32	49	19.25	-34.56	-7.5	3.38	3.63
19	2	28	3	33	49	19.31	-34.75	-8.69	3.38	3.69
19	2	28	3	34	49	19.31	-34.63	-7.13	3.38	3.63
19	2	28	3	35	49	19.31	-34.69	-8.25	3.38	3.69
19	2	28	3	36	50	19.25	-34.63	-6.88	3.38	3.69
19	2	28	3	37	50	19.25	-34.63	-7.75	3.38	3.69
19	2	28	3	38	50	19.25	-34.63	-6.56	3.38	3.69
19	2	28	3	39	50	19.19	-34.69	-7.31	3.38	3.69
19	2	28	3	40	51	19.19	-34.88	-8.44	3.38	3.75
19	2	28	3	41	51	19.12	-34.69	-6.81	3.38	3.69
19	2	28	3	42	51	19.12	-34.75	-7.94	3.38	3.75
19	2	28	3	43	51	19.12	-34.69	-6.63	3.38	3.69
19	2	28	3	44	52	19.12	-34.69	-7.5	3.38	3.69
19	2	28	3	45	52	19.12	-34.69	-6.31	3.38	3.75
19	2	28	3	46	52	19.12	-34.69	-7.19	3.38	3.69
19	2	28	3	47	52	19.12	-34.88	-8.31	3.38	3.75
19	2	28	3	48	53	19.12	-34.69	-6.69	3.38	3.75
19	2	28	3	49	53	19.12	-34.75	-7.69	3.38	3.75
19	2	28	3	50	53	19.12	-34.69	-6.38	3.38	3.75
19	2	28	3	51	53	19.12	-34.75	-7.25	3.31	3.75
19	2	28	3	52	54	19.12	-34.69	-6.13	3.31	3.75
19	2	28	3	53	54	19.19	-34.75	-6.88	3.38	3.75
19	2	28	3	54	54	19.19	-34.94	-7.75	3.38	3.75
19	2	28	3	55	54	19.19	-34.69	-6.44	3.31	3.75
19	2	28	3	56	55	19.19	-34.81	-7.44	3.38	3.75
19	2	28	3	57	55	19.19	-34.75	-6.13	3.31	3.75
19	2	28	3	58	55	19.19	-34.81	-6.94	3.31	3.81
19	2	28	3	59	56	19.19	-34.75	-5.81	3.38	3.81
19	2	28	4	0	56	19.19	-34.75	-6.56	3.31	3.81
19	2	28	4	1	56	19.25	-34.94	-7.25	3.31	3.81
19	2	28	4	2	56	19.25	-34.75	-6.06	3.38	3.81
19	2	28	4	3	57	19.25	-34.88	-7	3.31	3.81
19	2	28	4	4	57	19.25	-34.75	-5.75	3.38	3.81
19	2	28	4	5	57	19.25	-34.81	-6.56	3.38	3.81
19	2	28	4	6	57	19.25	-34.75	-5.5	3.38	3.88
19	2	28	4	7	58	19.25	-34.81	-6.13	3.38	3.88
19	2	28	4	8	58	19.25	-34.94	-7	3.38	3.88
19	2	28	4	9	58	19.25	-34.81	-5.69	3.31	3.88
19	2	28	4	10	58	19.25	-34.94	-6.56	3.38	3.88
19	2	28	4	11	59	19.25	-34.81	-5.38	3.38	3.88
19	2	28	4	12	59	19.31	-34.88	-6.13	3.38	3.88
19	2	28	4	13	59	19.25	-34.81	-5.13	3.38	3.88
19	2	28	4	14	59	19.25	-34.88	-5.81	3.38	3.88
19	2	28	4	16	0	19.19	-35.06	-6.88	3.38	3.94
19	2	28	4	17	0	19.19	-34.88	-5.44	3.38	3.94
19	2	28	4	18	0	19.19	-34.94	-6.13	3.38	3.94
19	2	28	4	19	0	19.12	-34.88	-5	3.38	3.94
19	2	28	4	20	1	19.12	-34.88	-5.69	3.38	3.94
19	2	28	4	21	1	19.12	-34.88	-4.75	3.38	3.94
19	2	28	4	22	1	19.06	-34.94	-5.38	3.38	3.94
19	2	28	4	23	1	19.06	-35	-6.06	3.38	4
19	2	28	4	24	2	19.06	-34.94	-4.88	3.38	4
19	2	28	4	25	2	19.12	-35	-5.69	3.38	4
19	2	28	4	26	2	19.12	-34.94	-4.69	3.38	4
19	2	28	4	27	2	19.12	-35	-5.25	3.38	4
19	2	28	4	28	3	19.12	-35	-4.38	3.38	4
19	2	28	4	29	3	19.12	-34.94	-4.69	3.38	4
19	2	28	4	30	3	19.19	-35	-5.31	3.38	4.06
19	2	28	4	31	3	19.19	-34.94	-4.25	3.38	4.06
19	2	28	4	32	4	19.19	-35	-5	3.38	4.06
19	2	28	4	33	4	19.25	-34.94	-4.13	3.38	4.06
19	2	28	4	34	4	19.25	-35	-4.63	3.38	4.06
19	2	28	4	35	5	19.25	-35.13	-5.44	3.38	4.13
19	2	28	4	36	5	19.25	-35	-4.25	3.38	4.13
19	2	28	4	37	5	19.25	-35.06	-4.88	3.38	4.13
19	2	28	4	38	5	19.31	-35.06	-3.94	3.38	4.13
19	2	28	4	39	6	19.31	-35.06	-4.38	3.38	4.13
19	2	28	4	40	6	19.31	-35	-3.63	3.38	4.13
19	2	28	4	41	6	19.31	-35.06	-4.06	3.38	4.13
19	2	28	4	42	6	19.31	-35.13	-4.69	3.38	4.19
19	2	28	4	43	7	19.31	-35.13	-3.69	3.38	4.19
19	2	28	4	44	7	19.31	-35.13	-4.31	3.38	4.19
19	2	28	4	45	7	19.37	-35.13	-3.5	3.44	4.19
19	2	28	4	46	7	19.37	-35.13	-3.94	3.38	4.19
19	2	28	4	47	8	19.37	-35.19	-4.44	3.44	4.19
19	2	28	4	48	8	19.37	-35.13	-3.44	3.38	4.19
19	2	28	4	49	8	19.37	-35.13	-3.94	3.44	4.25
19	2	28	4	50	8	19.37	-35.13	-3.13	3.44	4.25
19	2	28	4	51	9	19.37	-35.13	-3.63	3.44	4.25
19	2	28	4	52	9	19.37	-35.13	-2.94	3.44	4.25
19	2	28	4	53	9	19.31	-35.19	-3.38	3.44	4.25
19	2	28	4	54	9	19.25	-35.19	-3.81	3.44	4.31
19	2	28	4	55	10	19.25	-35.13	-2.94	3.44	4.31
19	2	28	4	56	13	19.19	-35.13	-3.38	3.44	4.31
19	2	28	4	57	14	19.19	-35.13	-2.69	3.44	4.31
19	2	28	4	58	14	19.12	-35.13	-3	3.44	4.31
19	2	28	4	59	14	19.12	-35.13	-3.31	3.44	4.38
19	2	28	5	0	14	19.12	-35.06	-2.56	3.44	4.38
19	2	28	5	1	15	19.12	-35.06	-2.94	3.44	4.38
19	2	28	5	2	15	19.12	-35.06	-2.31	3.44	4.38
19	2	28	5	3	15	19.12	-35.06	-2.63	3.44	4.38
19	2	28	5	4	15	19.12	-35.13	-3.06	3.44	4.44
19	2	28	5	5	16	19.12	-35.06	-2.25	3.44	4.44
19	2	28	5	6	16	19.12	-35.13	-2.75	3.44	4.44
19	2	28	5	7	16	19.12	-35.13	-2.06	3.44	4.44
19	2	28	5	8	17	19.12	-35.13	-2.31	3.44	4.44
19	2	28	5	9	17	19.12	-35.13	-1.81	3.44	4.5
19	2	28	5	10	17	19.12	-35.06	-1.5	3.5	4.5
19	2	28	5	11	17	19.19	-35.06	-1.75	3.5	4.56
19	2	28	5	12	18	19.19	-35.13	-1.25	3.5	4.56
19	2	28	5	13	18	19.19	-35.06	-1.56	3.5	4.56
19	2	28	5	14	18	19.19	-35.06	-1.12	3.5	4.56
19	2	28	5	15	18	19.19	-35.13	-1.44	3.5	4.63
19	2	28	5	16	19	19.19	-35.13	-1.75	3.5	4.63
19	2	28	5	17	19	19.25	-35.13	-1.12	3.5	4.63
19	2	28	5	18	19	19.25	-35.06	-1.37	3.5	4.63
19	2	28	5	19	19	19.25	-35.06	-0.88	3.5	4.63
19	2	28	5	20	20	19.25	-35.06	-1.06	3.5	4.63
19	2	28	5	21	20	19.25	-35.06	-1.19	3.56	4.69
19	2	28	5	22	20	19.25	-35	-0.63	3.5	4.69
19	2	28	5	23	20	19.31	-35	-0.88	3.56	4.69
19	2	28	5	24	21	19.31	-35	-0.44	3.56	4.75
19	2	28	5	25	21	19.31	-35	-0.69	3.56	4.75
19	2	28	5	26	21	19.31	-34.94	-0.5	3.56	4.75
19	2	28	5	27	21	19.31	-34.94	-0.06	3.56	4.75
19	2	28	5	28	22	19.31	-34.88	-0.19	3.56	4.81
19	2	28	5	29	22	19.25	-34.94	0.12	3.56	4.81
19	2	28	5	30	22	19.25	-34.88	0	3.56	4.81
19	2	28	5	31	22	19.19	-34.88	0.12	3.56	4.88
19	2	28	5	32	23	19.19	-34.81	0.44	3.56	4.88
19	2	28	5	33	23	19.12	-34.88	0.25	3.56	4.88
19	2	28	5	34	23	19.12	-34.88	0.56	3.63	4.94
19	2	28	5	35	23	19.06	-34.88	0.44	3.63	4.94
19	2	28	5	36	24	19.06	-34.88	0.75	3.63	4.94
19	2	28	5	37	24	19.06	-34.88	0.94	3.63	4.94
19	2	28	5	38	24	19.06	-34.81	0.94	3.63	5
19	2	28	5	39	25	19.06	-34.88	1.19	3.63	5
19	2	28	5	40	25	19.06	-34.88	1.25	3.63	5
19	2	28	5	41	25	19.06	-34.94	1.44	3.69	5.06
19	2	28	5	42	25	19.06	-34.94	1.56	3.69	5.06
19	2	28	5	43	26	19.06	-34.88	1.5	3.69	5.13
19	2	28	5	44	26	19.06	-34.94	1.69	3.69	5.13
19	2	28	5	45	26	19.12	-34.88	1.69	3.69	5.13
19	2	28	5	46	26	19.12	-34.94	1.81	3.69	5.19
19	2	28	5	47	27	19.12	-34.94	1.94	3.69	5.19
19	2	28	5	48	27	19.12	-34.94	2.06	3.69	5.19
19	2	28	5	49	27	19.12	-34.88	2.19	3.75	5.19
19	2	28	5	50	27	19.12	-34.88	2.25	3.69	5.25
19	2	28	5	51	28	19.12	-34.88	2.31	3.75	5.25
19	2	28	5	52	28	19.12	-34.81	2.38	3.75	5.25
19	2	28	5	53	28	19.12	-34.81	2.44	3.75	5.25
19	2	28	5	54	28	19.12	-34.75	2.5	3.75	5.25
19	2	28	5	55	29	19.06	-34.75	2.63	3.75	5.31
19	2	28	5	56	29	19.06	-34.69	2.63	3.75	5.31
19	2	28	5	57	29	19	-34.69	2.69	3.75	5.31
19	2	28	5	58	29	19	-34.69	2.75	3.81	5.31
19	2	28	5	59	30	18.94	-34.69	2.81	3.81	5.38
19	2	28	6	0	30	18.94	-34.63	2.88	3.81	5.38
19	2	28	6	1	30	18.94	-34.63	2.94	3.81	5.38
19	2	28	6	2	30	18.94	-34.63	2.94	3.81	5.38
19	2	28	6	3	31	18.94	-34.63	3	3.81	5.38
19	2	28	6	4	31	19	-34.63	3.06	3.81	5.38
19	2	28	6	5	31	19	-34.56	3.13	3.81	5.44
19	2	28	6	6	31	19	-34.56	3.19	3.81	5.44
19	2	28	6	7	32	19.06	-34.56	3.19	3.81	5.44
19	2	28	6	8	32	19.06	-34.56	3.25	3.81	5.44
19	2	28	6	9	32	19.06	-34.56	3.31	3.81	5.44
19	2	28	6	10	33	19.06	-34.56	3.31	3.81	5.44
19	2	28	6	11	33	19.12	-34.56	3.38	3.81	5.5
19	2	28	6	12	33	19.12	-34.56	3.44	3.88	5.5
19	2	28	6	13	33	19.12	-34.5	3.44	3.88	5.5
19	2	28	6	14	34	19.12	-34.5	3.5	3.88	5.5
19	2	28	6	15	34	19.19	-34.5	3.5	3.88	5.5
19	2	28	6	16	34	19.19	-34.5	3.56	3.88	5.5
19	2	28	6	17	34	19.19	-34.5	3.56	3.88	5.5
19	2	28	6	18	35	19.19	-34.5	3.63	3.88	5.5
19	2	28	6	19	35	19.19	-34.5	3.63	3.88	5.5
19	2	28	6	20	35	19.25	-34.5	3.69	3.88	5.5
19	2	28	6	21	35	19.25	-34.5	3.69	3.88	5.56
19	2	28	6	22	36	19.25	-34.5	3.75	3.88	5.5
19	2	28	6	23	36	19.25	-34.5	3.75	3.88	5.56
19	2	28	6	24	36	19.25	-34.44	3.81	3.88	5.56
19	2	28	6	25	36	19.31	-34.44	3.81	3.88	5.56
19	2	28	6	26	37	19.25	-34.44	3.88	3.88	5.56
19	2	28	6	27	37	19.31	-34.44	3.88	3.88	5.56
19	2	28	6	28	37	19.31	-34.44	3.94	3.88	5.56
19	2	28	6	29	37	19.31	-34.44	3.94	3.88	5.56
19	2	28	6	30	38	19.31	-34.44	4	3.88	5.56
19	2	28	6	31	38	19.37	-34.44	4	3.88	5.56
19	2	28	6	32	38	19.37	-34.44	4.06	3.88	5.56
19	2	28	6	33	38	19.37	-34.38	4.06	3.88	5.56
19	2	28	6	34	39	19.37	-34.38	4.06	3.88	5.56
19	2	28	6	35	39	19.37	-34.38	4.13	3.88	5.56
19	2	28	6	36	39	19.37	-34.38	4.13	3.88	5.63
19	2	28	6	37	39	19.31	-34.38	4.13	3.88	5.63
19	2	28	6	38	40	19.31	-34.38	4.19	3.88	5.63
19	2	28	6	39	40	19.31	-34.38	4.19	3.88	5.63
19	2	28	6	40	40	19.25	-34.38	4.19	3.88	5.63
19	2	28	6	41	41	19.25	-34.38	4.25	3.88	5.63
19	2	28	6	42	41	19.19	-34.38	4.25	3.88	5.63
19	2	28	6	43	41	19.19	-34.38	4.25	3.88	5.63
19	2	28	6	44	41	19.19	-34.38	4.31	3.88	5.63
19	2	28	6	45	42	19.12	-34.38	4.31	3.88	5.63
19	2	28	6	46	42	19.12	-34.38	4.31	3.94	5.63
19	2	28	6	47	42	19.12	-34.38	4.38	3.88	5.63
19	2	28	6	48	42	19.12	-34.31	4.38	3.94	5.63
19	2	28	6	49	43	19.19	-34.38	4.38	3.94	5.63
19	2	28	6	50	43	19.19	-34.38	4.44	3.94	5.63
19	2	28	6	51	43	19.25	-34.38	4.44	3.94	5.63
19	2	28	6	52	43	19.25	-34.38	4.44	3.94	5.63
19	2	28	6	53	44	19.25	-34.38	4.5	3.94	5.63
19	2	28	6	54	44	19.31	-34.38	4.5	3.94	5.63
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	2	28	6	56	54	19.37	-34.31	4.56	3.94	5.63
19	2	28	6	57	54	19.37	-34.31	4.56	3.94	5.63
19	2	28	6	58	54	19.37	-34.25	4.44	3.94	5.56
19	2	28	6	59	55	19.12	-34.31	4.5	3.94	5.63
19	2	28	7	0	55	19.44	-34.25	4.44	3.94	5.56
19	2	28	7	1	55	19.44	-34.31	4.5	4	5.56
19	2	28	7	2	55	19.44	-34.31	4.5	3.94	5.56
19	2	28	7	3	56	19.44	-34.31	4.44	4	5.56
19	2	28	7	4	56	19.44	-34.31	4.5	4	5.56
19	2	28	7	5	56	19.5	-34.31	4.44	3.94	5.56
19	2	28	7	6	57	19.5	-34.38	4.5	4	5.56
19	2	28	7	7	57	19.5	-34.31	4.5	4	5.56
19	2	28	7	8	57	19.5	-34.31	4.5	4	5.56
19	2	28	7	9	57	19.5	-34.31	4.56	4	5.56
19	2	28	7	10	58	19.5	-34.31	4.5	4	5.56
19	2	28	7	11	58	19.5	-34.38	4.56	4	5.56
19	2	28	7	12	58	19.5	-34.38	4.56	4	5.56
19	2	28	7	13	58	19.5	-34.31	4.5	4	5.56
19	2	28	7	14	59	19.5	-34.31	4.63	4	5.56
19	2	28	7	15	59	19.5	-34.31	4.56	4	5.56
19	2	28	7	16	59	19.56	-34.31	4.63	4	5.56
19	2	28	7	17	59	19.5	-34.31	4.63	4	5.56
19	2	28	7	19	0	19.5	-34.38	4.69	4.06	5.56
19	2	28	7	20	0	19.5	-34.38	4.75	4.06	5.56
19	2	28	7	21	0	19.44	-34.31	4.69	4.06	5.56
19	2	28	7	22	0	19.44	-34.38	4.75	4.06	5.56
19	2	28	7	23	1	19.37	-34.38	4.69	4.06	5.56
19	2	28	7	24	1	19.37	-34.38	4.75	4.06	5.56
19	2	28	7	25	1	19.37	-34.38	4.75	4.06	5.56
19	2	28	7	26	1	19.37	-34.44	4.81	4.06	5.56
19	2	28	7	27	2	19.37	-34.44	4.81	4.06	5.56
19	2	28	7	28	2	19.37	-34.38	4.88	4.06	5.56
19	2	28	7	29	2	19.37	-34.38	4.88	4.06	5.56
19	2	28	7	30	2	19.37	-34.31	4.88	4.06	5.56
19	2	28	7	31	3	19.37	-34.31	4.88	4.06	5.63
19	2	28	7	32	3	19.44	-34.25	4.94	4.06	5.63
19	2	28	7	33	3	19.44	-34.25	4.94	4.06	5.63
19	2	28	7	34	3	19.44	-34.19	4.94	4.06	5.63
19	2	28	7	35	4	19.44	-34.19	4.94	4.06	5.63
19	2	28	7	36	4	19.44	-34.19	4.94	4.06	5.63
19	2	28	7	37	4	19.44	-34.19	4.94	4.06	5.63
19	2	28	7	38	5	19.44	-34.19	5	4.06	5.63
19	2	28	7	39	5	19.37	-34.19	5	4.06	5.63
19	2	28	7	40	5	19.37	-34.19	5	4.06	5.63
19	2	28	7	41	5	19.37	-34.13	5	4.06	5.63
19	2	28	7	42	6	19.31	-34.19	5	4.06	5.63
19	2	28	7	43	6	19.31	-34.13	5	4.06	5.63
19	2	28	7	44	6	19.31	-34.13	5.06	4.06	5.63
19	2	28	7	45	6	19.31	-34.13	5.06	4.06	5.63
19	2	28	7	46	7	19.31	-34.13	5.06	4.06	5.63
19	2	28	7	47	7	19.31	-34.13	5.06	4.06	5.63
19	2	28	7	48	7	19.31	-34.13	5.06	4.06	5.63
19	2	28	7	49	7	19.31	-34.13	5.06	4.06	5.63
19	2	28	7	50	8	19.31	-34.13	5.13	4.06	5.69
19	2	28	7	51	8	19.37	-34.13	5.13	4.06	5.63
19	2	28	7	52	8	19.37	-34.13	5.13	4.06	5.63
19	2	28	7	53	8	19.37	-34.13	5.13	4.06	5.63
19	2	28	7	54	9	19.37	-34.13	5.13	4.06	5.63
19	2	28	7	55	9	19.44	-34.06	5.13	4.06	5.69
19	2	28	7	56	9	19.44	-34.13	5.13	4.06	5.69
19	2	28	7	57	9	19.5	-34.06	5.13	4.06	5.69
19	2	28	7	58	10	19.5	-34.06	5.13	4.06	5.69
19	2	28	7	59	10	19.5	-34.06	5.19	4.06	5.69
19	2	28	8	0	10	19.56	-34.06	5.19	4.06	5.69
19	2	28	8	1	10	19.56	-34.06	5.19	4.06	5.69
19	2	28	8	2	11	19.56	-34.06	5.19	4.06	5.69
19	2	28	8	3	11	19.56	-34.06	5.19	4.06	5.69
19	2	28	8	4	11	19.56	-34.06	5.19	4.06	5.69
19	2	28	8	5	12	19.56	-34.06	5.19	4.06	5.69
19	2	28	8	6	12	19.56	-34.06	5.19	4.06	5.69
19	2	28	8	7	12	19.56	-34.06	5.25	4.06	5.69
19	2	28	8	8	12	19.56	-34.06	5.25	4.06	5.69
19	2	28	8	9	13	19.56	-34.06	5.25	4.06	5.69
19	2	28	8	10	13	19.56	-34.06	5.25	4.06	5.69
19	2	28	8	11	13	19.56	-34.06	5.25	4.06	5.69
19	2	28	8	12	13	19.5	-34.06	5.25	4.13	5.69
19	2	28	8	13	14	19.5	-34.06	5.25	4.13	5.69
19	2	28	8	14	14	19.44	-34	5.25	4.06	5.69
19	2	28	8	15	14	19.44	-34	5.31	4.13	5.69
19	2	28	8	16	14	19.37	-34	5.31	4.13	5.69
19	2	28	8	17	15	19.37	-34	5.31	4.13	5.75
19	2	28	8	18	15	19.37	-34	5.31	4.13	5.75
19	2	28	8	19	15	19.37	-34	5.31	4.13	5.75
19	2	28	8	20	15	19.37	-34	5.31	4.13	5.75
19	2	28	8	21	16	19.37	-34	5.31	4.13	5.75
19	2	28	8	22	16	19.37	-34	5.31	4.13	5.75
19	2	28	8	23	16	19.37	-34	5.31	4.13	5.75
19	2	28	8	24	16	19.37	-34	5.31	4.13	5.75
19	2	28	8	25	17	19.37	-34	5.38	4.13	5.75
19	2	28	8	26	17	19.37	-34	5.38	4.13	5.75
19	2	28	8	27	17	19.37	-34	5.38	4.13	5.75
19	2	28	8	28	17	19.37	-34	5.38	4.13	5.75
19	2	28	8	29	18	19.44	-34	5.38	4.13	5.75
19	2	28	8	30	18	19.44	-34	5.38	4.13	5.75
19	2	28	8	31	18	19.44	-34	5.38	4.13	5.75
19	2	28	8	32	18	19.37	-34	5.38	4.13	5.75
19	2	28	8	33	19	19.37	-34	5.38	4.13	5.75
19	2	28	8	34	19	19.37	-34	5.38	4.13	5.75
19	2	28	8	35	19	19.31	-34	5.38	4.13	5.75
19	2	28	8	36	20	19.31	-34	5.38	4.13	5.75
19	2	28	8	37	20	19.25	-34	5.44	4.13	5.75
19	2	28	8	38	20	19.25	-33.94	5.44	4.13	5.75
19	2	28	8	39	20	19.25	-33.94	5.44	4.19	5.75
19	2	28	8	40	21	19.25	-33.94	5.44	4.13	5.81
19	2	28	8	41	21	19.25	-33.94	5.44	4.13	5.75
19	2	28	8	42	21	19.25	-33.94	5.44	4.13	5.75
19	2	28	8	43	21	19.25	-33.94	5.44	4.19	5.81
19	2	28	8	44	22	19.25	-33.94	5.44	4.13	5.81
19	2	28	8	45	22	19.25	-33.94	5.44	4.19	5.81
19	2	28	8	46	22	19.31	-33.94	5.44	4.19	5.81
19	2	28	8	47	22	19.31	-33.94	5.44	4.13	5.81
19	2	28	8	48	23	19.31	-33.94	5.5	4.19	5.81
19	2	28	8	49	23	19.37	-33.94	5.5	4.19	5.81
19	2	28	8	50	23	19.37	-33.94	5.5	4.19	5.81
19	2	28	8	51	23	19.37	-33.94	5.5	4.19	5.81
19	2	28	8	52	24	19.37	-33.94	5.5	4.19	5.81
19	2	28	8	53	24	19.37	-33.94	5.5	4.19	5.81
19	2	28	8	54	24	19.44	-33.94	5.5	4.19	5.81
19	2	28	8	55	24	19.37	-33.94	5.5	4.19	5.81
19	2	28	8	56	25	19.37	-33.94	5.5	4.19	5.81
19	2	28	8	57	25	19.37	-33.94	5.5	4.19	5.81
19	2	28	8	58	25	19.37	-33.94	5.5	4.19	5.81
19	2	28	8	59	25	19.31	-33.94	5.56	4.19	5.81
19	2	28	9	0	26	19.25	-33.94	5.5	4.19	5.81
19	2	28	9	1	26	19.25	-33.94	5.56	4.19	5.81
19	2	28	9	2	26	19.25	-33.88	5.56	4.19	5.81
19	2	28	9	3	26	19.19	-33.94	5.56	4.19	5.81
19	2	28	9	4	27	19.19	-33.94	5.56	4.19	5.81
19	2	28	9	5	27	19.19	-33.88	5.56	4.19	5.81
19	2	28	9	6	27	19.19	-33.88	5.56	4.19	5.88
19	2	28	9	7	28	19.19	-33.88	5.56	4.19	5.81
19	2	28	9	8	28	19.19	-33.88	5.56	4.19	5.81
19	2	28	9	9	28	19.19	-33.88	5.56	4.19	5.88
19	2	28	9	10	28	19.19	-33.88	5.56	4.19	5.88
19	2	28	9	11	29	19.19	-33.88	5.56	4.19	5.88
19	2	28	9	12	29	19.19	-33.88	5.56	4.19	5.88
19	2	28	9	13	29	19.19	-33.88	5.56	4.19	5.88
19	2	28	9	14	29	19.25	-33.88	5.56	4.19	5.88
19	2	28	9	15	30	19.25	-33.88	5.56	4.19	5.88
19	2	28	9	16	30	19.25	-33.88	5.56	4.25	5.88
19	2	28	9	17	30	19.25	-33.88	5.63	4.19	5.88
19	2	28	9	18	30	19.25	-33.88	5.63	4.25	5.88
19	2	28	9	19	31	19.25	-33.88	5.63	4.19	5.88
19	2	28	9	20	31	19.25	-33.88	5.63	4.25	5.88
19	2	28	9	21	31	19.25	-33.88	5.63	4.19	5.88
19	2	28	9	22	31	19.25	-33.88	5.63	4.19	5.88
19	2	28	9	23	32	19.25	-33.88	5.63	4.25	5.88
19	2	28	9	24	32	19.25	-33.88	5.63	4.25	5.88
19	2	28	9	25	32	19.25	-33.88	5.63	4.25	5.88
19	2	28	9	26	32	19.19	-33.88	5.63	4.25	5.88
19	2	28	9	27	33	19.19	-33.88	5.63	4.25	5.88
19	2	28	9	28	33	19.19	-33.88	5.63	4.25	5.88
19	2	28	9	29	33	19.19	-33.88	5.63	4.25	5.88
19	2	28	9	30	33	19.12	-33.88	5.63	4.25	5.88
19	2	28	9	31	34	19.12	-33.88	5.63	4.25	5.88
19	2	28	9	32	34	19.12	-33.88	5.63	4.25	5.88
19	2	28	9	33	34	19.12	-33.81	5.63	4.25	5.88
19	2	28	9	34	34	19.19	-33.81	5.63	4.25	5.88
19	2	28	9	35	35	19.19	-33.88	5.69	4.25	5.88
19	2	28	9	36	35	19.19	-33.81	5.63	4.25	5.88
19	2	28	9	37	35	19.19	-33.81	5.69	4.25	5.88
19	2	28	9	38	35	19.25	-33.81	5.69	4.25	5.94
19	2	28	9	39	36	19.25	-33.81	5.69	4.25	5.88
19	2	28	9	40	36	19.25	-33.81	5.69	4.25	5.88
19	2	28	9	41	36	19.25	-33.81	5.69	4.25	5.94
19	2	28	9	42	37	19.25	-33.81	5.69	4.25	5.94
19	2	28	9	43	37	21.94	-33.63	5.69	4	5.13
19	2	28	9	44	37	19.62	-33.75	5.69	4.25	5.75
19	2	28	9	45	37	19.31	-33.81	5.69	4.25	5.94
19	2	28	9	46	38	19.31	-33.81	5.69	4.25	5.94
19	2	28	9	47	38	19.31	-33.81	5.69	4.25	5.94
19	2	28	9	48	38	16.94	-33.88	5.69	4.38	6.31
19	2	28	9	49	38	19.37	-33.81	5.69	4.31	5.94
19	2	28	9	50	39	19.37	-33.81	5.69	4.31	5.94
19	2	28	9	51	39	19.37	-33.81	5.69	4.25	5.94
19	2	28	9	52	39	19.37	-33.81	5.69	4.25	5.94
19	2	28	9	53	39	19.37	-33.81	5.69	4.31	5.94
19	2	28	9	54	40	16.81	-33.88	5.69	4.31	6.25
19	2	28	9	55	40	20.69	-33.75	5.69	4.19	5.69
19	2	28	9	56	40	19.44	-33.81	5.75	4.31	5.94
19	2	28	9	57	40	19.44	-33.81	5.69	4.25	5.94
19	2	28	9	58	41	19.44	-33.81	5.75	4.31	5.94
19	2	28	9	59	41	19.44	-33.81	5.75	4.31	5.94
19	2	28	10	0	41	19.5	-33.75	5.75	4.31	5.94
19	2	28	10	1	41	19.5	-33.75	5.75	4.31	5.94
19	2	28	10	2	42	19.5	-33.81	5.75	4.31	5.94
19	2	28	10	3	42	19.31	-33.75	5.75	4.31	5.94
19	2	28	10	4	42	19.56	-33.75	5.75	4.31	5.94
19	2	28	10	5	42	19.56	-33.75	5.75	4.31	6
19	2	28	10	6	43	19.5	-33.81	5.75	4.31	5.94
19	2	28	10	7	43	19.37	-33.75	5.75	4.31	6
19	2	28	10	8	43	19.06	-33.75	5.75	4.31	6
19	2	28	10	9	43	19.62	-33.75	5.75	4.31	6
19	2	28	10	10	44	19.56	-33.75	5.75	4.31	5.94
19	2	28	10	11	44	19.56	-33.75	5.75	4.31	5.94
19	2	28	10	12	44	19.56	-33.75	5.75	4.31	6
19	2	28	10	13	45	19.62	-33.75	5.75	4.31	6
19	2	28	10	14	45	19.62	-33.75	5.75	4.31	6
19	2	28	10	15	45	19.56	-33.75	5.75	4.31	6
19	2	28	10	16	45	19.62	-33.75	5.75	4.31	6
19	2	28	10	17	46	19.62	-33.75	5.75	4.31	6
19	2	28	10	18	46	19.62	-33.75	5.81	4.31	6
19	2	28	10	19	46	19.69	-33.75	5.75	4.31	6
19	2	28	10	20	46	19.69	-33.75	5.75	4.31	6
19	2	28	10	21	47	19.69	-33.75	5.81	4.31	6
19	2	28	10	22	47	19.75	-33.75	5.75	4.31	6
19	2	28	10	23	47	19.62	-33.75	5.81	4.31	6
19	2	28	10	24	47	19.69	-33.75	5.81	4.31	6
19	2	28	10	25	48	19.81	-33.75	5.81	4.31	6
19	2	28	10	26	48	19.87	-33.75	5.81	4.31	6
19	2	28	10	27	48	19.87	-33.75	5.81	4.31	6
19	2	28	10	28	48	19.56	-33.75	5.81	4.31	6
19	2	28	10	29	49	19.94	-33.75	5.81	4.31	6
19	2	28	10	30	49	19.87	-33.75	5.81	4.31	6
19	2	28	10	31	49	19.94	-33.75	5.81	4.31	6
19	2	28	10	32	49	19.87	-33.75	5.81	4.38	6
19	2	28	10	33	50	19.87	-33.75	5.81	4.38	6
19	2	28	10	34	50	19.87	-33.69	5.81	4.38	6.06
19	2	28	10	35	50	19.87	-33.69	5.81	4.38	6.06
19	2	28	10	36	50	19.81	-33.69	5.81	4.38	6.06
19	2	28	10	37	51	19.69	-33.75	5.81	4.31	6
19	2	28	10	38	51	19.75	-33.69	5.81	4.38	6
19	2	28	10	39	51	19.75	-33.69	5.81	4.38	6
19	2	28	10	40	52	19.69	-33.75	5.81	4.38	6.06
19	2	28	10	41	52	19.56	-33.69	5.81	4.38	6.06
19	2	28	10	42	52	19.56	-33.69	5.88	4.38	6.06
19	2	28	10	43	52	19.75	-33.69	5.81	4.38	6.06
19	2	28	10	44	57	19.75	-33.69	5.81	4.38	6.06
19	2	28	10	45	57	19.56	-33.69	5.88	4.38	6.06
19	2	28	10	46	57	19.75	-33.69	5.88	4.38	6.06
19	2	28	10	47	57	19.81	-33.69	5.88	4.38	6.06
19	2	28	10	48	58	19.69	-33.69	5.88	4.38	6.06
19	2	28	10	49	58	19.75	-33.69	5.88	4.38	6.06
19	2	28	10	50	58	19.69	-33.69	5.88	4.38	6.06
19	2	28	10	51	58	19.69	-33.69	5.88	4.38	6.06
19	2	28	10	52	59	19.75	-33.69	5.88	4.38	6.06
19	2	28	10	53	59	19.75	-33.69	5.88	4.38	6.06
19	2	28	10	54	59	19.75	-33.69	5.88	4.38	6.06
19	2	28	10	55	59	19.75	-33.69	5.88	4.38	6.06
19	2	28	10	57	0	19.75	-33.69	5.88	4.38	6.06
19	2	28	10	58	0	19.75	-33.69	5.88	4.38	6.06
19	2	28	10	59	0	19.75	-33.69	5.88	4.38	6.06
19	2	28	11	0	0	19.75	-33.69	5.88	4.38	6.06
19	2	28	11	1	1	19.75	-33.69	5.88	4.38	6.06
19	2	28	11	2	1	19.75	-33.69	5.88	4.38	6.06
19	2	28	11	3	1	19.75	-33.63	5.94	4.38	6.06
19	2	28	11	4	1	19.75	-33.69	5.88	4.38	6.06
19	2	28	11	5	2	19.81	-33.63	5.94	4.38	6.13
19	2	28	11	6	2	19.81	-33.69	5.88	4.38	6.06
19	2	28	11	7	2	19.81	-33.63	5.94	4.38	6.13
19	2	28	11	8	3	19.81	-33.63	5.94	4.38	6.13
19	2	28	11	9	3	19.87	-33.63	5.94	4.44	6.13
19	2	28	11	10	3	19.87	-33.63	5.94	4.44	6.13
19	2	28	11	11	3	19.94	-33.63	5.94	4.44	6.13
19	2	28	11	12	4	19.94	-33.63	5.94	4.38	6.13
19	2	28	11	13	4	19.94	-33.63	5.94	4.44	6.13
19	2	28	11	14	4	19.94	-33.63	5.94	4.38	6.13
19	2	28	11	15	4	20	-33.63	5.94	4.44	6.13
19	2	28	11	16	5	20	-33.63	5.94	4.44	6.13
19	2	28	11	17	5	20	-33.63	5.94	4.44	6.13
19	2	28	11	18	5	20.06	-33.63	5.94	4.44	6.13
19	2	28	11	19	5	20.06	-33.63	5.94	4.44	6.13
19	2	28	11	20	6	20.06	-33.63	5.94	4.44	6.13
19	2	28	11	21	6	20.06	-33.63	5.94	4.44	6.13
19	2	28	11	22	6	20.06	-33.63	5.94	4.44	6.13
19	2	28	11	23	6	20.06	-33.63	5.94	4.44	6.13
19	2	28	11	24	7	20.06	-33.63	5.94	4.44	6.13
19	2	28	11	25	7	20.06	-33.63	5.94	4.44	6.13
19	2	28	11	26	7	20.06	-33.63	5.94	4.44	6.13
19	2	28	11	27	7	20.06	-33.63	5.94	4.44	6.13
19	2	28	11	28	8	20.06	-33.63	6	4.44	6.13
19	2	28	11	29	8	20.06	-33.63	5.94	4.44	6.19
19	2	28	11	30	8	19.94	-33.63	5.94	4.44	6.13
19	2	28	11	31	8	20.06	-33.63	5.94	4.44	6.13
19	2	28	11	32	9	20.06	-33.56	6	4.44	6.19
19	2	28	11	33	9	20.06	-33.56	6	4.44	6.19
19	2	28	11	34	9	20.06	-33.56	6	4.44	6.19
19	2	28	11	35	10	20.06	-33.63	6	4.44	6.19
19	2	28	11	36	10	20.06	-33.56	6	4.44	6.19
19	2	28	11	37	10	20.06	-33.56	6	4.44	6.19
19	2	28	11	38	10	20.06	-33.56	6	4.44	6.19
19	2	28	11	39	11	20.06	-33.56	6	4.44	6.19
19	2	28	11	40	11	20.06	-33.56	6	4.5	6.19
19	2	28	11	41	11	20.12	-33.56	6	4.44	6.19
19	2	28	11	42	11	20.12	-33.56	6	4.5	6.19
19	2	28	11	43	12	20.12	-33.56	6	4.5	6.19
19	2	28	11	44	12	20.12	-33.56	6	4.44	6.19
19	2	28	11	45	12	20.12	-33.56	6	4.44	6.19
19	2	28	11	46	12	20.12	-33.56	6	4.5	6.19
19	2	28	11	47	13	20.19	-33.56	6	4.5	6.19
19	2	28	11	48	13	20.12	-33.56	6	4.5	6.19
19	2	28	11	49	13	20.19	-33.56	6	4.5	6.19
19	2	28	11	50	13	20.19	-33.56	6	4.5	6.19
19	2	28	11	51	14	20.19	-33.56	6.06	4.5	6.19
19	2	28	11	52	14	20.19	-33.56	6	4.5	6.19
19	2	28	11	53	14	20.19	-33.56	6	4.5	6.19
19	2	28	11	54	14	20.19	-33.56	6	4.5	6.19
19	2	28	11	55	15	20.25	-33.56	6	4.5	6.25
19	2	28	11	56	15	20.25	-33.56	6.06	4.5	6.19
19	2	28	11	57	15	20.25	-33.56	6.06	4.5	6.19
19	2	28	11	58	15	20.25	-33.56	6.06	4.5	6.25
19	2	28	11	59	16	20.25	-33.56	6.06	4.5	6.19
19	2	28	12	0	16	20.25	-33.5	6.06	4.5	6.25
19	2	28	12	1	16	20.25	-33.56	6.06	4.5	6.25
19	2	28	12	2	16	20.25	-33.56	6.06	4.5	6.25
19	2	28	12	3	17	20.31	-33.5	6.06	4.5	6.25
19	2	28	12	4	17	20.25	-33.5	6.06	4.5	6.25
19	2	28	12	5	17	20.25	-33.5	6.06	4.5	6.25
19	2	28	12	6	18	20.31	-33.5	6.06	4.5	6.25
19	2	28	12	7	18	20.31	-33.5	6.06	4.5	6.25
19	2	28	12	8	18	20.31	-33.5	6.06	4.5	6.25
19	2	28	12	9	18	20.31	-33.5	6.06	4.56	6.25
19	2	28	12	10	19	20.37	-33.5	6.06	4.5	6.25
19	2	28	12	11	19	20.37	-33.5	6.06	4.5	6.25
19	2	28	12	12	19	20.37	-33.5	6.06	4.5	6.25
19	2	28	12	13	19	20.44	-33.5	6.06	4.5	6.25
19	2	28	12	14	20	20.44	-33.5	6.06	4.56	6.25
19	2	28	12	15	20	20.44	-33.5	6.06	4.56	6.25
19	2	28	12	16	20	20.5	-33.5	6.06	4.5	6.25
19	2	28	12	17	20	20.56	-33.5	6.13	4.56	6.25
19	2	28	12	18	21	20.56	-33.5	6.06	4.56	6.25
19	2	28	12	19	21	20.62	-33.5	6.13	4.5	6.25
19	2	28	12	20	21	20.62	-33.5	6.13	4.56	6.25
19	2	28	12	21	21	20.69	-33.5	6.13	4.56	6.25
19	2	28	12	22	22	20.69	-33.5	6.13	4.56	6.25
19	2	28	12	23	22	20.75	-33.5	6.13	4.56	6.25
19	2	28	12	24	22	20.69	-33.5	6.13	4.56	6.25
19	2	28	12	25	22	20.75	-33.5	6.13	4.56	6.31
19	2	28	12	26	23	20.75	-33.5	6.13	4.56	6.31
19	2	28	12	27	23	20.75	-33.5	6.13	4.56	6.31
19	2	28	12	28	23	20.81	-33.5	6.13	4.56	6.31
19	2	28	12	29	23	20.81	-33.5	6.13	4.56	6.31
19	2	28	12	30	24	20.81	-33.5	6.13	4.56	6.31
19	2	28	12	31	24	20.87	-33.5	6.13	4.56	6.31
19	2	28	12	32	24	20.81	-33.5	6.13	4.56	6.31
19	2	28	12	33	24	20.87	-33.5	6.13	4.56	6.31
19	2	28	12	34	25	20.87	-33.5	6.13	4.56	6.31
19	2	28	12	35	25	20.87	-33.5	6.13	4.56	6.31
19	2	28	12	36	25	20.87	-33.44	6.19	4.56	6.31
19	2	28	12	37	26	20.94	-33.44	6.13	4.56	6.31
19	2	28	12	38	26	20.94	-33.44	6.19	4.56	6.31
19	2	28	12	39	26	20.94	-33.44	6.19	4.56	6.31
19	2	28	12	40	26	20.94	-33.44	6.13	4.56	6.31
19	2	28	12	41	27	21	-33.44	6.13	4.56	6.31
19	2	28	12	42	27	21	-33.44	6.19	4.56	6.31
19	2	28	12	43	27	21	-33.44	6.19	4.56	6.31
19	2	28	12	44	27	21.06	-33.44	6.19	4.56	6.31
19	2	28	12	45	28	21.06	-33.44	6.19	4.56	6.31
19	2	28	12	46	28	21	-33.44	6.19	4.63	6.31
19	2	28	12	47	28	21.06	-33.44	6.19	4.63	6.31
19	2	28	12	48	28	21.06	-33.44	6.19	4.63	6.38
19	2	28	12	49	29	21	-33.44	6.19	4.63	6.31
19	2	28	12	50	29	21.06	-33.44	6.19	4.56	6.31
19	2	28	12	51	29	21.06	-33.44	6.19	4.63	6.31
19	2	28	12	52	29	21	-33.44	6.19	4.63	6.38
19	2	28	12	53	30	21.06	-33.44	6.19	4.63	6.38
19	2	28	12	54	30	21.06	-33.44	6.19	4.63	6.38
19	2	28	12	55	30	21.06	-33.44	6.19	4.63	6.38
19	2	28	12	56	30	21.06	-33.44	6.19	4.63	6.38
19	2	28	12	57	31	21.06	-33.44	6.19	4.63	6.38
19	2	28	12	58	31	21.06	-33.44	6.19	4.63	6.38
19	2	28	12	59	31	21.06	-33.44	6.19	4.63	6.38
19	2	28	13	0	31	21.12	-33.44	6.19	4.63	6.38
19	2	28	13	1	32	21.06	-33.38	6.25	4.63	6.38
19	2	28	13	2	32	21.06	-33.44	6.25	4.63	6.38
19	2	28	13	3	32	21.12	-33.38	6.25	4.63	6.38
19	2	28	13	4	33	21.12	-33.44	6.25	4.63	6.38
19	2	28	13	5	33	21.12	-33.44	6.25	4.63	6.38
19	2	28	13	6	33	21.12	-33.38	6.25	4.63	6.38
19	2	28	13	7	33	21.12	-33.38	6.25	4.63	6.38
19	2	28	13	8	34	21.12	-33.38	6.25	4.63	6.38
19	2	28	13	9	34	21.12	-33.38	6.25	4.63	6.38
19	2	28	13	10	34	21.19	-33.38	6.25	4.63	6.38
19	2	28	13	11	34	21.12	-33.38	6.25	4.63	6.38
19	2	28	13	12	35	21.12	-33.38	6.25	4.63	6.38
19	2	28	13	13	35	21.19	-33.38	6.25	4.63	6.38
19	2	28	13	14	35	21.19	-33.38	6.25	4.63	6.38
19	2	28	13	15	35	21.12	-33.38	6.25	4.63	6.38
19	2	28	13	16	36	21.06	-33.38	6.25	4.63	6.38
19	2	28	13	17	36	21.12	-33.38	6.25	4.63	6.38
19	2	28	13	18	36	21.12	-33.38	6.25	4.63	6.38
19	2	28	13	19	36	21.19	-33.38	6.25	4.63	6.38
19	2	28	13	20	37	21.25	-33.38	6.25	4.63	6.38
19	2	28	13	21	37	21.19	-33.38	6.25	4.69	6.44
19	2	28	13	22	37	21.25	-33.38	6.25	4.63	6.44
19	2	28	13	23	37	21.25	-33.38	6.31	4.69	6.44
19	2	28	13	24	38	21.25	-33.38	6.25	4.69	6.44
19	2	28	13	25	38	21.31	-33.38	6.25	4.63	6.44
19	2	28	13	26	38	21.31	-33.38	6.25	4.69	6.44
19	2	28	13	27	38	21.37	-33.38	6.31	4.69	6.44
19	2	28	13	28	39	21.37	-33.38	6.31	4.69	6.44
19	2	28	13	29	39	21.44	-33.38	6.31	4.69	6.44
19	2	28	13	30	39	21.5	-33.38	6.31	4.69	6.44
19	2	28	13	31	40	21.5	-33.31	6.31	4.69	6.44
19	2	28	13	32	40	21.56	-33.31	6.31	4.69	6.44
19	2	28	13	33	40	21.56	-33.31	6.31	4.69	6.44
19	2	28	13	34	40	21.62	-33.31	6.31	4.69	6.44
19	2	28	13	35	41	21.62	-33.31	6.31	4.69	6.44
19	2	28	13	36	41	21.69	-33.31	6.31	4.69	6.44
19	2	28	13	37	41	21.37	-33.31	6.31	4.69	6.44
19	2	28	13	38	41	21.62	-33.31	6.31	4.69	6.44
19	2	28	13	39	42	21.69	-33.31	6.31	4.69	6.44
19	2	28	13	40	42	21.5	-33.31	6.31	4.69	6.5
19	2	28	13	41	42	21.44	-33.31	6.31	4.69	6.44
19	2	28	13	42	42	21.25	-33.31	6.31	4.69	6.44
19	2	28	13	43	43	21.44	-33.31	6.31	4.69	6.44
19	2	28	13	44	43	21.37	-33.31	6.31	4.69	6.5
19	2	28	13	45	43	21.37	-33.31	6.38	4.69	6.5
19	2	28	13	46	43	21.31	-33.31	6.38	4.69	6.5
19	2	28	13	47	44	21.25	-33.31	6.38	4.69	6.5
19	2	28	13	48	44	21.25	-33.31	6.38	4.69	6.5
19	2	28	13	49	44	21.19	-33.31	6.31	4.69	6.5
19	2	28	13	50	44	21.12	-33.31	6.38	4.69	6.5
19	2	28	13	51	45	21.06	-33.31	6.38	4.75	6.5
19	2	28	13	52	45	21.06	-33.31	6.38	4.69	6.5
19	2	28	13	53	45	21	-33.31	6.38	4.69	6.5
19	2	28	13	54	45	20.94	-33.31	6.38	4.75	6.5
19	2	28	13	55	46	20.94	-33.31	6.38	4.75	6.5
19	2	28	13	56	46	20.94	-33.25	6.38	4.75	6.5
19	2	28	13	57	46	20.87	-33.25	6.38	4.75	6.5
19	2	28	13	58	47	20.87	-33.25	6.38	4.75	6.5
19	2	28	13	59	47	20.94	-33.25	6.38	4.75	6.5
19	2	28	14	0	47	20.87	-33.25	6.38	4.75	6.5
19	2	28	14	1	47	20.94	-33.25	6.38	4.75	6.5
19	2	28	14	2	48	20.94	-33.25	6.38	4.75	6.5
19	2	28	14	3	48	20.94	-33.25	6.38	4.75	6.5
19	2	28	14	4	48	20.94	-33.25	6.38	4.75	6.5
19	2	28	14	5	48	21	-33.25	6.38	4.75	6.5
19	2	28	14	6	49	21	-33.25	6.38	4.75	6.5
19	2	28	14	7	49	20.94	-33.25	6.38	4.75	6.5
19	2	28	14	8	49	20.94	-33.25	6.38	4.75	6.56
19	2	28	14	9	49	20.94	-33.25	6.44	4.75	6.5
19	2	28	14	10	50	20.87	-33.25	6.44	4.75	6.5
19	2	28	14	11	50	20.81	-33.25	6.44	4.75	6.56
19	2	28	14	12	50	20.81	-33.25	6.44	4.75	6.56
19	2	28	14	13	50	20.75	-33.25	6.44	4.75	6.56
19	2	28	14	14	51	20.75	-33.25	6.44	4.75	6.56
19	2	28	14	15	51	20.69	-33.25	6.44	4.75	6.56
19	2	28	14	16	51	20.69	-33.25	6.44	4.75	6.56
19	2	28	14	17	51	20.69	-33.25	6.44	4.75	6.56
19	2	28	14	18	52	20.69	-33.19	6.44	4.75	6.56
19	2	28	14	19	52	20.69	-33.19	6.44	4.75	6.56
19	2	28	14	20	52	20.69	-33.19	6.44	4.75	6.56
19	2	28	14	21	52	20.69	-33.19	6.44	4.75	6.56
19	2	28	14	22	53	20.69	-33.19	6.44	4.75	6.56
19	2	28	14	23	53	20.69	-33.19	6.44	4.75	6.56
19	2	28	14	24	53	20.69	-33.19	6.44	4.75	6.56
19	2	28	14	25	54	20.75	-33.19	6.44	4.81	6.56
19	2	28	14	26	54	20.75	-33.19	6.44	4.81	6.56
19	2	28	14	27	54	20.75	-33.19	6.44	4.81	6.56
19	2	28	14	28	54	20.75	-33.19	6.44	4.81	6.56
19	2	28	14	29	55	20.75	-33.19	6.44	4.81	6.56
19	2	28	14	30	55	20.75	-33.19	6.44	4.81	6.56
19	2	28	14	31	55	20.75	-33.19	6.5	4.81	6.56
19	2	28	14	32	55	20.75	-33.19	6.44	4.81	6.56
19	2	28	14	33	56	20.75	-33.19	6.44	4.81	6.56
19	2	28	14	34	59	20.75	-33.19	6.44	4.81	6.56
19	2	28	14	36	0	20.81	-33.19	6.5	4.81	6.56
19	2	28	14	37	0	20.81	-33.19	6.44	4.81	6.56
19	2	28	14	38	0	20.87	-33.19	6.5	4.81	6.56
19	2	28	14	39	0	20.87	-33.19	6.5	4.81	6.56
19	2	28	14	40	1	20.81	-33.19	6.44	4.81	6.56
19	2	28	14	41	1	20.87	-33.13	6.5	4.81	6.63
19	2	28	14	42	1	20.87	-33.19	6.5	4.81	6.56
19	2	28	14	43	1	20.87	-33.19	6.5	4.81	6.63
19	2	28	14	44	2	20.87	-33.19	6.5	4.81	6.56
19	2	28	14	45	2	20.81	-33.13	6.5	4.81	6.63
19	2	28	14	46	2	20.81	-33.13	6.5	4.81	6.63
19	2	28	14	47	2	20.75	-33.13	6.5	4.81	6.63
19	2	28	14	48	3	20.75	-33.13	6.5	4.81	6.63
19	2	28	14	49	3	20.69	-33.13	6.5	4.81	6.63
19	2	28	14	50	3	20.69	-33.13	6.5	4.81	6.63
19	2	28	14	51	3	20.69	-33.13	6.5	4.81	6.63
19	2	28	14	52	4	20.69	-33.13	6.5	4.81	6.63
19	2	28	14	53	4	20.69	-33.13	6.5	4.81	6.63
19	2	28	14	54	4	20.69	-33.13	6.5	4.88	6.63
19	2	28	14	55	5	20.69	-33.13	6.5	4.81	6.63
19	2	28	14	56	5	20.69	-33.13	6.5	4.81	6.63
19	2	28	14	57	5	20.69	-33.13	6.5	4.88	6.63
19	2	28	14	58	5	20.69	-33.13	6.5	4.88	6.63
19	2	28	14	59	6	20.69	-33.13	6.5	4.81	6.63
19	2	28	15	0	6	20.69	-33.13	6.5	4.88	6.63
19	2	28	15	1	6	20.69	-33.13	6.5	4.81	6.63
19	2	28	15	2	6	20.69	-33.13	6.5	4.88	6.63
19	2	28	15	3	7	20.75	-33.13	6.5	4.88	6.63
19	2	28	15	4	7	20.69	-33.13	6.5	4.88	6.63
19	2	28	15	5	7	20.69	-33.06	6.5	4.88	6.63
19	2	28	15	6	7	20.62	-33.13	6.56	4.88	6.63
19	2	28	15	7	8	20.62	-33.13	6.56	4.88	6.63
19	2	28	15	8	8	20.56	-33.13	6.5	4.88	6.63
19	2	28	15	9	8	20.56	-33.13	6.56	4.88	6.63
19	2	28	15	10	8	20.5	-33.06	6.56	4.88	6.63
19	2	28	15	11	9	20.44	-33.06	6.56	4.88	6.69
19	2	28	15	12	9	20	-33.13	6.56	4.88	6.69
19	2	28	15	13	9	20.37	-33.06	6.56	4.88	6.69
19	2	28	15	14	9	20.19	-33.06	6.56	4.88	6.69
19	2	28	15	15	10	20.5	-33.06	6.56	4.88	6.69
19	2	28	15	16	10	20.31	-33.06	6.56	4.88	6.69
19	2	28	15	17	10	20.44	-33.06	6.56	4.88	6.69
19	2	28	15	18	10	19.81	-33.06	6.56	4.88	6.69
19	2	28	15	19	11	20.31	-33.06	6.56	4.88	6.69
19	2	28	15	20	11	20.31	-33.06	6.56	4.88	6.63
19	2	28	15	21	11	20.31	-33.06	6.56	4.88	6.69
19	2	28	15	22	11	20.25	-33.06	6.56	4.88	6.69
19	2	28	15	23	12	20.06	-33.06	6.56	4.88	6.69
19	2	28	15	24	12	20.31	-33.06	6.56	4.88	6.69
19	2	28	15	25	12	20.37	-33.06	6.56	4.88	6.69
19	2	28	15	26	13	20.44	-33.06	6.56	4.88	6.69
19	2	28	15	27	13	20.44	-33.06	6.56	4.94	6.69
19	2	28	15	28	13	20.44	-33.06	6.56	4.94	6.69
19	2	28	15	29	13	20.44	-33.06	6.56	4.94	6.69
19	2	28	15	30	14	20.44	-33.06	6.56	4.94	6.69
19	2	28	15	31	14	20.37	-33.06	6.56	4.88	6.69
19	2	28	15	32	14	20.37	-33.06	6.56	4.88	6.69
19	2	28	15	33	14	20.37	-33.06	6.56	4.94	6.69
19	2	28	15	34	15	20.44	-33.06	6.56	4.88	6.69
19	2	28	15	35	15	20.37	-33.06	6.56	4.88	6.69
19	2	28	15	36	15	20.44	-33.06	6.56	4.94	6.69
19	2	28	15	37	15	20.44	-33.06	6.56	4.88	6.69
19	2	28	15	38	16	20.44	-33.06	6.56	4.94	6.69
19	2	28	15	39	16	20.37	-33	6.56	4.88	6.69
19	2	28	15	40	16	20.37	-33	6.56	4.94	6.69
19	2	28	15	41	16	20.37	-33.06	6.63	4.94	6.69
19	2	28	15	42	17	20.31	-33	6.63	4.94	6.69
19	2	28	15	43	17	20.31	-33	6.63	4.94	6.69
19	2	28	15	44	17	20.31	-33	6.63	4.94	6.69
19	2	28	15	45	17	20.25	-33	6.63	4.94	6.69
19	2	28	15	46	18	20.25	-33	6.63	4.94	6.69
19	2	28	15	47	18	20.25	-33	6.63	4.94	6.69
19	2	28	15	48	18	20.25	-33	6.63	4.94	6.69
19	2	28	15	49	18	20.25	-33	6.63	4.94	6.69
19	2	28	15	50	19	20.25	-33	6.63	4.94	6.69
19	2	28	15	51	19	20.25	-33	6.63	4.94	6.75
19	2	28	15	52	19	20.31	-33	6.63	4.94	6.69
19	2	28	15	53	20	20.31	-33	6.63	4.94	6.75
19	2	28	15	54	20	20.37	-33	6.63	4.94	6.75
19	2	28	15	55	20	20.37	-33	6.63	4.94	6.75
19	2	28	15	56	20	20.44	-33	6.63	4.94	6.75
19	2	28	15	57	21	20.44	-33	6.63	4.94	6.75
19	2	28	15	58	21	20.44	-33	6.63	4.94	6.75
19	2	28	15	59	21	20.5	-33	6.63	4.94	6.75
19	2	28	16	0	21	20.5	-33	6.63	4.94	6.75
19	2	28	16	1	22	20.5	-33	6.63	4.94	6.75
19	2	28	16	2	22	20.56	-33	6.63	5	6.75
19	2	28	16	3	22	20.56	-32.94	6.63	5	6.75
19	2	28	16	4	22	20.5	-33	6.63	4.94	6.75
19	2	28	16	5	23	20.5	-32.94	6.63	5	6.75
19	2	28	16	6	23	20.5	-33	6.63	5	6.75
19	2	28	16	7	23	20.44	-33	6.63	5	6.75
19	2	28	16	8	23	20.37	-32.94	6.63	4.94	6.75
19	2	28	16	9	24	20.37	-32.94	6.63	5	6.75
19	2	28	16	10	24	20.37	-32.94	6.63	5	6.75
19	2	28	16	11	24	20.31	-32.94	6.63	5	6.75
19	2	28	16	12	24	20.31	-32.94	6.63	5	6.75
19	2	28	16	13	25	20.25	-32.94	6.63	5	6.75
19	2	28	16	14	25	20.25	-32.94	6.63	5	6.75
19	2	28	16	15	25	20.25	-32.94	6.63	5	6.75
19	2	28	16	16	25	20.25	-32.94	6.69	5	6.75
19	2	28	16	17	26	20.25	-32.94	6.69	5	6.75
19	2	28	16	18	26	20.25	-32.94	6.63	5	6.75
19	2	28	16	19	26	20.25	-32.94	6.69	5	6.75
19	2	28	16	20	27	20.25	-32.94	6.69	5	6.75
19	2	28	16	21	27	20.25	-32.94	6.69	5	6.81
19	2	28	16	22	27	20.31	-32.94	6.69	5	6.81
19	2	28	16	23	27	20.31	-32.94	6.69	5	6.81
19	2	28	16	24	28	20.31	-32.94	6.69	5	6.75
19	2	28	16	25	28	20.31	-32.94	6.69	5	6.81
19	2	28	16	26	28	20.31	-32.88	6.69	5	6.81
19	2	28	16	27	28	20.31	-32.94	6.69	5	6.81
19	2	28	16	28	29	20.37	-32.94	6.69	5	6.81
19	2	28	16	29	29	20.31	-32.94	6.69	5	6.81
19	2	28	16	30	29	20.31	-32.94	6.69	5	6.81
19	2	28	16	31	29	20.31	-32.88	6.69	5.06	6.81
19	2	28	16	32	30	20.31	-32.88	6.69	5	6.81
19	2	28	16	33	30	20.25	-32.94	6.69	5.06	6.81
19	2	28	16	34	30	20.19	-32.88	6.69	5.06	6.81
19	2	28	16	35	30	20.19	-32.88	6.69	5.06	6.81
19	2	28	16	36	31	20.19	-32.88	6.69	5.06	6.81
19	2	28	16	37	31	20.12	-32.88	6.69	5	6.81
19	2	28	16	38	31	20.12	-32.88	6.69	5.06	6.81
19	2	28	16	39	31	20.12	-32.88	6.69	5.06	6.81
19	2	28	16	40	32	20.12	-32.88	6.69	5.06	6.81
19	2	28	16	41	32	20.19	-32.88	6.69	5.06	6.81
19	2	28	16	42	32	20.12	-32.88	6.69	5.06	6.81
19	2	28	16	43	32	20.12	-32.88	6.69	5.06	6.81
19	2	28	16	44	33	20.19	-32.88	6.69	5.06	6.81
19	2	28	16	45	33	20.12	-32.88	6.69	5.06	6.81
19	2	28	16	46	33	20.19	-32.88	6.69	5.06	6.81
19	2	28	16	47	34	20.19	-32.88	6.69	5.06	6.81
19	2	28	16	48	34	20.19	-32.88	6.75	5.06	6.81
19	2	28	16	49	34	20.19	-32.88	6.75	5.06	6.81
19	2	28	16	50	34	20.12	-32.88	6.69	5.06	6.81
19	2	28	16	51	35	20.19	-32.88	6.75	5.06	6.81
19	2	28	16	52	35	20.19	-32.88	6.75	5.06	6.81
19	2	28	16	53	35	20.19	-32.88	6.75	5.06	6.81
19	2	28	16	54	35	20.19	-32.88	6.75	5.06	6.81
19	2	28	16	55	36	20.19	-32.88	6.75	5.06	6.81
19	2	28	16	56	36	20.19	-32.88	6.75	5.06	6.81
19	2	28	16	57	36	20.19	-32.88	6.75	5.06	6.81
19	2	28	16	58	36	20.25	-32.88	6.75	5.06	6.88
19	2	28	16	59	37	20.25	-32.88	6.75	5.06	6.81
19	2	28	17	0	37	20.19	-32.81	6.75	5.06	6.81
19	2	28	17	1	37	20.19	-32.88	6.75	5.06	6.81
19	2	28	17	2	37	20.19	-32.81	6.75	5.06	6.81
19	2	28	17	3	38	20.19	-32.81	6.75	5.06	6.81
19	2	28	17	4	38	20.12	-32.88	6.75	5.06	6.88
19	2	28	17	5	38	20.12	-32.81	6.75	5.06	6.88
19	2	28	17	6	39	20.12	-32.81	6.75	5.06	6.88
19	2	28	17	7	39	20.12	-32.81	6.75	5.06	6.88
19	2	28	17	8	39	20.12	-32.81	6.75	5.06	6.88
19	2	28	17	9	39	20.06	-32.81	6.75	5.06	6.88
19	2	28	17	10	40	20.12	-32.81	6.75	5.06	6.88
19	2	28	17	11	40	20.12	-32.81	6.75	5.06	6.88
19	2	28	17	12	40	20.12	-32.81	6.75	5.06	6.88
19	2	28	17	13	40	20.12	-32.81	6.75	5.06	6.88
19	2	28	17	14	41	20.12	-32.81	6.75	5.06	6.88
19	2	28	17	15	41	20.19	-32.81	6.75	5.13	6.88
19	2	28	17	16	41	20.19	-32.81	6.75	5.13	6.88
19	2	28	17	17	41	20.19	-32.81	6.75	5.13	6.88
19	2	28	17	18	42	20.25	-32.81	6.75	5.13	6.88
19	2	28	17	19	42	20.25	-32.81	6.75	5.13	6.88
19	2	28	17	20	42	20.25	-32.81	6.75	5.13	6.88
19	2	28	17	21	42	20.25	-32.81	6.75	5.13	6.88
19	2	28	17	22	43	20.31	-32.81	6.75	5.13	6.88
19	2	28	17	23	43	20.31	-32.81	6.75	5.06	6.88
19	2	28	17	24	43	20.31	-32.81	6.75	5.13	6.88
19	2	28	17	25	43	20.31	-32.81	6.75	5.13	6.88
19	2	28	17	26	44	20.31	-32.81	6.81	5.13	6.88
19	2	28	17	27	44	20.31	-32.81	6.75	5.13	6.88
19	2	28	17	28	44	20.31	-32.81	6.81	5.13	6.88
19	2	28	17	29	44	20.25	-32.81	6.81	5.13	6.88
19	2	28	17	30	45	20.25	-32.81	6.81	5.13	6.88
19	2	28	17	31	45	20.19	-32.81	6.81	5.13	6.88
19	2	28	17	32	45	20.19	-32.81	6.81	5.13	6.88
19	2	28	17	33	46	20.19	-32.75	6.81	5.13	6.88
19	2	28	17	34	46	20.12	-32.81	6.81	5.13	6.94
19	2	28	17	35	46	20.12	-32.75	6.81	5.13	6.88
19	2	28	17	36	46	20.12	-32.81	6.81	5.13	6.94
19	2	28	17	37	47	20.12	-32.75	6.81	5.13	6.94
19	2	28	17	38	47	20.12	-32.75	6.81	5.13	6.94
19	2	28	17	39	47	20.12	-32.75	6.81	5.13	6.94
19	2	28	17	40	47	20.12	-32.75	6.81	5.13	6.94
19	2	28	17	41	48	20.12	-32.75	6.81	5.13	6.94
19	2	28	17	42	48	20.19	-32.75	6.81	5.13	6.94
19	2	28	17	43	48	20.19	-32.75	6.81	5.13	6.94
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	2	28	17	47	32	20.31	-32.81	6.81	5.19	6.88
19	2	28	17	48	32	20.31	-32.81	6.81	5.13	6.94
19	2	28	17	49	32	20.31	-32.75	6.75	5.19	6.88
19	2	28	17	50	32	21.44	-32.38	6.75	5.19	6.13
19	2	28	17	51	33	20.37	-32.69	6.75	5.19	6.94
19	2	28	17	52	33	20.37	-32.69	6.75	5.19	6.88
19	2	28	17	53	33	20.37	-32.69	6.75	5.19	6.88
19	2	28	17	54	34	20.31	-32.69	6.69	5.19	6.94
19	2	28	17	55	34	20.31	-32.75	6.75	5.19	6.88
19	2	28	17	56	34	20.31	-32.69	6.69	5.19	6.88
19	2	28	17	57	34	20.25	-32.75	6.75	5.19	6.94
19	2	28	17	58	35	20.25	-32.75	6.75	5.19	6.94
19	2	28	17	59	35	20.25	-32.69	6.69	5.19	6.94
19	2	28	18	0	35	20.19	-32.75	6.69	5.19	6.94
19	2	28	18	1	35	20.19	-32.75	6.69	5.19	6.94
19	2	28	18	2	36	12.75	-33.56	6.75	5.38	8.94
19	2	28	18	3	36	20.12	-32.75	6.69	5.19	6.94
19	2	28	18	4	36	20.12	-32.75	6.69	5.19	6.94
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	2	28	18	8	13	20.06	-32.88	6.75	5.19	6.94
19	2	28	18	9	13	20	-32.88	6.81	5.19	6.94
19	2	28	18	10	14	20	-32.81	6.75	5.19	6.94
19	2	28	18	11	14	19.94	-32.81	6.75	5.19	6.94
19	2	28	18	12	14	19.94	-32.75	6.75	5.19	6.94
19	2	28	18	13	14	19.94	-32.81	6.75	5.19	6.94
19	2	28	18	14	15	19.94	-32.75	6.75	5.19	6.94
19	2	28	18	15	15	19.94	-32.75	6.69	5.19	6.94
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	2	28	18	18	44	19.94	-32.81	6.75	5.19	7
19	2	28	18	19	44	19.94	-32.88	6.81	5.19	7
19	2	28	18	20	45	19.94	-32.81	6.75	5.25	7
19	2	28	18	21	45	19.94	-32.81	6.75	5.25	6.94
19	2	28	18	22	45	19.87	-32.75	6.75	5.25	7
19	2	28	18	23	45	19.94	-32.81	6.75	5.25	7
19	2	28	18	24	46	19.94	-32.75	6.75	5.25	7
19	2	28	18	25	46	19.87	-32.69	6.69	5.25	7
19	2	28	18	26	46	19.87	-32.75	6.75	5.25	7
19	2	28	18	27	47	19.87	-32.69	6.69	5.25	7
19	2	28	18	28	47	19.87	-32.75	6.75	5.25	7
19	2	28	18	29	47	19.87	-32.75	6.75	5.25	7
19	2	28	18	30	47	19.87	-32.69	6.69	5.25	7
19	2	28	18	31	48	19.87	-32.75	6.75	5.25	7
19	2	28	18	32	48	19.87	-32.69	6.69	5.25	7
19	2	28	18	33	48	19.87	-32.75	6.75	5.25	7
19	2	28	18	34	48	19.87	-32.75	6.75	5.25	7
19	2	28	18	35	49	19.81	-32.69	6.69	5.25	7
19	2	28	18	36	49	19.81	-32.75	6.75	5.25	7
19	2	28	18	37	49	19.81	-32.69	6.75	5.25	7
19	2	28	18	38	49	19.81	-32.75	6.75	5.25	7
19	2	28	18	39	50	19.75	-32.75	6.75	5.25	7
19	2	28	18	40	50	19.75	-32.69	6.69	5.25	7
19	2	28	18	41	50	19.69	-32.75	6.75	5.25	7
19	2	28	18	42	50	19.69	-32.69	6.69	5.25	7
19	2	28	18	43	51	19.69	-32.75	6.75	5.25	7
19	2	28	18	44	51	19.62	-32.69	6.75	5.25	7
19	2	28	18	45	51	19.62	-32.69	6.69	5.25	7.06
19	2	28	18	46	52	19.62	-32.75	6.75	5.25	7.06
19	2	28	18	47	52	19.56	-32.69	6.69	5.25	7.06
19	2	28	18	48	52	19.56	-32.75	6.75	5.25	7.06
19	2	28	18	49	52	19.56	-32.69	6.75	5.25	7.06
19	2	28	18	50	53	19.56	-32.69	6.69	5.25	7.06
19	2	28	18	51	53	19.5	-32.75	6.75	5.25	7.06
19	2	28	18	52	53	19.5	-32.75	6.69	5.31	7.06
19	2	28	18	53	53	19.5	-32.75	6.75	5.25	7.06
19	2	28	18	54	54	19.44	-32.75	6.69	5.31	7.06
19	2	28	18	55	54	19.44	-32.69	6.69	5.31	7.06
19	2	28	18	56	54	19.37	-32.75	6.75	5.31	7.06
19	2	28	18	57	54	19.37	-32.63	6.38	5.31	7.06
19	2	28	18	58	55	19.37	-32.69	6.44	5.31	7.06
19	2	28	18	59	55	19.37	-32.63	6.19	5.31	7.06
19	2	28	19	0	55	19.37	-32.63	6.19	5.31	7.06
19	2	28	19	1	55	19.37	-32.69	6.31	5.31	7.06
19	2	28	19	2	56	19.44	-32.63	6	5.31	7
19	2	28	19	3	56	19.44	-32.69	6.13	5.31	7
19	2	28	19	4	56	19.44	-32.63	5.88	5.31	7
19	2	28	19	5	56	19.44	-32.63	5.94	5.31	7
19	2	28	19	6	57	19.44	-32.69	6.06	5.31	7
19	2	28	19	7	57	19.44	-32.63	5.75	5.31	7
19	2	28	19	8	57	19.44	-32.69	5.94	5.31	7
19	2	28	19	9	57	19.37	-32.63	5.63	5.31	7
19	2	28	19	10	58	19.44	-32.69	5.25	5.31	7
19	2	28	19	11	58	19.37	-32.69	5.56	5.31	6.94
19	2	28	19	12	58	19.37	-32.75	5.25	5.31	6.94
19	2	28	19	13	59	19.37	-32.75	5.44	5.31	6.94
19	2	28	19	14	59	19.37	-32.75	5.19	5.31	6.94
19	2	28	19	16	3	19.37	-32.75	4.88	5.31	6.94
19	2	28	19	17	3	19.37	-32.81	5.19	5.31	6.94
19	2	28	19	18	4	19.37	-32.75	4.69	5.25	6.94
19	2	28	19	19	4	19.37	-32.75	4.94	5.25	6.94
19	2	28	19	20	4	19.37	-32.75	4.5	5.31	6.88
19	2	28	19	21	4	19.37	-32.81	3.94	5.25	6.88
19	2	28	19	22	5	19.31	-32.81	4.44	5.25	6.88
19	2	28	19	23	5	19.31	-32.88	3.94	5.25	6.81
19	2	28	19	24	5	19.31	-32.88	4.31	5.25	6.81
19	2	28	19	25	5	19.31	-32.94	3.88	5.25	6.81
19	2	28	19	26	6	19.31	-32.94	4.25	5.25	6.81
19	2	28	19	27	6	19.31	-32.94	4.5	5.25	6.81
19	2	28	19	28	6	19.25	-32.94	3.63	5.25	6.81
19	2	28	19	29	6	19.31	-33	4.06	5.25	6.75
19	2	28	19	30	7	19.31	-33	3.38	5.25	6.75
19	2	28	19	31	7	19.31	-33.06	3.75	5.25	6.75
19	2	28	19	32	7	19.31	-33	3.94	5.25	6.75
19	2	28	19	33	7	19.31	-33.06	3.19	5.25	6.75
19	2	28	19	34	8	19.31	-33.06	3.63	5.25	6.75
19	2	28	19	35	8	19.31	-33.06	3	5.25	6.75
19	2	28	19	36	8	19.25	-33.13	3.44	5.25	6.69
19	2	28	19	37	9	19.25	-33.19	2.69	5.25	6.69
19	2	28	19	38	9	19.25	-33.25	1.75	5.25	6.69
19	2	28	19	39	9	19.19	-33.19	2.5	5.25	6.63
19	2	28	19	40	9	19.19	-33.25	1.75	5.19	6.63
19	2	28	19	41	10	19.19	-33.19	2.38	5.19	6.63
19	2	28	19	42	10	19.19	-33.25	1.75	5.19	6.56
19	2	28	19	43	10	19.12	-33.31	0.88	5.19	6.56
19	2	28	19	44	10	19.12	-33.25	1.75	5.19	6.56
19	2	28	19	45	11	19.12	-33.25	1.06	5.19	6.5
19	2	28	19	46	11	19.12	-33.19	1.75	5.19	6.5
19	2	28	19	47	11	19.12	-33.25	1	5.19	6.5
19	2	28	19	48	11	19.12	-33.19	1.63	5.19	6.44
19	2	28	19	49	12	19.12	-33.19	1.31	5.19	6.44
19	2	28	19	50	12	19.06	-33.25	0.31	5.19	6.44
19	2	28	19	51	12	19.06	-33.19	1.19	5.13	6.44
19	2	28	19	52	12	19.06	-33.19	0.44	5.19	6.44
19	2	28	19	53	13	19	-33.19	1.12	5.19	6.38
19	2	28	19	54	13	19	-33.19	0.5	5.13	6.38
19	2	28	19	55	13	19	-33.25	-0.5	5.13	6.38
19	2	28	19	56	14	19	-33.19	0.56	5.13	6.31
19	2	28	19	57	14	19	-33.19	-0.37	5.13	6.31
19	2	28	19	58	14	19	-33.13	0.44	5.13	6.31
19	2	28	19	59	14	19	-33.19	-0.31	5.13	6.25
19	2	28	20	0	15	19	-33.13	0.37	5.13	6.25
19	2	28	20	1	15	19	-33.13	0.19	5.13	6.25
19	2	28	20	2	15	19	-33.13	-0.88	5.06	6.25
19	2	28	20	3	15	19	-33.06	0.06	5.06	6.25
19	2	28	20	4	16	19	-33.13	-0.75	5.06	6.25
19	2	28	20	5	16	19	-33.06	0.06	5.06	6.19
19	2	28	20	6	16	19	-33.06	-0.63	5.06	6.19
19	2	28	20	7	16	19	-33.19	-1.94	5.06	6.19
19	2	28	20	8	17	19	-33.06	-0.75	5.06	6.13
19	2	28	20	9	17	19	-33.13	-1.69	5.06	6.13
19	2	28	20	10	17	19	-33.06	-0.75	5.06	6.13
19	2	28	20	11	17	19	-33.06	-1.56	5.06	6.06
19	2	28	20	12	18	19	-33	-0.75	5.06	6.06
19	2	28	20	13	18	19	-33	-1.44	5.06	6
19	2	28	20	14	18	19	-33.13	-2.56	5	6
19	2	28	20	15	18	19	-33	-1.37	5	6
19	2	28	20	16	19	19	-33.06	-2.19	5	6
19	2	28	20	17	19	19	-33	-1.25	5	5.94
19	2	28	20	18	19	19	-33	-2.13	5	5.94
19	2	28	20	19	19	19	-32.94	-1.31	5	5.94
19	2	28	20	20	20	18.94	-33	-2.06	5	5.88
19	2	28	20	21	20	18.94	-33.06	-3.13	4.94	5.88
19	2	28	20	22	20	18.94	-33	-1.94	4.94	5.88
19	2	28	20	23	20	18.87	-33	-2.88	4.94	5.88
19	2	28	20	24	21	18.94	-32.94	-1.88	4.94	5.81
19	2	28	20	25	21	18.94	-33	-2.69	4.94	5.81
19	2	28	20	26	21	18.87	-33.06	-3.06	4.94	5.81
19	2	28	20	27	22	18.87	-32.94	-2.38	4.94	5.75
19	2	28	20	28	22	18.87	-33	-3.31	4.94	5.75
19	2	28	20	29	22	18.87	-32.94	-2.25	4.94	5.75
19	2	28	20	30	22	18.87	-32.94	-3	4.94	5.75
19	2	28	20	31	23	18.94	-32.88	-2.06	4.94	5.69
19	2	28	20	32	23	18.87	-32.94	-2.81	4.88	5.69
19	2	28	20	33	23	18.87	-33	-3.94	4.94	5.69
19	2	28	20	34	23	18.87	-32.88	-2.56	4.94	5.69
19	2	28	20	35	24	18.87	-32.94	-3.44	4.88	5.69
19	2	28	20	36	24	18.87	-32.88	-2.38	4.88	5.69
19	2	28	20	37	24	18.87	-32.88	-2.88	4.88	5.69
19	2	28	20	38	24	18.87	-32.88	-2	4.88	5.69
19	2	28	20	39	25	18.87	-32.88	-2.44	4.88	5.69
19	2	28	20	40	25	18.87	-32.94	-3.19	4.88	5.69
19	2	28	20	41	25	18.81	-32.88	-2.19	4.88	5.69
19	2	28	20	42	25	17.44	-33	-2.75	4.88	5.75
19	2	28	20	43	26	18.87	-32.88	-1.94	4.88	5.69
19	2	28	20	44	26	18.81	-32.94	-2.44	4.88	5.69
19	2	28	20	45	26	18.81	-33	-2.94	4.88	5.69
19	2	28	20	46	27	18.75	-32.94	-2.19	4.88	5.69
19	2	28	20	47	27	18.75	-32.94	-2.5	4.88	5.69
19	2	28	20	48	27	18.75	-32.94	-1.69	4.88	5.75
19	2	28	20	49	27	18.75	-32.94	-2	4.88	5.75
19	2	28	20	50	28	18.75	-32.94	-1.31	4.88	5.75
19	2	28	20	51	28	18.75	-32.94	-1.56	4.88	5.75
19	2	28	20	52	28	18.75	-33	-2.13	4.88	5.75
19	2	28	20	53	28	18.75	-32.94	-1.31	4.88	5.75
19	2	28	20	54	29	18.75	-33	-1.75	4.88	5.81
19	2	28	20	55	29	18.75	-32.94	-1.06	4.88	5.81
19	2	28	20	56	29	18.75	-33	-1.44	4.88	5.81
19	2	28	20	57	29	18.75	-32.94	-1.37	4.88	5.81
19	2	28	20	58	30	18.75	-32.88	-0.69	4.88	5.88
19	2	28	20	59	30	18.75	-32.81	-0.81	4.88	5.88
19	2	28	21	0	30	18.75	-32.81	-0.31	4.88	5.88
19	2	28	21	1	30	18.75	-32.81	-0.44	4.88	5.94
19	2	28	21	2	31	18.75	-32.88	-0.75	4.88	5.94
19	2	28	21	3	31	18.69	-32.88	-0.12	4.88	5.94
19	2	28	21	4	31	18.75	-32.81	-0.31	4.88	6
19	2	28	21	5	31	18.75	-32.75	0.19	4.88	6
19	2	28	21	6	32	18.75	-32.81	-0.06	4.88	6
19	2	28	21	7	32	18.75	-32.81	0.37	4.88	6
19	2	28	21	8	32	18.75	-32.81	0.75	4.88	6.06
19	2	28	21	9	33	18.75	-32.69	0.88	4.94	6.06
19	2	28	21	10	33	18.75	-32.75	1.19	4.94	6.13
19	2	28	21	11	33	18.75	-32.69	1.37	4.94	6.13
19	2	28	21	12	33	18.75	-32.75	1.44	4.94	6.19
19	2	28	21	13	34	18.69	-32.75	1.69	4.94	6.19
19	2	28	21	14	34	18.69	-32.69	1.81	4.94	6.25
19	2	28	21	15	34	18.69	-32.75	2	4.94	6.25
19	2	28	21	16	34	18.69	-32.75	2.13	5	6.31
19	2	28	21	17	35	18.69	-32.75	2.31	5	6.31
19	2	28	21	18	35	18.69	-32.75	2.44	5	6.31
19	2	28	21	19	35	18.69	-32.75	2.56	5	6.38
19	2	28	21	20	35	18.69	-32.75	2.69	5	6.38
19	2	28	21	21	36	18.69	-32.75	2.81	5	6.38
19	2	28	21	22	36	18.69	-32.69	2.94	5	6.44
19	2	28	21	23	36	18.69	-32.69	3	5	6.44
19	2	28	21	24	36	18.69	-32.63	3.13	5.06	6.44
19	2	28	21	25	37	18.69	-32.63	3.19	5.06	6.5
19	2	28	21	26	37	18.62	-32.63	3.25	5.06	6.5
19	2	28	21	27	37	18.62	-32.63	3.38	5.06	6.56
19	2	28	21	28	37	18.62	-32.56	3.44	5.06	6.56
19	2	28	21	29	38	18.62	-32.56	3.5	5.06	6.56
19	2	28	21	30	38	18.56	-32.56	3.56	5.06	6.56
19	2	28	21	31	38	18.56	-32.56	3.69	5.06	6.56
19	2	28	21	32	38	18.56	-32.56	3.75	5.06	6.63
19	2	28	21	33	39	18.56	-32.56	3.81	5.06	6.63
19	2	28	21	34	39	18.56	-32.5	3.88	5.06	6.63
19	2	28	21	35	39	18.62	-32.56	3.94	5.06	6.63
19	2	28	21	36	40	18.62	-32.5	4	5.06	6.63
19	2	28	21	37	40	18.62	-32.5	4.06	5.06	6.63
19	2	28	21	38	40	18.62	-32.5	4.13	5.13	6.69
19	2	28	21	39	40	18.62	-32.5	4.06	5.13	6.69
19	2	28	21	40	41	18.62	-32.5	4.19	5.13	6.69
19	2	28	21	41	41	18.62	-32.5	4.13	5.13	6.69
19	2	28	21	42	41	18.62	-32.56	4.25	5.13	6.69
19	2	28	21	43	41	18.62	-32.63	4.31	5.13	6.69
19	2	28	21	44	42	18.62	-32.56	4.25	5.13	6.75
19	2	28	21	45	42	18.62	-32.63	4.38	5.13	6.75
19	2	28	21	46	42	18.69	-32.63	4.38	5.13	6.75
19	2	28	21	47	42	18.69	-32.63	4.31	5.13	6.75
19	2	28	21	48	43	18.69	-32.63	4.44	5.13	6.75
19	2	28	21	49	43	18.69	-32.63	4.13	5.13	6.75
19	2	28	21	50	43	18.75	-32.63	4.31	5.13	6.75
19	2	28	21	51	43	18.75	-32.63	4	5.13	6.75
19	2	28	21	52	44	18.75	-32.63	4.06	5.13	6.75
19	2	28	21	53	44	18.69	-32.63	4.25	5.13	6.75
19	2	28	21	54	44	18.69	-32.63	3.94	5.13	6.75
19	2	28	21	55	44	18.69	-32.69	4.19	5.13	6.75
19	2	28	21	56	45	18.69	-32.69	3.75	5.13	6.69
19	2	28	21	57	45	18.69	-32.69	3.69	5.13	6.69
19	2	28	21	58	45	18.69	-32.69	3.75	5.13	6.69
19	2	28	21	59	46	18.69	-32.75	3.25	5.13	6.69
19	2	28	22	0	46	18.69	-32.81	3.63	5.13	6.63
19	2	28	22	1	46	18.69	-32.81	3.19	5.13	6.63
19	2	28	22	2	46	18.69	-32.88	3.5	5.13	6.63
19	2	28	22	3	47	18.69	-32.81	3.63	5.13	6.63
19	2	28	22	4	47	18.69	-32.75	3.38	5.13	6.63
19	2	28	22	5	47	18.62	-32.81	3.69	5.13	6.63
19	2	28	22	6	47	18.62	-32.81	3	5.13	6.56
19	2	28	22	7	48	18.62	-32.88	3.31	5.13	6.56
19	2	28	22	8	48	18.62	-32.94	2.75	5.13	6.56
19	2	28	22	9	48	18.62	-33	2.25	5.13	6.56
19	2	28	22	10	48	18.62	-33	2.81	5.13	6.5
19	2	28	22	11	49	18.62	-33	2.19	5.13	6.5
19	2	28	22	12	49	18.62	-33.06	2.69	5.13	6.5
19	2	28	22	13	49	18.62	-33.06	2.13	5.13	6.44
19	2	28	22	14	49	18.62	-33.13	1.69	5.13	6.44
19	2	28	22	15	50	18.56	-33.13	2.31	5.06	6.44
19	2	28	22	16	50	18.56	-33.13	1.56	5.06	6.44
19	2	28	22	17	50	18.56	-33.13	2.13	5.06	6.38
19	2	28	22	18	50	18.56	-33.19	1.44	5.06	6.38
19	2	28	22	19	51	18.56	-33.25	0.94	5.06	6.38
19	2	28	22	20	51	18.56	-33.19	1.56	5.06	6.38
19	2	28	22	21	51	18.56	-33.25	0.81	5.06	6.31
19	2	28	22	22	52	18.56	-33.25	1.5	5.06	6.31
19	2	28	22	23	52	18.56	-33.25	0.88	5.06	6.31
19	2	28	22	24	52	18.56	-33.19	1.44	5.06	6.25
19	2	28	22	25	52	18.56	-33.25	0.88	5	6.25
19	2	28	22	26	53	18.56	-33.31	0.19	5.06	6.25
19	2	28	22	27	53	18.56	-33.31	0.94	5	6.19
19	2	28	22	28	53	18.56	-33.31	0.19	5	6.19
19	2	28	22	29	53	18.56	-33.25	0.88	5	6.19
19	2	28	22	30	54	18.62	-33.25	0.19	5	6.13
19	2	28	22	31	54	18.62	-33.38	-0.63	5	6.13
19	2	28	22	32	54	18.62	-33.25	0.37	5	6.13
19	2	28	22	33	54	18.62	-33.31	-0.44	5	6.13
19	2	28	22	34	55	18.69	-33.25	0.37	5	6.06
19	2	28	22	35	55	18.69	-33.31	-0.31	5	6.06
19	2	28	22	36	55	18.69	-33.25	0.37	5	6.06
19	2	28	22	37	55	18.75	-33.25	-0.25	4.94	6.06
19	2	28	22	38	56	18.75	-33.31	-1	4.94	6.06
19	2	28	22	39	56	18.75	-33.25	-0.12	4.94	6
19	2	28	22	40	56	18.81	-33.25	-0.94	4.94	6
19	2	28	22	41	56	18.81	-33.19	-0.19	4.94	6
19	2	28	22	42	57	18.81	-33.25	-0.88	4.94	5.94
19	2	28	22	43	57	18.81	-33.31	-1.75	4.94	5.94
19	2	28	22	44	57	18.87	-33.19	-0.63	4.94	5.94
19	2	28	22	46	1	18.87	-33.25	-1.44	4.94	5.94
19	2	28	22	47	1	18.87	-33.19	-0.56	4.94	5.88
19	2	28	22	48	2	18.94	-33.19	-1.12	4.94	5.88
19	2	28	22	49	2	18.94	-33.25	-1.5	4.94	5.88
19	2	28	22	50	2	18.94	-33.25	-0.88	4.94	5.88
19	2	28	22	51	2	18.94	-33.25	-1.56	4.88	5.88
19	2	28	22	52	3	18.94	-33.19	-0.69	4.88	5.88
19	2	28	22	53	3	18.94	-33.25	-1.25	4.88	5.88
19	2	28	22	54	3	18.87	-33.19	-0.5	4.88	5.88
19	2	28	22	55	3	18.87	-33.19	-1	4.88	5.88
19	2	28	22	56	4	18.87	-33.25	-1.81	4.88	5.88
19	2	28	22	57	4	18.81	-33.19	-0.88	4.88	5.81
19	2	28	22	58	4	18.81	-33.13	-1.12	4.88	5.81
19	2	28	22	59	4	18.81	-33.13	-0.44	4.88	5.81
19	2	28	23	0	5	18.75	-33.13	-0.69	4.88	5.81
19	2	28	23	1	5	18.75	-33.19	-1.19	4.88	5.88
19	2	28	23	2	5	18.75	-33.06	-0.31	4.88	5.88
19	2	28	23	3	5	18.69	-33.13	-0.69	4.88	5.88
19	2	28	23	4	6	18.69	-33.06	-0.06	4.88	5.88
19	2	28	23	5	6	18.62	-33.06	-0.37	4.88	5.94
19	2	28	23	6	6	18.62	-33.06	-0.25	4.88	5.94
19	2	28	23	7	7	18.56	-33	0.19	4.88	5.94
19	2	28	23	8	7	18.56	-32.94	0.19	4.88	5.94
19	2	28	23	9	7	18.56	-32.94	0.63	4.88	6
19	2	28	23	10	7	18.56	-32.88	0.63	4.88	6
19	2	28	23	11	8	18.5	-32.88	1	4.88	6
19	2	28	23	12	8	18.5	-32.88	0.5	4.88	6
19	2	28	23	13	8	18.5	-32.88	0.44	4.88	6.06
19	2	28	23	14	8	18.5	-32.88	0.88	4.94	6.06
19	2	28	23	15	9	18.5	-32.81	0.88	4.88	6.06
19	2	28	23	16	9	18.5	-32.81	1.25	4.94	6.06
19	2	28	23	17	9	18.5	-32.75	1.5	4.94	6.13
19	2	28	23	18	9	18.5	-32.75	1.75	4.94	6.13
19	2	28	23	19	10	18.5	-32.75	2	4.94	6.19
19	2	28	23	20	10	18.5	-32.69	2.19	4.94	6.19
19	2	28	23	21	10	18.5	-32.69	2.38	4.94	6.19
19	2	28	23	22	10	18.5	-32.63	2.5	4.94	6.25
19	2	28	23	23	11	18.56	-32.63	2.63	5	6.25
19	2	28	23	24	11	18.56	-32.63	2.81	4.94	6.31
19	2	28	23	25	11	18.56	-32.63	2.88	5	6.31
19	2	28	23	26	11	18.56	-32.63	3	5	6.31
19	2	28	23	27	12	18.56	-32.63	3.13	5	6.38
19	2	28	23	28	12	18.62	-32.63	3.19	5	6.38
19	2	28	23	29	12	18.62	-32.63	3.31	5	6.44
19	2	28	23	30	12	18.62	-32.63	3.38	5	6.44
19	2	28	23	31	13	18.62	-32.63	3.44	5	6.44
19	2	28	23	32	13	18.62	-32.56	3.56	5	6.44
19	2	28	23	33	13	18.69	-32.56	3.63	5	6.5
19	2	28	23	34	14	18.69	-32.56	3.69	5	6.5
19	2	28	23	35	14	18.69	-32.56	3.75	5	6.5
19	2	28	23	36	14	18.69	-32.56	3.81	5	6.5
19	2	28	23	37	14	18.75	-32.56	3.88	5.06	6.56
19	2	28	23	38	15	18.75	-32.56	3.94	5.06	6.56
19	2	28	23	39	15	18.75	-32.56	4	5.06	6.56
19	2	28	23	40	15	18.75	-32.56	4.06	5.06	6.63
19	2	28	23	41	15	18.75	-32.5	4.13	5.06	6.63
19	2	28	23	42	16	18.75	-32.5	4.19	5.06	6.63
19	2	28	23	43	16	18.75	-32.5	4.25	5.06	6.63
19	2	28	23	44	16	18.75	-32.5	4.31	5.06	6.63
19	2	28	23	45	16	18.69	-32.5	4.31	5.06	6.63
19	2	28	23	46	17	18.69	-32.5	4.38	5.06	6.63
19	2	28	23	47	17	18.62	-32.5	4.44	5.06	6.69
19	2	28	23	48	17	18.62	-32.5	4.5	5.06	6.69
19	2	28	23	49	17	18.56	-32.5	4.56	5.06	6.69
19	2	28	23	50	18	18.56	-32.5	4.56	5.06	6.69
19	2	28	23	51	18	18.56	-32.5	4.63	5.13	6.69
19	2	28	23	52	18	18.56	-32.44	4.69	5.06	6.69
19	2	28	23	53	18	18.56	-32.44	4.69	5.06	6.69
19	2	28	23	54	19	18.5	-32.44	4.75	5.06	6.75
19	2	28	23	55	19	18.5	-32.44	4.81	5.06	6.69
19	2	28	23	56	19	18.5	-32.44	4.81	5.06	6.75
19	2	28	23	57	20	18.5	-32.44	4.88	5.06	6.75
19	2	28	23	58	20	18.44	-32.44	4.94	5.13	6.75
19	2	28	23	59	20	18.44	-32.44	4.94	5.13	6.75
19	3	1	0	0	20	18.44	-32.44	5	5.13	6.75
19	3	1	0	1	21	18.44	-32.44	5.06	5.13	6.75
19	3	1	0	2	21	18.37	-32.44	5.06	5.13	6.75
19	3	1	0	3	21	18.37	-32.44	5.13	5.13	6.75
19	3	1	0	4	21	18.37	-32.44	5.13	5.13	6.75
19	3	1	0	5	22	18.37	-32.44	5.19	5.13	6.75
19	3	1	0	6	22	18.37	-32.44	5.19	5.13	6.75
19	3	1	0	7	22	18.37	-32.44	5.25	5.13	6.81
19	3	1	0	8	22	18.37	-32.44	5.25	5.13	6.75
19	3	1	0	9	23	18.37	-32.44	5.31	5.13	6.81
19	3	1	0	10	23	18.37	-32.38	5.31	5.13	6.81
19	3	1	0	11	23	18.37	-32.38	5.31	5.13	6.81
19	3	1	0	12	23	18.37	-32.44	5.38	5.13	6.81
19	3	1	0	13	24	18.37	-32.38	5.44	5.13	6.81
19	3	1	0	14	24	18.44	-32.38	5.44	5.13	6.81
19	3	1	0	15	24	18.44	-32.38	5.44	5.13	6.81
19	3	1	0	16	24	18.44	-32.38	5.5	5.13	6.81
19	3	1	0	17	25	18.44	-32.38	5.5	5.13	6.81
19	3	1	0	18	25	18.5	-32.38	5.5	5.13	6.81
19	3	1	0	19	25	18.5	-32.38	5.56	5.13	6.81
19	3	1	0	20	26	18.5	-32.31	5.5	5.13	6.81
19	3	1	0	21	26	18.5	-32.38	5.56	5.13	6.81
19	3	1	0	22	26	18.5	-32.44	5.56	5.13	6.88
19	3	1	0	23	26	18.5	-32.44	5.63	5.13	6.81
19	3	1	0	24	27	18.5	-32.44	5.63	5.13	6.88
19	3	1	0	25	27	18.5	-32.44	5.56	5.13	6.81
19	3	1	0	26	27	18.5	-32.44	5.63	5.13	6.88
19	3	1	0	27	27	18.5	-32.44	5.63	5.13	6.88
19	3	1	0	28	28	18.5	-32.38	5.56	5.13	6.88
19	3	1	0	29	28	18.44	-32.44	5.63	5.13	6.88
19	3	1	0	30	28	18.44	-32.44	5.31	5.13	6.88
19	3	1	0	31	28	18.44	-32.44	5.44	5.13	6.88
19	3	1	0	32	29	18.44	-32.5	5.13	5.13	6.81
19	3	1	0	33	29	18.44	-32.5	5.19	5.13	6.81
19	3	1	0	34	29	18.44	-32.56	5.31	5.19	6.81
19	3	1	0	35	29	18.37	-32.5	5	5.13	6.81
19	3	1	0	36	30	18.37	-32.56	5.19	5.13	6.81
19	3	1	0	37	30	18.37	-32.56	4.69	5.13	6.81
19	3	1	0	38	30	18.37	-32.63	4.38	5.13	6.75
19	3	1	0	39	30	18.31	-32.63	4.69	5.13	6.75
19	3	1	0	40	31	18.31	-32.63	4.25	5.13	6.75
19	3	1	0	41	31	18.31	-32.69	4.5	5.13	6.75
19	3	1	0	42	31	18.31	-32.69	4.06	5.13	6.69
19	3	1	0	43	31	18.31	-32.75	3.75	5.13	6.69
19	3	1	0	44	32	18.31	-32.81	4.13	5.13	6.69
19	3	1	0	45	32	18.31	-32.81	3.69	5.13	6.63
19	3	1	0	46	32	18.31	-32.81	4	5.13	6.63
19	3	1	0	47	33	18.31	-32.81	3.44	5.13	6.63
19	3	1	0	48	33	18.25	-32.88	2.88	5.13	6.63
19	3	1	0	49	33	18.31	-32.88	3.44	5.13	6.56
19	3	1	0	50	33	18.25	-32.88	2.75	5.13	6.56
19	3	1	0	51	34	18.25	-32.94	3.25	5.13	6.5
19	3	1	0	52	34	18.31	-32.94	2.69	5.13	6.5
19	3	1	0	53	34	18.25	-33	2.38	5.13	6.5
19	3	1	0	54	34	18.25	-33	2.81	5.06	6.5
19	3	1	0	55	35	18.25	-33	2.19	5.06	6.44
19	3	1	0	56	35	18.25	-33	2.75	5.13	6.44
19	3	1	0	57	35	18.25	-33	2.06	5.06	6.44
19	3	1	0	58	35	18.19	-33	2.56	5.06	6.44
19	3	1	0	59	36	18.19	-33	2.13	5.06	6.38
19	3	1	1	0	36	18.19	-33.06	1.5	5.06	6.38
19	3	1	1	1	36	18.19	-33	2.13	5.06	6.38
19	3	1	1	2	36	18.12	-33.06	1.44	5.06	6.31
19	3	1	1	3	37	18.12	-33.06	2	5.06	6.31
19	3	1	1	4	37	18.12	-33.06	1.44	5.06	6.31
19	3	1	1	5	37	18.12	-33.13	0.81	5.06	6.31
19	3	1	1	6	38	18.12	-33.06	1.63	5.06	6.31
19	3	1	1	7	38	18.12	-33.13	0.81	5.06	6.25
19	3	1	1	8	38	18.19	-33.13	1.44	5.06	6.25
19	3	1	1	9	38	18.19	-33.13	0.75	5	6.25
19	3	1	1	10	39	18.19	-33.19	0.44	5	6.25
19	3	1	1	11	39	18.25	-33.13	0.88	5	6.19
19	3	1	1	12	39	18.25	-33.19	0.12	5	6.19
19	3	1	1	13	39	18.25	-33.13	0.88	5	6.19
19	3	1	1	14	40	18.25	-33.13	0.25	5	6.13
19	3	1	1	15	40	18.25	-33.13	0.88	5	6.13
19	3	1	1	16	40	18.25	-33.19	0.31	5	6.13
19	3	1	1	17	40	18.25	-33.19	-0.56	5	6.13
19	3	1	1	18	41	18.19	-33.13	0.37	5	6.06
19	3	1	1	19	41	18.19	-33.19	-0.5	4.94	6.06
19	3	1	1	20	41	18.25	-33.13	0.31	4.94	6.06
19	3	1	1	21	41	18.19	-33.13	-0.44	4.94	6
19	3	1	1	22	42	18.19	-33.19	-1	4.94	6
19	3	1	1	23	42	18.25	-33.13	-0.25	4.94	6
19	3	1	1	24	42	18.25	-33.19	-1.06	4.94	6
19	3	1	1	25	42	18.19	-33.06	-0.19	4.94	5.94
19	3	1	1	26	43	18.19	-33.13	-0.69	4.94	5.94
19	3	1	1	27	43	18.19	-33.06	-0.06	4.94	5.94
19	3	1	1	28	43	18.19	-33.13	-0.56	4.94	5.94
19	3	1	1	29	43	18.19	-33.19	-1.31	4.94	5.94
19	3	1	1	30	44	18.19	-33.06	-0.37	4.88	5.94
19	3	1	1	31	44	18.19	-33.13	-1	4.88	5.94
19	3	1	1	32	44	18.19	-33.06	-0.25	4.88	5.94
19	3	1	1	33	44	18.19	-33.06	-0.75	4.88	5.88
19	3	1	1	34	45	18.19	-33.13	-1.56	4.88	5.88
19	3	1	1	35	45	18.19	-33	-0.56	4.88	5.88
19	3	1	1	36	45	18.12	-33.06	-0.88	4.88	5.88
19	3	1	1	37	46	18.12	-33	-0.19	4.88	5.88
19	3	1	1	38	46	18.12	-33	-0.44	4.88	5.88
19	3	1	1	39	46	18.12	-33	0.19	4.88	5.88
19	3	1	1	40	46	18.12	-32.88	0.5	4.88	5.94
19	3	1	1	41	47	18.12	-32.88	0	4.88	6
19	3	1	1	42	47	18.12	-32.88	0.56	4.88	6
19	3	1	1	43	47	18.12	-32.94	0.19	4.88	6
19	3	1	1	44	47	18.06	-32.94	0.69	4.88	6
19	3	1	1	45	48	18.06	-32.94	0.31	4.88	6
19	3	1	1	46	48	18.06	-32.88	0.25	4.88	6
19	3	1	1	47	48	18.06	-32.88	0.81	4.88	6
19	3	1	1	48	48	18.06	-32.81	0.75	4.88	6.06
19	3	1	1	49	49	18.06	-32.81	1.12	4.88	6.06
19	3	1	1	50	49	18.06	-32.81	1.06	4.94	6.06
19	3	1	1	51	49	18.06	-32.75	0.94	4.94	6.06
19	3	1	1	52	49	18.06	-32.75	1.37	4.88	6.13
19	3	1	1	53	50	18.06	-32.75	1.25	4.94	6.13
19	3	1	1	54	50	18.06	-32.75	1.63	4.94	6.13
19	3	1	1	55	50	18.06	-32.69	1.5	4.94	6.13
19	3	1	1	56	50	18.06	-32.69	1.81	4.94	6.19
19	3	1	1	57	51	18.06	-32.69	2.06	4.94	6.19
19	3	1	1	58	51	18.06	-32.69	2.31	4.94	6.19
19	3	1	1	59	51	18.06	-32.63	2.44	4.94	6.25
19	3	1	2	0	52	18.06	-32.63	2.63	4.94	6.25
19	3	1	2	1	52	18.06	-32.56	2.81	5	6.25
19	3	1	2	2	52	18.06	-32.56	2.94	5	6.31
19	3	1	2	3	52	18.06	-32.56	3	5	6.31
19	3	1	2	4	53	18.06	-32.5	3.13	5	6.38
19	3	1	2	5	53	18.06	-32.5	3.25	5	6.38
19	3	1	2	6	53	18.06	-32.5	3.31	5	6.38
19	3	1	2	7	53	18.06	-32.5	3.44	5	6.44
19	3	1	2	8	54	18.06	-32.5	3.5	5	6.44
19	3	1	2	9	54	18.06	-32.5	3.56	5	6.44
19	3	1	2	10	54	18.06	-32.5	3.69	5.06	6.44
19	3	1	2	11	54	18.06	-32.5	3.75	5	6.5
19	3	1	2	12	55	18	-32.44	3.81	5.06	6.5
19	3	1	2	13	55	18	-32.44	3.88	5.06	6.5
19	3	1	2	14	55	18	-32.44	3.94	5.06	6.5
19	3	1	2	15	59	18	-32.44	4	5.06	6.56
19	3	1	2	16	59	18	-32.44	4.06	5.06	6.56
19	3	1	2	17	59	17.94	-32.44	4.13	5.06	6.56
19	3	1	2	19	0	17.94	-32.38	4.19	5.06	6.56
19	3	1	2	20	0	17.94	-32.38	4.25	5.06	6.63
19	3	1	2	21	0	18	-32.38	4.25	5.06	6.63
19	3	1	2	22	0	18	-32.38	4.31	5.06	6.63
19	3	1	2	23	1	18	-32.38	4.38	5.06	6.63
19	3	1	2	24	1	18	-32.38	4.44	5.06	6.63
19	3	1	2	25	1	18	-32.38	4.5	5.06	6.63
19	3	1	2	26	2	18	-32.38	4.5	5.06	6.69
19	3	1	2	27	2	18	-32.38	4.56	5.06	6.69
19	3	1	2	28	2	18	-32.38	4.63	5.06	6.69
19	3	1	2	29	2	18	-32.31	4.56	5.06	6.69
19	3	1	2	30	3	18	-32.38	4.63	5.06	6.69
19	3	1	2	31	3	18.06	-32.44	4.69	5.13	6.69
19	3	1	2	32	3	18.06	-32.44	4.75	5.06	6.69
19	3	1	2	33	3	18.06	-32.44	4.81	5.13	6.69
19	3	1	2	34	4	18.06	-32.44	4.75	5.13	6.69
19	3	1	2	35	4	18.12	-32.44	4.81	5.13	6.75
19	3	1	2	36	4	18.12	-32.44	4.88	5.13	6.75
19	3	1	2	37	4	18.12	-32.44	4.81	5.13	6.75
19	3	1	2	38	5	18.12	-32.44	4.88	5.13	6.75
19	3	1	2	39	5	18.12	-32.44	4.56	5.13	6.75
19	3	1	2	40	5	18.06	-32.5	4.69	5.13	6.69
19	3	1	2	41	5	18.12	-32.5	4.44	5.13	6.69
19	3	1	2	42	6	18.06	-32.5	4.5	5.13	6.69
19	3	1	2	43	6	18.06	-32.56	4.69	5.13	6.69
19	3	1	2	44	6	18.06	-32.56	4.44	5.13	6.69
19	3	1	2	45	6	18.06	-32.63	4.56	5.13	6.69
19	3	1	2	46	7	18.06	-32.69	4.31	5.13	6.69
19	3	1	2	47	7	18.06	-32.69	4	5.13	6.69
19	3	1	2	48	7	18.06	-32.69	4.31	5.13	6.69
19	3	1	2	49	8	18.06	-32.75	3.88	5.13	6.63
19	3	1	2	50	8	18.06	-32.69	4.13	5.13	6.63
19	3	1	2	51	8	18.06	-32.63	3.75	5.13	6.63
19	3	1	2	52	8	18.06	-32.69	3.44	5.13	6.63
19	3	1	2	53	9	18.06	-32.69	3.81	5.13	6.56
19	3	1	2	54	9	18	-32.75	3.38	5.13	6.56
19	3	1	2	55	9	18	-32.75	3.69	5.13	6.56
19	3	1	2	56	9	18	-32.81	3.38	5.06	6.56
19	3	1	2	57	10	18	-32.88	2.81	5.13	6.56
19	3	1	2	58	10	18	-32.88	3.31	5.13	6.5
19	3	1	2	59	10	18	-32.88	2.69	5.06	6.5
19	3	1	3	0	10	18	-32.88	3.13	5.06	6.5
19	3	1	3	1	11	18	-32.94	2.56	5.06	6.44
19	3	1	3	2	11	18.06	-33	2.06	5.06	6.44
19	3	1	3	3	11	18.06	-32.94	2.69	5.06	6.44
19	3	1	3	4	11	18.06	-33	2.06	5.06	6.44
19	3	1	3	5	12	18.06	-33	2.63	5.06	6.38
19	3	1	3	6	12	18.06	-33	2.06	5.06	6.38
19	3	1	3	7	12	18.06	-33	2.56	5.06	6.38
19	3	1	3	8	12	18.06	-33.06	2	5.06	6.38
19	3	1	3	9	13	18.06	-33.06	1.44	5.06	6.31
19	3	1	3	10	13	18.06	-33.06	2	5.06	6.31
19	3	1	3	11	13	18.12	-33.06	1.37	5.06	6.31
19	3	1	3	12	13	18.12	-33.06	1.88	5.06	6.25
19	3	1	3	13	14	18.12	-33.13	1.37	5.06	6.25
19	3	1	3	14	14	18.12	-33.19	0.75	5	6.25
19	3	1	3	15	14	18.12	-33.13	1.5	5	6.25
19	3	1	3	16	15	18.12	-33.13	0.88	5	6.25
19	3	1	3	17	15	18.06	-33.13	1.44	5	6.19
19	3	1	3	18	15	18.06	-33.19	0.75	5	6.19
19	3	1	3	19	15	18.06	-33.25	0	5	6.19
19	3	1	3	20	16	18.06	-33.19	0.88	5	6.13
19	3	1	3	21	16	18.06	-33.25	0.06	5	6.13
19	3	1	3	22	16	18.06	-33.19	0.81	5	6.13
19	3	1	3	23	16	18.06	-33.25	0.12	5	6.13
19	3	1	3	24	17	18.06	-33.19	0.75	5	6.13
19	3	1	3	25	17	18.06	-33.19	0.37	4.94	6.06
19	3	1	3	26	17	18.06	-33.25	-0.31	4.94	6.06
19	3	1	3	27	17	18.06	-33.19	0.5	4.94	6.06
19	3	1	3	28	18	18.06	-33.25	0	4.94	6.06
19	3	1	3	29	18	18.06	-33.19	0.63	4.94	6.06
19	3	1	3	30	18	18.06	-33.19	0.19	4.94	6.06
19	3	1	3	31	18	18	-33.31	-0.63	4.94	6.06
19	3	1	3	32	19	18	-33.19	0.31	4.94	6
19	3	1	3	33	19	18	-33.19	-0.31	4.94	6
19	3	1	3	34	19	18	-33.13	0.37	4.94	6
19	3	1	3	35	19	18	-33.19	-0.06	4.94	6
19	3	1	3	36	20	18	-33.13	-0.06	4.94	6
19	3	1	3	37	20	18.06	-33.13	0.44	4.94	6
19	3	1	3	38	20	18.06	-33.06	0.25	4.94	6
19	3	1	3	39	21	18.06	-33	0.81	4.94	6
19	3	1	3	40	21	18	-33	0.63	4.94	6
19	3	1	3	41	21	18.06	-32.94	1.06	4.94	6
19	3	1	3	42	21	18	-32.88	1.44	4.94	6.06
19	3	1	3	43	22	18.06	-32.88	1	4.94	6.06
19	3	1	3	44	22	18	-32.88	1.5	4.94	6.06
19	3	1	3	45	22	18	-32.88	1.19	4.94	6.06
19	3	1	3	46	22	18	-32.88	1.56	4.94	6.06
19	3	1	3	47	23	18	-32.81	1.75	4.94	6.13
19	3	1	3	48	23	18	-32.75	1.88	4.94	6.13
19	3	1	3	49	23	18	-32.81	2.19	4.94	6.19
19	3	1	3	50	23	18	-32.75	2.25	5	6.19
19	3	1	3	51	24	18	-32.75	2.5	5	6.19
19	3	1	3	52	24	18	-32.75	2.56	5	6.25
19	3	1	3	53	24	17.94	-32.75	2.63	5	6.25
19	3	1	3	54	24	17.94	-32.75	2.81	5	6.25
19	3	1	3	55	25	17.94	-32.75	2.88	5	6.31
19	3	1	3	56	25	17.94	-32.75	3.06	5	6.38
19	3	1	3	57	25	17.94	-32.75	3.19	5	6.38
19	3	1	3	58	25	17.94	-32.75	3.31	5	6.38
19	3	1	3	59	26	17.94	-32.69	3.44	5	6.38
19	3	1	4	0	26	17.94	-32.69	3.5	5	6.44
19	3	1	4	1	26	17.94	-32.63	3.56	5	6.44
19	3	1	4	2	26	17.94	-32.63	3.69	5	6.44
19	3	1	4	3	27	18	-32.56	3.75	5.06	6.5
19	3	1	4	4	27	18	-32.5	3.81	5.06	6.5
19	3	1	4	5	27	18	-32.5	3.88	5.06	6.5
19	3	1	4	6	28	18	-32.5	3.94	5.06	6.5
19	3	1	4	7	28	18	-32.44	4	5.06	6.56
19	3	1	4	8	28	18.06	-32.44	4.06	5.06	6.56
19	3	1	4	9	28	18.06	-32.44	4.13	5.06	6.56
19	3	1	4	10	29	18.06	-32.38	4.19	5.06	6.56
19	3	1	4	11	29	18.06	-32.38	4.25	5.06	6.56
19	3	1	4	12	29	18.06	-32.38	4.31	5.06	6.56
19	3	1	4	13	29	18.06	-32.38	4.38	5.06	6.63
19	3	1	4	14	30	18.06	-32.38	4.38	5.06	6.63
19	3	1	4	15	30	18.06	-32.38	4.44	5.06	6.63
19	3	1	4	16	30	18	-32.38	4.5	5.06	6.63
19	3	1	4	17	30	18	-32.31	4.56	5.06	6.63
19	3	1	4	18	31	18	-32.31	4.56	5.06	6.63
19	3	1	4	19	31	18	-32.31	4.63	5.06	6.63
19	3	1	4	20	31	17.94	-32.31	4.56	5.06	6.69
19	3	1	4	21	31	17.94	-32.38	4.69	5.13	6.69
19	3	1	4	22	32	17.94	-32.44	4.69	5.06	6.69
19	3	1	4	23	32	17.94	-32.44	4.75	5.13	6.69
19	3	1	4	24	32	17.87	-32.44	4.81	5.13	6.69
19	3	1	4	25	32	17.94	-32.38	4.75	5.06	6.69
19	3	1	4	26	33	17.87	-32.44	4.81	5.13	6.69
19	3	1	4	27	33	17.87	-32.38	4.81	5.13	6.69
19	3	1	4	28	33	17.87	-32.38	4.75	5.13	6.75
19	3	1	4	29	33	17.87	-32.38	4.88	5.13	6.75
19	3	1	4	30	34	17.87	-32.44	4.56	5.13	6.69
19	3	1	4	31	34	17.87	-32.5	4.69	5.13	6.69
19	3	1	4	32	34	17.87	-32.56	4.44	5.13	6.69
19	3	1	4	33	35	17.87	-32.56	4.5	5.13	6.69
19	3	1	4	34	35	17.87	-32.56	4.63	5.13	6.69
19	3	1	4	35	35	17.87	-32.56	4.38	5.13	6.69
19	3	1	4	36	35	17.81	-32.63	4.5	5.13	6.69
19	3	1	4	37	36	17.81	-32.56	4.06	5.13	6.69
19	3	1	4	38	36	17.81	-32.56	3.75	5.13	6.63
19	3	1	4	39	36	17.81	-32.63	4.06	5.13	6.63
19	3	1	4	40	36	17.81	-32.56	3.56	5.13	6.63
19	3	1	4	41	37	17.81	-32.63	3.88	5.06	6.63
19	3	1	4	42	37	17.87	-32.69	3.5	5.13	6.56
19	3	1	4	43	37	17.87	-32.69	3.25	5.13	6.56
19	3	1	4	44	37	17.87	-32.75	3.63	5.13	6.56
19	3	1	4	45	38	17.81	-32.75	3.25	5.13	6.56
19	3	1	4	46	38	17.87	-32.75	3.56	5.13	6.56
19	3	1	4	47	38	17.81	-32.81	2.94	5.06	6.5
19	3	1	4	48	38	17.81	-32.88	2.44	5.06	6.5
19	3	1	4	49	39	17.81	-32.88	3	5.06	6.5
19	3	1	4	50	39	17.81	-32.94	2.38	5.06	6.44
19	3	1	4	51	39	17.81	-32.94	2.88	5.06	6.44
19	3	1	4	52	39	17.81	-33	2.31	5.06	6.44
19	3	1	4	53	40	17.81	-33	2	5.06	6.38
19	3	1	4	54	40	17.81	-33	2.44	5.06	6.38
19	3	1	4	55	40	17.81	-33	1.88	5.06	6.38
19	3	1	4	56	41	17.81	-33	2.38	5.06	6.38
19	3	1	4	57	41	17.81	-33.06	1.69	5.06	6.31
19	3	1	4	58	41	17.75	-33.06	2.19	5.06	6.31
19	3	1	4	59	41	17.75	-33.06	1.81	5	6.31
19	3	1	5	0	42	17.75	-33.13	1.19	5	6.31
19	3	1	5	1	42	17.75	-33.13	1.81	5	6.25
19	3	1	5	2	42	17.75	-33.19	1.19	5	6.25
19	3	1	5	3	42	17.75	-33.13	1.75	5	6.25
19	3	1	5	4	43	17.75	-33.19	1.19	5	6.25
19	3	1	5	5	43	17.75	-33.25	0.56	5	6.25
19	3	1	5	6	43	17.75	-33.19	1.31	5	6.19
19	3	1	5	7	43	17.75	-33.19	0.5	5	6.19
19	3	1	5	8	44	17.75	-33.19	1.19	5	6.19
19	3	1	5	9	44	17.69	-33.25	0.44	5	6.13
19	3	1	5	10	44	17.69	-33.25	0.19	5	6.13
19	3	1	5	11	44	17.75	-33.25	0.63	4.94	6.13
19	3	1	5	12	45	17.75	-33.31	-0.19	4.94	6.13
19	3	1	5	13	45	17.75	-33.25	0.63	4.94	6.06
19	3	1	5	14	45	17.75	-33.25	-0.06	4.94	6.06
19	3	1	5	15	45	17.81	-33.19	0.63	4.94	6.06
19	3	1	5	16	46	17.81	-33.25	0.06	4.94	6
19	3	1	5	17	46	17.87	-33.31	-0.56	4.94	6
19	3	1	5	18	46	17.87	-33.25	0.25	4.94	6
19	3	1	5	19	46	17.87	-33.25	-0.25	4.94	6
19	3	1	5	20	47	17.94	-33.19	0.44	4.94	6
19	3	1	5	21	47	17.94	-33.19	-0.06	4.94	6
19	3	1	5	22	47	18	-33.31	-0.81	4.94	6
19	3	1	5	23	48	18	-33.19	0.12	4.88	5.94
19	3	1	5	24	48	18.06	-33.25	-0.44	4.88	5.94
19	3	1	5	25	48	18.06	-33.19	0.31	4.88	5.94
19	3	1	5	26	48	18.06	-33.19	0.12	4.88	5.94
19	3	1	5	27	49	18.12	-33.13	0.56	4.88	6
19	3	1	5	28	49	18.12	-33.13	0.12	4.88	5.94
19	3	1	5	29	49	18.12	-33.13	-0.06	4.88	5.94
19	3	1	5	30	49	18.12	-33.06	0.56	4.88	5.94
19	3	1	5	31	50	18.19	-33.06	0.37	4.88	6
19	3	1	5	32	50	18.19	-33	0.88	4.88	6
19	3	1	5	33	50	18.19	-33	0.69	4.88	6
19	3	1	5	34	50	18.19	-33	0.44	4.88	6
19	3	1	5	35	51	18.25	-33	1	4.88	6
19	3	1	5	36	51	18.25	-32.88	1.19	4.88	6
19	3	1	5	37	51	18.25	-32.88	1.56	4.88	6.06
19	3	1	5	38	51	18.25	-32.75	1.75	4.94	6.06
19	3	1	5	39	52	18.25	-32.69	1.88	4.94	6.13
19	3	1	5	40	52	18.25	-32.75	2.13	4.94	6.13
19	3	1	5	41	52	18.19	-32.69	2.25	4.94	6.19
19	3	1	5	42	52	18.19	-32.75	2.44	4.94	6.19
19	3	1	5	43	53	18.19	-32.69	2.56	4.94	6.25
19	3	1	5	44	53	18.12	-32.69	2.56	4.94	6.25
19	3	1	5	45	57	18.12	-32.69	2.81	4.94	6.31
19	3	1	5	46	57	18.12	-32.75	2.94	4.94	6.31
19	3	1	5	47	57	18.06	-32.75	3.06	5	6.31
19	3	1	5	48	58	18	-32.69	3.19	4.94	6.31
19	3	1	5	49	58	18	-32.69	3.31	5	6.38
19	3	1	5	50	58	18	-32.63	3.38	5	6.38
19	3	1	5	51	58	17.94	-32.56	3.5	5	6.38
19	3	1	5	52	59	17.94	-32.56	3.56	5	6.44
19	3	1	5	53	59	17.87	-32.5	3.63	5	6.44
19	3	1	5	54	59	17.87	-32.5	3.69	5	6.44
19	3	1	5	55	59	17.87	-32.44	3.81	5	6.5
19	3	1	5	57	0	17.87	-32.44	3.88	5	6.5
19	3	1	5	58	0	17.81	-32.44	3.94	5	6.5
19	3	1	5	59	0	17.87	-32.38	4	5	6.5
19	3	1	6	0	0	17.87	-32.38	4.06	5.06	6.5
19	3	1	6	1	1	17.87	-32.38	4.13	5	6.56
19	3	1	6	2	1	17.87	-32.38	4.13	5	6.56
19	3	1	6	3	1	17.87	-32.31	4.19	5	6.56
19	3	1	6	4	1	17.87	-32.31	4.25	5.06	6.56
19	3	1	6	5	2	17.87	-32.31	4.31	5.06	6.56
19	3	1	6	6	2	17.94	-32.31	4.38	5.06	6.63
19	3	1	6	7	2	17.94	-32.25	4.38	5.06	6.63
19	3	1	6	8	2	17.94	-32.25	4.44	5.06	6.63
19	3	1	6	9	3	17.94	-32.25	4.5	5.06	6.63
19	3	1	6	10	3	17.94	-32.25	4.56	5.06	6.63
19	3	1	6	11	3	18	-32.25	4.56	5.06	6.63
19	3	1	6	12	4	18	-32.25	4.63	5.06	6.63
19	3	1	6	13	4	18	-32.25	4.69	5.06	6.63
19	3	1	6	14	4	18	-32.25	4.75	5.06	6.69
19	3	1	6	15	4	18	-32.25	4.75	5.06	6.69
19	3	1	6	16	5	18.06	-32.25	4.81	5.06	6.69
19	3	1	6	17	5	18.06	-32.25	4.81	5.06	6.69
19	3	1	6	18	5	18	-32.19	4.88	5.06	6.69
19	3	1	6	19	5	18.06	-32.19	4.94	5.06	6.69
19	3	1	6	20	6	18	-32.19	4.94	5.06	6.69
19	3	1	6	21	6	18	-32.19	5	5.06	6.69
19	3	1	6	22	6	17.94	-32.19	5	5.13	6.69
19	3	1	6	23	6	17.94	-32.19	5.06	5.06	6.69
19	3	1	6	24	7	17.94	-32.19	5.06	5.06	6.69
19	3	1	6	25	7	17.87	-32.19	5.13	5.06	6.75
19	3	1	6	26	7	17.87	-32.19	5.13	5.06	6.75
19	3	1	6	27	7	17.87	-32.19	5.19	5.13	6.75
19	3	1	6	28	8	17.81	-32.19	5.19	5.13	6.75
19	3	1	6	29	8	17.81	-32.19	5.25	5.13	6.75
19	3	1	6	30	8	17.81	-32.19	5.25	5.13	6.75
19	3	1	6	31	8	17.81	-32.19	5.31	5.13	6.75
19	3	1	6	32	9	17.81	-32.13	5.31	5.13	6.75
19	3	1	6	33	9	17.81	-32.13	5.38	5.13	6.75
19	3	1	6	34	9	17.81	-32.13	5.38	5.13	6.75
19	3	1	6	35	9	17.87	-32.13	5.44	5.13	6.75
19	3	1	6	36	10	17.87	-32.13	5.44	5.13	6.75
19	3	1	6	37	10	17.87	-32.13	5.38	5.13	6.75
19	3	1	6	38	10	17.87	-32.19	5.38	5.13	6.75
19	3	1	6	39	11	17.81	-32.19	5.44	5.13	6.75
19	3	1	6	40	11	17.81	-32.25	5.5	5.13	6.75
19	3	1	6	41	11	17.81	-32.25	5.5	5.13	6.81
19	3	1	6	42	11	17.81	-32.19	5.44	5.13	6.81
19	3	1	6	43	12	17.81	-32.25	5.5	5.13	6.81
19	3	1	6	44	12	17.81	-32.25	5.56	5.13	6.81
19	3	1	6	45	12	17.81	-32.25	5.56	5.13	6.81
19	3	1	6	46	12	17.81	-32.25	5.63	5.13	6.81
19	3	1	6	47	13	17.81	-32.19	5.25	5.13	6.81
19	3	1	6	48	13	17.81	-32.25	5.38	5.13	6.75
19	3	1	6	49	13	17.75	-32.25	5.13	5.13	6.75
19	3	1	6	50	13	17.75	-32.31	5.13	5.13	6.75
19	3	1	6	51	14	17.75	-32.38	5.25	5.13	6.75
19	3	1	6	52	14	17.75	-32.44	4.94	5.13	6.75
19	3	1	6	53	14	17.75	-32.44	5.13	5.13	6.75
19	3	1	6	54	14	17.75	-32.44	4.88	5.13	6.75
19	3	1	6	55	15	17.75	-32.44	4.94	5.13	6.75
19	3	1	6	56	15	17.75	-32.5	5.13	5.13	6.69
19	3	1	6	57	15	17.75	-32.44	4.56	5.13	6.69
19	3	1	6	58	15	17.75	-32.44	4.81	5.13	6.69
19	3	1	6	59	16	17.75	-32.5	4.38	5.13	6.69
19	3	1	7	0	16	17.75	-32.56	4.13	5.13	6.69
19	3	1	7	1	16	17.75	-32.56	4.44	5.13	6.69
19	3	1	7	2	17	17.75	-32.63	3.94	5.13	6.63
19	3	1	7	3	17	17.75	-32.63	4.25	5.13	6.63
19	3	1	7	4	17	17.75	-32.69	3.81	5.13	6.63
19	3	1	7	5	17	17.81	-32.69	3.5	5.13	6.56
19	3	1	7	6	18	17.81	-32.69	3.94	5.13	6.56
19	3	1	7	7	18	17.81	-32.75	3.25	5.13	6.56
19	3	1	7	8	18	17.81	-32.75	3.69	5.13	6.5
19	3	1	7	9	18	17.87	-32.81	3.06	5.13	85
19	3	1	7	10	19	17.87	-32.88	2.56	5.06	85
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	3	1	7	14	26	17.94	-32.69	3.94	5.06	6.5
19	3	1	7	15	27	18	-32.63	4.13	5.06	6.5
19	3	1	7	16	27	18	-32.5	4.13	5.06	6.5
19	3	1	7	17	27	18	-32.56	4.31	5.06	6.56
19	3	1	7	18	27	18	-32.5	4.31	5.06	6.56
19	3	1	7	19	28	18	-32.56	4.44	5.06	6.56
19	3	1	7	20	28	18	-32.5	4.44	5.13	6.63
19	3	1	7	21	28	18	-32.56	4.56	5.06	6.63
19	3	1	7	22	28	18	-32.56	4.69	5.13	6.63
19	3	1	7	23	29	18	-32.5	4.63	5.13	6.63
19	3	1	7	24	29	18	-32.5	4.75	5.13	6.63
19	3	1	7	25	29	18	-32.5	4.81	5.13	6.63
19	3	1	7	26	29	18	-32.5	4.75	5.13	6.69
19	3	1	7	27	30	18	-32.5	4.88	5.13	6.69
19	3	1	7	28	30	18	-32.44	4.81	5.13	6.69
19	3	1	7	29	30	18	-32.5	4.94	5.13	6.69
19	3	1	7	30	30	18	-32.5	4.94	5.13	6.69
19	3	1	7	31	31	18.06	-32.5	5	5.13	6.69
19	3	1	7	32	31	18.06	-32.5	5.06	5.13	6.75
19	3	1	7	33	31	18.06	-32.5	5.13	5.13	6.75
19	3	1	7	34	31	18.06	-32.44	5.19	5.13	6.75
19	3	1	7	35	32	18.06	-32.38	5.19	5.13	6.75
19	3	1	7	36	32	18.06	-32.38	5.25	5.13	6.75
19	3	1	7	37	32	18.06	-32.31	5.31	5.13	6.75
19	3	1	7	38	33	18.12	-32.31	5.31	5.13	6.75
19	3	1	7	39	33	18.12	-32.25	5.31	5.13	6.75
19	3	1	7	40	33	18.12	-32.25	5.38	5.13	6.75
19	3	1	7	41	33	18.12	-32.19	5.38	5.13	6.81
19	3	1	7	42	34	18.12	-32.19	5.44	5.13	6.81
19	3	1	7	43	34	18.19	-32.19	5.44	5.19	6.81
19	3	1	7	44	34	18.19	-32.13	5.5	5.13	6.81
19	3	1	7	45	34	18.19	-32.13	5.5	5.19	6.81
19	3	1	7	46	35	18.19	-32.13	5.5	5.19	6.81
19	3	1	7	47	35	18.25	-32.13	5.56	5.19	6.81
19	3	1	7	48	35	18.25	-32.13	5.56	5.19	6.81
19	3	1	7	49	35	18.25	-32.13	5.63	5.19	6.81
19	3	1	7	50	36	18.25	-32.13	5.63	5.19	6.81
19	3	1	7	51	36	18.25	-32.06	5.63	5.19	6.88
19	3	1	7	52	36	18.31	-32.06	5.63	5.19	6.81
19	3	1	7	53	36	18.31	-32.06	5.56	5.19	6.88
19	3	1	7	54	37	18.25	-32.13	5.63	5.19	6.88
19	3	1	7	55	37	18.25	-32.13	5.69	5.19	6.88
19	3	1	7	56	37	18.31	-32.19	5.69	5.19	6.88
19	3	1	7	57	37	18.25	-32.13	5.75	5.19	6.88
19	3	1	7	58	38	18.31	-32.13	5.63	5.19	6.88
19	3	1	7	59	38	18.31	-32.13	5.69	5.19	6.88
19	3	1	8	0	38	18.31	-32.19	5.75	5.19	6.88
19	3	1	8	1	39	18.31	-32.13	5.63	5.19	6.88
19	3	1	8	2	39	18.25	-32.13	5.75	5.19	6.88
19	3	1	8	3	39	18.25	-32.13	5.38	5.19	6.88
19	3	1	8	4	39	18.25	-32.25	5.5	5.19	6.88
19	3	1	8	5	40	18.25	-32.31	5.25	5.19	6.88
19	3	1	8	6	40	18.25	-32.31	5.25	5.19	6.88
19	3	1	8	7	40	18.25	-32.31	5.38	5.19	6.88
19	3	1	8	8	40	18.25	-32.31	5.13	5.19	6.88
19	3	1	8	9	41	18.25	-32.38	5.25	5.19	6.81
19	3	1	8	10	41	18.25	-32.44	5	5.19	6.81
19	3	1	8	11	41	18.25	-32.5	4.69	5.19	85
19	3	1	8	12	41	18.25	-32.5	4.94	5.19	6.81
19	3	1	8	13	42	18.25	-32.5	4.5	5.19	6.81
19	3	1	8	14	42	18.25	-32.5	4.75	5.19	6.75
19	3	1	8	15	42	18.25	-32.5	4.31	5.19	6.75
19	3	1	8	16	42	18.25	-32.5	4	5.19	6.75
19	3	1	8	17	43	18.25	-32.5	4.38	5.19	6.69
19	3	1	8	18	43	18.31	-32.56	3.94	5.19	6.69
19	3	1	8	19	43	18.25	-32.56	4.25	5.19	6.69
19	3	1	8	20	43	18.31	-32.56	3.88	5.19	6.69
19	3	1	8	21	44	18.31	-32.56	3.31	5.19	6.63
19	3	1	8	22	44	18.31	-32.56	3.81	5.19	6.63
19	3	1	8	23	44	18.31	-32.63	3.13	5.19	6.63
19	3	1	8	24	44	18.31	-32.63	3.56	5.19	6.56
19	3	1	8	25	45	18.31	-32.69	3.06	5.19	6.56
19	3	1	8	26	45	18.31	-32.75	2.5	5.19	6.56
19	3	1	8	27	45	18.31	-32.75	3.13	5.13	6.56
19	3	1	8	28	46	18.25	-32.81	2.5	5.13	6.56
19	3	1	8	29	46	18.25	-32.81	3	5.13	6.5
19	3	1	8	30	46	18.25	-32.88	2.5	5.13	6.5
19	3	1	8	31	46	18.25	-32.81	2.94	5.13	6.5
19	3	1	8	32	47	18.25	-32.94	2.38	5.13	6.5
19	3	1	8	33	47	18.25	-32.94	1.81	5.13	6.44
19	3	1	8	34	47	18.25	-32.94	2.38	5.13	6.44
19	3	1	8	35	47	18.25	-33	1.75	5.13	6.44
19	3	1	8	36	48	18.25	-32.94	2.25	5.13	6.38
19	3	1	8	37	48	18.25	-33	1.69	5.13	6.38
19	3	1	8	38	48	18.25	-33.06	1.12	5.13	6.38
19	3	1	8	39	48	18.19	-33	1.88	5.13	6.38
19	3	1	8	40	49	18.25	-33.06	1.19	5.13	6.38
19	3	1	8	41	49	18.19	-33	1.81	5.06	6.31
19	3	1	8	42	49	18.25	-33	1.25	5.06	6.31
19	3	1	8	43	49	18.25	-33.06	0.44	5.06	6.31
19	3	1	8	44	50	18.25	-33	1.31	5.06	6.31
19	3	1	8	45	50	18.25	-33.06	0.5	5.06	6.25
19	3	1	8	46	50	18.25	-33	1.19	5.06	6.25
19	3	1	8	47	50	18.31	-33.06	0.5	5.06	6.25
19	3	1	8	48	51	18.25	-33	1.19	5.06	6.25
19	3	1	8	49	51	18.31	-33	0.88	5.06	6.19
19	3	1	8	50	51	18.25	-33.06	0.19	5.06	6.19
19	3	1	8	51	52	18.25	-33	0.94	5.06	6.19
19	3	1	8	52	52	18.25	-33	0.31	5.06	6.19
19	3	1	8	53	52	18.25	-32.94	0.94	5.06	6.19
19	3	1	8	54	52	18.25	-33	0.19	5.06	6.13
19	3	1	8	55	53	18.31	-33.06	-0.75	5	6.13
19	3	1	8	56	53	18.31	-33	0.25	5	6.13
19	3	1	8	57	53	18.31	-33.06	-0.63	5	6.13
19	3	1	8	58	53	18.31	-33	0.19	5	6.06
19	3	1	8	59	54	18.25	-33	-0.5	5	6.06
19	3	1	9	0	54	18.25	-32.94	0.19	5	6.06
19	3	1	9	1	54	18.25	-33	-0.81	5	6
19	3	1	9	2	54	18.25	-33.06	-1.75	5	6
19	3	1	9	3	55	18.25	-33	-0.69	4.94	5.94
19	3	1	9	4	55	18.25	-33.06	-1.5	4.94	5.94
19	3	1	9	5	55	18.25	-32.94	-0.63	4.94	5.94
19	3	1	9	6	55	18.25	-33	-1.31	4.94	5.88
19	3	1	9	7	56	18.25	-33	-1.25	4.94	5.88
19	3	1	9	8	56	18.25	-33	-1.37	4.94	5.88
19	3	1	9	9	56	18.25	-33.06	-2.13	4.94	5.88
19	3	1	9	11	0	18.25	-32.94	-1.06	4.94	5.81
19	3	1	9	12	1	18.25	-33	-1.63	4.94	5.81
19	3	1	9	13	1	18.25	-32.94	-0.81	4.94	5.81
19	3	1	9	14	1	18.25	-32.94	-1.31	4.94	5.81
19	3	1	9	15	2	18.31	-33	-2.13	4.88	5.81
19	3	1	9	16	2	18.31	-32.88	-1.12	4.88	5.81
19	3	1	9	17	2	18.31	-32.94	-1.81	4.88	5.81
19	3	1	9	18	2	18.37	-32.88	-1	4.88	5.75
19	3	1	9	19	3	18.37	-32.94	-1.56	4.88	5.75
19	3	1	9	20	3	18.37	-32.94	-1.69	4.88	5.75
19	3	1	9	21	3	18.37	-32.88	-1.69	4.88	5.75
19	3	1	9	22	3	18.37	-32.94	-2.06	4.88	5.75
19	3	1	9	23	4	18.37	-32.88	-1.19	4.88	5.75
19	3	1	9	24	4	18.37	-32.94	-1.5	4.88	5.75
19	3	1	9	25	4	18.37	-32.81	-0.81	4.88	5.75
19	3	1	9	26	4	18.37	-32.88	-1.12	4.88	5.75
19	3	1	9	27	5	18.37	-32.94	-1.63	4.88	5.75
19	3	1	9	28	5	18.37	-32.88	-0.81	4.88	5.75
19	3	1	9	29	5	18.37	-32.94	-1.19	4.88	5.81
19	3	1	9	30	5	18.37	-32.88	-0.56	4.88	5.81
19	3	1	9	31	6	18.37	-32.94	-0.94	4.88	5.81
19	3	1	9	32	6	18.37	-32.88	-1.12	4.88	5.81
19	3	1	9	33	6	18.37	-32.81	-0.37	4.88	5.81
19	3	1	9	34	6	18.37	-32.81	-0.5	4.88	5.88
19	3	1	9	35	7	18.37	-32.75	0.06	4.88	5.88
19	3	1	9	36	7	18.37	-32.75	-0.06	4.88	5.88
19	3	1	9	37	7	18.37	-32.75	-0.37	4.88	5.88
19	3	1	9	38	7	18.37	-32.69	0.25	4.88	5.94
19	3	1	9	39	8	18.37	-32.69	0	4.88	5.94
19	3	1	9	40	8	18.37	-32.69	0.5	4.88	5.94
19	3	1	9	41	8	18.37	-32.69	0.31	4.88	5.94
19	3	1	9	42	9	18.44	-32.63	0.75	4.88	6
19	3	1	9	43	9	18.44	-32.69	0.44	4.88	6
19	3	1	9	44	9	18.44	-32.56	0.69	4.88	6
19	3	1	9	45	9	18.44	-32.56	1.06	4.88	6
19	3	1	9	46	10	18.5	-32.5	1.25	4.94	6.06
19	3	1	9	47	10	18.5	-32.5	1.56	4.94	6.06
19	3	1	9	48	10	18.5	-32.44	1.69	4.94	6.13
19	3	1	9	49	10	18.5	-32.44	1.81	4.94	6.13
19	3	1	9	50	11	18.56	-32.44	2	4.94	6.19
19	3	1	9	51	11	18.56	-32.44	2.13	4.94	6.19
19	3	1	9	52	11	18.56	-32.5	2.31	4.94	6.25
19	3	1	9	53	11	18.62	-32.44	2.5	4.94	6.25
19	3	1	9	54	12	18.62	-32.44	2.63	4.94	6.25
19	3	1	9	55	12	18.62	-32.38	2.75	5	6.31
19	3	1	9	56	12	18.69	-32.31	2.88	5	6.31
19	3	1	9	57	12	18.69	-32.31	3	5	6.38
19	3	1	9	58	13	18.69	-32.25	3.06	5	6.38
19	3	1	9	59	13	18.75	-32.19	3.19	5	6.38
19	3	1	10	0	13	18.75	-32.19	3.25	5	6.38
19	3	1	10	1	14	18.75	-32.13	3.31	5	6.44
19	3	1	10	2	14	18.75	-32.13	3.44	5	6.44
19	3	1	10	3	14	18.81	-32.13	3.5	5	6.44
19	3	1	10	4	14	18.81	-32.06	3.56	5	6.5
19	3	1	10	5	15	18.81	-32.06	3.63	5	6.5
19	3	1	10	6	15	18.81	-32.06	3.69	5.06	6.5
19	3	1	10	7	15	18.87	-32.06	3.75	5.06	6.5
19	3	1	10	8	15	18.87	-32	3.81	5.06	6.56
19	3	1	10	9	16	18.87	-32	3.88	5.06	6.56
19	3	1	10	10	16	18.87	-32	4	5.06	6.56
19	3	1	10	11	16	18.94	-32	4	5.06	6.56
19	3	1	10	12	16	18.94	-31.94	4.06	5.06	6.56
19	3	1	10	13	17	18.94	-31.94	4.13	5.06	6.56
19	3	1	10	14	17	18.94	-31.94	4.19	5.06	6.63
19	3	1	10	15	17	18.94	-31.94	4.25	5.06	6.63
19	3	1	10	16	17	18.94	-31.94	4.31	5.06	6.63
19	3	1	10	17	18	18.94	-31.94	4.38	5.06	6.63
19	3	1	10	18	18	18.94	-31.94	4.38	5.06	6.63
19	3	1	10	19	18	18.94	-31.94	4.44	5.06	6.69
19	3	1	10	20	18	18.94	-31.87	4.5	5.06	6.69
19	3	1	10	21	19	18.87	-31.87	4.56	5.06	6.69
19	3	1	10	22	19	18.87	-31.87	4.56	5.13	6.69
19	3	1	10	23	19	18.81	-31.87	4.63	5.06	6.69
19	3	1	10	24	19	18.81	-31.87	4.69	5.06	6.69
19	3	1	10	25	20	18.81	-31.87	4.75	5.06	6.69
19	3	1	10	26	20	18.81	-31.87	4.75	5.13	6.69
19	3	1	10	27	20	18.75	-31.87	4.81	5.06	6.69
19	3	1	10	28	21	18.75	-31.87	4.88	5.13	6.75
19	3	1	10	29	21	18.75	-31.81	4.88	5.13	6.75
19	3	1	10	30	21	18.69	-31.81	4.94	5.13	6.75
19	3	1	10	31	21	18.69	-31.81	4.94	5.13	6.75
19	3	1	10	32	22	18.69	-31.81	5	5.13	6.75
19	3	1	10	33	22	18.62	-31.81	5.06	5.13	6.75
19	3	1	10	34	22	18.62	-31.81	5.06	5.13	6.75
19	3	1	10	35	22	18.62	-31.81	5.13	5.13	6.75
19	3	1	10	36	23	18.62	-31.81	5.13	5.13	6.75
19	3	1	10	37	23	18.62	-31.81	5.19	5.13	6.81
19	3	1	10	38	23	18.56	-31.81	5.19	5.13	6.81
19	3	1	10	39	23	18.62	-31.81	5.25	5.13	6.81
19	3	1	10	40	24	18.62	-31.81	5.25	5.13	6.81
19	3	1	10	41	24	18.62	-31.81	5.31	5.13	6.81
19	3	1	10	42	24	18.62	-31.81	5.31	5.13	6.81
19	3	1	10	43	24	18.62	-31.81	5.38	5.13	6.81
19	3	1	10	44	25	18.62	-31.75	5.38	5.13	6.81
19	3	1	10	45	25	18.69	-31.81	5.38	5.13	6.81
19	3	1	10	46	25	18.69	-31.81	5.31	5.13	6.81
19	3	1	10	47	25	18.62	-31.81	5.38	5.13	6.81
19	3	1	10	48	26	18.62	-31.87	5.44	5.13	6.81
19	3	1	10	49	26	18.62	-31.87	5.44	5.13	6.81
19	3	1	10	50	26	18.62	-31.87	5.5	5.13	6.81
19	3	1	10	51	27	18.62	-31.87	5.44	5.13	6.81
19	3	1	10	52	27	18.62	-31.87	5.5	5.13	6.81
19	3	1	10	53	27	18.62	-31.87	5.5	5.13	6.81
19	3	1	10	54	27	18.69	-31.87	5.44	5.13	6.81
19	3	1	10	55	28	18.69	-31.94	5.5	5.13	6.81
19	3	1	10	56	28	18.69	-31.87	5.19	5.13	6.88
19	3	1	10	57	28	18.69	-31.94	5.31	5.13	6.81
19	3	1	10	58	28	18.75	-32	5.06	5.13	6.81
19	3	1	10	59	29	18.75	-32	5.13	5.13	6.81
19	3	1	11	0	29	18.81	-32.06	5.25	5.13	6.81
19	3	1	11	1	29	18.81	-32.06	4.94	5.13	6.81
19	3	1	11	2	29	18.87	-32.06	5.13	5.19	6.81
19	3	1	11	3	30	18.87	-32.06	4.69	5.13	6.75
19	3	1	11	4	30	18.87	-32.13	4.38	5.13	6.75
19	3	1	11	5	30	18.87	-32.19	4.69	5.13	6.75
19	3	1	11	6	30	18.87	-32.25	4.19	5.13	6.75
19	3	1	11	7	31	18.87	-32.31	4.5	5.13	6.69
19	3	1	11	8	31	18.87	-32.31	4.06	5.13	6.69
19	3	1	11	9	31	18.81	-32.31	3.75	5.13	6.69
19	3	1	11	10	31	18.81	-32.31	4.13	5.13	6.63
19	3	1	11	11	32	18.81	-32.38	3.5	5.13	6.63
19	3	1	11	12	32	18.75	-32.44	3.88	5.13	6.63
19	3	1	11	13	32	18.75	-32.56	3.25	5.13	6.63
19	3	1	11	14	32	18.75	-32.63	2.75	5.13	6.56
19	3	1	11	15	33	18.75	-32.56	3.31	5.13	6.56
19	3	1	11	16	33	18.75	-32.63	2.69	5.13	6.56
19	3	1	11	17	33	18.75	-32.63	3.19	5.13	6.5
19	3	1	11	18	34	18.81	-32.69	2.63	5.13	6.5
19	3	1	11	19	34	18.81	-32.69	2.44	5.13	6.5
19	3	1	11	20	34	18.87	-32.75	2.75	5.06	6.44
19	3	1	11	21	34	18.87	-32.75	2	5.06	6.44
19	3	1	11	22	35	18.87	-32.69	2.56	5.06	6.44
19	3	1	11	23	35	18.94	-32.81	1.88	5.06	6.38
19	3	1	11	24	35	18.94	-32.75	2.38	5.06	6.38
19	3	1	11	25	35	19	-32.81	1.81	5.06	6.38
19	3	1	11	26	36	19	-32.88	1.25	5.06	6.38
19	3	1	11	27	36	19	-32.81	1.94	5.06	6.31
19	3	1	11	28	36	18.94	-32.88	1.31	5.06	6.31
19	3	1	11	29	36	19	-32.81	1.88	5.06	6.31
19	3	1	11	30	37	18.94	-32.88	1.31	5.06	6.31
19	3	1	11	31	37	19	-32.94	0.5	5.06	6.25
19	3	1	11	32	37	19	-32.81	1.37	5.06	6.25
19	3	1	11	33	37	19.06	-32.88	0.5	5.06	6.25
19	3	1	11	34	38	19.06	-32.81	1.25	5.06	6.25
19	3	1	11	35	38	19.12	-32.81	0.56	5	6.19
19	3	1	11	36	38	19.12	-32.81	0.81	5	6.19
19	3	1	11	37	38	19.12	-32.81	0.31	5	6.13
19	3	1	11	38	39	19.12	-32.88	-0.25	5	6.13
19	3	1	11	39	39	19.12	-32.81	0.56	5	6.13
19	3	1	11	40	39	19.12	-32.81	-0.06	5	6.13
19	3	1	11	41	40	19.12	-32.81	0.63	5	6.06
19	3	1	11	42	40	19.19	-32.81	-0.06	5	6.06
19	3	1	11	43	40	19.19	-32.94	-0.94	5	6.06
19	3	1	11	44	40	19.19	-32.81	0.06	5	6.06
19	3	1	11	45	41	19.19	-32.81	-0.75	4.94	6.06
19	3	1	11	46	41	19.19	-32.81	0.06	4.94	6
19	3	1	11	47	41	19.19	-32.81	-0.63	4.94	6
19	3	1	11	48	41	19.19	-32.75	-0.06	4.94	6
19	3	1	11	49	42	19.19	-32.81	-0.44	4.94	6
19	3	1	11	50	42	19.19	-32.81	-1.12	4.94	5.94
19	3	1	11	51	42	19.19	-32.75	-0.25	4.94	5.94
19	3	1	11	52	42	19.19	-32.81	-0.81	4.94	5.94
19	3	1	11	53	43	19.19	-32.75	-0.06	4.94	5.94
19	3	1	11	54	43	19.12	-32.81	-0.56	4.94	5.94
19	3	1	11	55	43	19.12	-32.88	-1.44	4.94	5.94
19	3	1	11	56	43	19.12	-32.75	-0.44	4.94	5.94
19	3	1	11	57	44	19.12	-32.81	-1	4.94	5.94
19	3	1	11	58	44	19.12	-32.75	-0.25	4.94	5.94
19	3	1	11	59	44	19.12	-32.81	-0.81	4.94	5.88
19	3	1	12	0	44	19.19	-32.75	-0.69	4.88	5.88
19	3	1	12	1	45	19.12	-32.75	-1	4.88	5.88
19	3	1	12	2	45	19.12	-32.81	-1.25	4.88	5.88
19	3	1	12	3	45	19.12	-32.75	-0.5	4.88	5.88
19	3	1	12	4	46	19.12	-32.75	-0.75	4.88	5.88
19	3	1	12	5	46	19.19	-32.69	-0.12	4.88	5.88
19	3	1	12	6	46	19.19	-32.75	-0.31	4.88	5.88
19	3	1	12	7	46	19.19	-32.75	-0.75	4.88	5.94
19	3	1	12	8	47	19.25	-32.69	0	4.88	5.94
19	3	1	12	9	47	19.25	-32.69	-0.31	4.88	5.94
19	3	1	12	10	47	19.25	-32.63	0.25	4.88	5.94
19	3	1	12	11	47	19.25	-32.56	0.69	4.88	5.94
19	3	1	12	12	48	19.25	-32.5	0.5	4.94	6
19	3	1	12	13	48	19.25	-32.44	1	4.94	6
19	3	1	12	14	48	19.25	-32.38	0.94	4.88	6.06
19	3	1	12	15	48	19.25	-32.38	1.31	4.94	6.06
19	3	1	12	16	49	19.19	-32.31	1.25	4.94	6.06
19	3	1	12	17	49	19.19	-32.31	1	4.94	6.13
19	3	1	12	18	49	19.19	-32.31	1.5	4.94	6.13
19	3	1	12	19	49	19.25	-32.31	1.37	4.94	6.13
19	3	1	12	20	50	19.25	-32.38	1.69	4.94	6.13
19	3	1	12	21	50	19.25	-32.38	1.94	4.94	6.19
19	3	1	12	22	50	19.25	-32.31	2.19	4.94	6.19
19	3	1	12	23	50	19.25	-32.25	2.38	4.94	6.19
19	3	1	12	24	51	19.25	-32.19	2.56	4.94	6.25
19	3	1	12	25	51	19.31	-32.13	2.69	5	6.25
19	3	1	12	26	51	19.31	-32.06	2.81	5	6.31
19	3	1	12	27	52	19.37	-32.06	2.94	5	6.31
19	3	1	12	28	52	19.37	-32	3.06	5	6.38
19	3	1	12	29	52	19.44	-32	3.13	5	6.38
19	3	1	12	30	52	19.44	-32	3.25	5	6.38
19	3	1	12	31	53	19.44	-31.94	3.31	5	6.44
19	3	1	12	32	53	19.5	-31.94	3.44	5	6.44
19	3	1	12	33	53	19.5	-31.94	3.5	5	6.44
19	3	1	12	34	53	19.5	-31.87	3.56	5	6.5
19	3	1	12	35	54	19.56	-31.87	3.69	5	6.5
19	3	1	12	36	54	19.56	-31.87	3.75	5.06	6.5
19	3	1	12	37	54	19.56	-31.87	3.81	5.06	6.56
19	3	1	12	38	54	19.62	-31.81	3.88	5.06	6.56
19	3	1	12	39	55	19.62	-31.81	3.94	5.06	6.56
19	3	1	12	40	55	19.62	-31.81	4	5.06	6.56
19	3	1	12	41	55	19.69	-31.81	4.06	5.06	6.56
19	3	1	12	42	55	19.69	-31.81	4.13	5.06	6.63
19	3	1	12	43	56	19.69	-31.75	4.19	5.06	6.63
19	3	1	12	44	56	19.69	-31.75	4.25	5.06	6.63
19	3	1	12	45	56	19.75	-31.75	4.31	5.06	6.63
19	3	1	12	46	56	19.75	-31.75	4.31	5.06	6.69
19	3	1	12	47	57	19.75	-31.75	4.38	5.06	6.69
19	3	1	12	48	57	19.75	-31.75	4.44	5.06	6.69
19	3	1	12	49	57	19.75	-31.75	4.5	5.06	6.69
19	3	1	12	51	1	19.75	-31.69	4.56	5.06	6.69
19	3	1	12	52	1	19.75	-31.75	4.5	5.06	6.69
19	3	1	12	53	2	19.75	-31.81	4.56	5.13	6.69
19	3	1	12	54	2	19.75	-31.81	4.63	5.06	6.75
19	3	1	12	55	2	19.75	-31.75	4.56	5.13	6.75
19	3	1	12	56	2	19.69	-31.87	4.69	5.13	6.75
19	3	1	12	57	3	19.69	-31.94	4.63	5.13	6.75
19	3	1	12	58	3	19.69	-31.94	4.75	5.13	6.75
19	3	1	12	59	3	19.69	-31.94	4.81	5.13	6.75
19	3	1	13	0	3	19.69	-31.94	4.75	5.13	6.75
19	3	1	13	1	4	19.69	-31.94	4.81	5.13	6.75
19	3	1	13	2	4	19.69	-31.94	4.56	5.13	6.75
19	3	1	13	3	4	19.62	-32	4.56	5.13	6.75
19	3	1	13	4	4	19.62	-32	4.69	5.13	6.75
19	3	1	13	5	5	19.62	-32.06	4.44	5.13	6.81
19	3	1	13	6	5	19.62	-32.06	4.63	5.13	6.75
19	3	1	13	7	5	19.62	-32.06	4.38	5.13	6.75
19	3	1	13	8	5	19.62	-32.06	4.44	5.13	6.75
19	3	1	13	9	6	19.62	-32.13	4.63	5.13	6.75
19	3	1	13	10	6	19.62	-32.06	4.31	5.13	6.75
19	3	1	13	11	6	19.62	-32.13	4.5	5.13	6.75
19	3	1	13	12	7	19.69	-32.13	4.13	5.13	6.75
19	3	1	13	13	7	19.69	-32.06	3.81	5.13	6.75
19	3	1	13	14	7	19.75	-32.06	4.13	5.13	6.69
19	3	1	13	15	7	19.75	-32.13	3.69	5.13	6.69
19	3	1	13	16	8	19.75	-32.13	4	5.13	6.69
19	3	1	13	17	8	19.75	-32.19	3.63	5.13	6.69
19	3	1	13	18	8	19.75	-32.25	3.25	5.13	6.63
19	3	1	13	19	8	19.75	-32.25	3.75	5.13	6.63
19	3	1	13	20	9	19.75	-32.31	3.25	5.13	6.63
19	3	1	13	21	9	19.69	-32.31	3.63	5.13	6.63
19	3	1	13	22	9	19.75	-32.44	3.06	5.13	6.56
19	3	1	13	23	9	19.69	-32.38	2.81	5.13	6.56
19	3	1	13	24	10	19.69	-32.5	3.13	5.13	6.56
19	3	1	13	25	10	19.69	-32.5	2.44	5.13	6.5
19	3	1	13	26	10	19.69	-32.5	2.94	5.06	6.5
19	3	1	13	27	10	19.69	-32.63	2.38	5.06	6.5
19	3	1	13	28	11	19.69	-32.56	2.88	5.06	6.5
19	3	1	13	29	11	19.69	-32.56	2.56	5.06	6.44
19	3	1	13	30	11	19.62	-32.56	2.25	5.06	6.5
19	3	1	13	31	11	19.62	-32.56	2.69	5.06	6.44
19	3	1	13	32	12	19.62	-32.63	1.94	5.06	6.44
19	3	1	13	33	12	19.62	-32.56	2.5	5.06	6.44
19	3	1	13	34	12	19.62	-32.63	1.81	5.06	6.44
19	3	1	13	35	13	19.62	-32.75	1.25	5.06	6.38
19	3	1	13	36	13	19.56	-32.69	1.94	5.06	6.38
19	3	1	13	37	13	19.56	-32.75	1.25	5.06	6.38
19	3	1	13	38	13	19.62	-32.69	1.81	5.06	6.38
19	3	1	13	39	14	19.62	-32.75	1.25	5.06	6.31
19	3	1	13	40	14	19.62	-32.81	0.69	5.06	6.31
19	3	1	13	41	14	19.69	-32.75	1.5	5.06	6.31
19	3	1	13	42	14	19.69	-32.81	0.63	5.06	6.31
19	3	1	13	43	15	19.69	-32.75	1.31	5	6.25
19	3	1	13	44	15	19.69	-32.81	0.63	5	6.25
19	3	1	13	45	15	19.69	-32.75	1.25	5	6.25
19	3	1	13	46	15	19.69	-32.81	0.81	5	6.19
19	3	1	13	47	16	19.62	-32.88	0.12	5	6.19
19	3	1	13	48	16	19.62	-32.75	0.94	5	6.19
19	3	1	13	49	16	19.62	-32.75	0.19	5	6.19
19	3	1	13	50	16	19.62	-32.75	0.88	5	6.19
19	3	1	13	51	17	19.56	-32.75	0.25	5	6.19
19	3	1	13	52	17	19.56	-32.88	-0.63	5	6.13
19	3	1	13	53	17	19.5	-32.81	0.37	5	6.13
19	3	1	13	54	17	19.5	-32.88	-0.5	5	6.13
19	3	1	13	55	18	19.5	-32.75	0.31	5	6.06
19	3	1	13	56	18	19.5	-32.75	-0.44	4.94	6.06
19	3	1	13	57	18	19.56	-32.69	0.25	4.94	6.06
19	3	1	13	58	19	19.56	-32.75	0.06	4.94	6.06
19	3	1	13	59	19	19.56	-32.81	-0.69	4.94	6.06
19	3	1	14	0	19	19.56	-32.69	0.12	4.94	6.06
19	3	1	14	1	19	19.56	-32.75	-0.63	4.94	6
19	3	1	14	2	20	19.56	-32.69	0.12	4.94	6
19	3	1	14	3	20	19.62	-32.69	-0.37	4.94	6
19	3	1	14	4	20	19.62	-32.75	-1.25	4.94	6
19	3	1	14	5	20	19.69	-32.69	-0.25	4.94	6
19	3	1	14	6	21	19.69	-32.75	-0.81	4.94	6
19	3	1	14	7	21	19.75	-32.69	-0.06	4.94	6
19	3	1	14	8	21	19.75	-32.75	-0.56	4.94	6
19	3	1	14	9	21	19.75	-32.63	0.06	4.94	6
19	3	1	14	10	22	19.75	-32.63	0.06	4.94	6
19	3	1	14	11	22	19.75	-32.69	-0.37	4.94	6
19	3	1	14	12	22	19.75	-32.63	0.31	4.94	6
19	3	1	14	13	22	19.75	-32.69	0	4.94	6
19	3	1	14	14	23	19.75	-32.63	0.56	4.94	6
19	3	1	14	15	23	19.75	-32.69	0.25	4.94	6
19	3	1	14	16	23	19.75	-32.75	-0.25	4.94	6
19	3	1	14	17	23	19.81	-32.69	0.44	4.94	6.06
19	3	1	14	18	24	19.81	-32.69	0.06	4.94	6.06
19	3	1	14	19	24	19.87	-32.69	0.63	4.94	6.06
19	3	1	14	20	24	19.87	-32.69	0.31	4.88	6.06
19	3	1	14	21	25	19.87	-32.63	0.25	4.94	6.06
19	3	1	14	22	25	19.94	-32.56	0.81	4.94	6.06
19	3	1	14	23	25	19.94	-32.5	0.81	4.94	6.13
19	3	1	14	24	25	19.94	-32.44	1.25	4.94	6.13
19	3	1	14	25	26	19.94	-32.44	1.19	4.94	6.13
19	3	1	14	26	26	19.94	-32.38	1	4.94	6.19
19	3	1	14	27	26	19.94	-32.31	1.5	4.94	6.19
19	3	1	14	28	26	20	-32.25	1.44	4.94	6.19
19	3	1	14	29	27	20	-32.19	1.75	4.94	6.19
19	3	1	14	30	27	20.06	-32.19	1.69	4.94	6.19
19	3	1	14	31	27	20.06	-32.13	2	5	6.25
19	3	1	14	32	27	20.06	-32.13	2.25	5	6.25
19	3	1	14	33	28	20	-32.06	2.44	5	6.31
19	3	1	14	34	28	20	-32	2.63	5	6.31
19	3	1	14	35	28	20	-31.94	2.75	5	6.31
19	3	1	14	36	28	20	-31.94	2.88	5	6.38
19	3	1	14	37	29	20	-31.87	3	5	6.38
19	3	1	14	38	29	20	-31.87	3.13	5	6.44
19	3	1	14	39	29	20	-31.81	3.25	5	6.44
19	3	1	14	40	29	20	-31.81	3.38	5.06	6.44
19	3	1	14	41	30	20	-31.81	3.44	5	6.5
19	3	1	14	42	30	20	-31.81	3.5	5.06	6.5
19	3	1	14	43	30	20	-31.75	3.5	5.06	6.5
19	3	1	14	44	31	20	-31.81	3.63	5.06	6.56
19	3	1	14	45	31	20	-31.87	3.63	5.06	6.56
19	3	1	14	46	31	20	-31.94	3.75	5.06	6.56
19	3	1	14	47	31	19.94	-32	3.81	5.06	6.56
19	3	1	14	48	32	19.94	-31.94	3.81	5.06	6.56
19	3	1	14	49	32	19.94	-32	3.94	5.06	6.63
19	3	1	14	50	32	19.87	-32	3.88	5.06	6.63
19	3	1	14	51	32	19.87	-32	3.88	5.06	6.63
19	3	1	14	52	33	19.87	-32	4.06	5.06	6.69
19	3	1	14	53	33	19.87	-32	3.75	5.06	6.69
19	3	1	14	54	33	19.87	-32	3.94	5.06	6.69
19	3	1	14	55	33	19.87	-32	3.69	5.06	6.69
19	3	1	14	56	34	19.87	-32	3.69	5.06	6.69
19	3	1	14	57	34	19.87	-32.06	3.94	5.06	6.69
19	3	1	14	58	34	19.87	-32.13	3.63	5.13	6.69
19	3	1	14	59	34	19.87	-32.13	3.88	5.06	6.69
19	3	1	15	0	35	19.87	-32.13	3.69	5.06	6.69
19	3	1	15	1	35	19.87	-32.13	3.5	5.13	6.69
19	3	1	15	2	35	19.94	-32.13	3.69	5.13	6.69
19	3	1	15	3	35	19.94	-32.19	3.25	5.13	6.63
19	3	1	15	4	36	19.94	-32.19	3.56	5.13	6.63
19	3	1	15	5	36	19.94	-32.25	3.19	5.13	6.63
19	3	1	15	6	36	20	-32.31	3.5	5.06	6.63
19	3	1	15	7	37	20	-32.25	3.69	5.06	6.63
19	3	1	15	8	37	20	-32.19	3.44	5.13	6.63
19	3	1	15	9	37	20	-32.13	3.69	5.06	6.63
19	3	1	15	10	37	19.94	-32.19	3.38	5.13	6.63
19	3	1	15	11	38	19.94	-32.25	3.63	5.06	6.63
19	3	1	15	12	38	19.94	-32.38	3.06	5.13	6.56
19	3	1	15	13	38	19.94	-32.44	2.63	5.06	6.56
19	3	1	15	14	38	19.87	-32.44	3.06	5.06	6.56
19	3	1	15	15	39	19.87	-32.5	2.5	5.06	6.56
19	3	1	15	16	39	19.87	-32.5	2.88	5.06	6.56
19	3	1	15	17	39	19.87	-32.63	2.44	5.06	6.5
19	3	1	15	18	39	19.87	-32.63	2	5.06	6.5
19	3	1	15	19	40	19.87	-32.56	2.56	5.06	6.5
19	3	1	15	20	40	19.81	-32.63	2.06	5.06	6.44
19	3	1	15	21	40	19.81	-32.56	2.5	5.06	6.44
19	3	1	15	22	40	19.81	-32.69	1.88	5.06	6.44
19	3	1	15	23	41	19.75	-32.75	1.31	5.06	6.44
19	3	1	15	24	41	19.75	-32.69	2	5.06	6.44
19	3	1	15	25	41	19.75	-32.81	1.31	5.06	6.38
19	3	1	15	26	41	19.75	-32.75	1.88	5.06	6.38
19	3	1	15	27	42	19.75	-32.88	1.31	5.06	6.38
19	3	1	15	28	42	19.69	-32.81	1.81	5	6.31
19	3	1	15	29	42	19.69	-32.69	1.81	5.06	6.31
19	3	1	15	30	43	19.69	-32.75	1.31	5.06	6.31
19	3	1	15	31	43	19.69	-32.63	1.88	5	6.31
19	3	1	15	32	43	19.69	-32.81	1.12	5	6.31
19	3	1	15	33	43	19.69	-32.75	1.69	5	6.31
19	3	1	15	34	44	19.69	-32.88	1	5	6.31
19	3	1	15	35	44	19.69	-33	0.25	5	6.25
19	3	1	15	36	44	19.69	-32.88	1.06	5	6.25
19	3	1	15	37	44	19.69	-33	0.31	5	6.25
19	3	1	15	38	45	19.69	-32.94	1	5	6.25
19	3	1	15	39	45	19.69	-33.06	0.37	5	6.19
19	3	1	15	40	45	19.69	-33	0.31	5	6.19
19	3	1	15	41	45	19.69	-33	0.56	5	6.19
19	3	1	15	42	46	19.69	-33.13	-0.37	5	6.19
19	3	1	15	43	46	19.69	-33	0.5	4.94	6.19
19	3	1	15	44	46	19.69	-33.13	-0.31	4.94	6.13
19	3	1	15	45	46	19.69	-33	0.37	5	6.13
19	3	1	15	46	47	19.69	-33.06	-0.25	4.94	6.13
19	3	1	15	47	47	19.69	-33.19	-1.06	4.94	6.06
19	3	1	15	48	47	19.69	-33	-0.12	4.94	6.06
19	3	1	15	49	47	19.75	-33.13	-0.88	4.94	6.06
19	3	1	15	50	48	19.75	-33	-0.06	4.94	6.06
19	3	1	15	51	48	19.81	-33.06	-0.63	4.94	6
19	3	1	15	52	48	19.81	-32.94	-0.63	4.94	6
19	3	1	15	53	49	19.75	-33	-0.75	4.94	6
19	3	1	15	54	49	19.81	-33.13	-1.37	4.94	5.94
19	3	1	15	55	49	19.81	-32.94	-0.5	4.94	5.94
19	3	1	15	56	49	19.75	-33.06	-1	4.88	5.94
19	3	1	15	57	50	19.75	-32.94	-0.25	4.88	5.94
19	3	1	15	58	50	19.75	-33	-0.69	4.88	5.94
19	3	1	15	59	50	19.75	-33.13	-1.44	4.88	5.94
19	3	1	16	0	50	19.75	-33	-0.5	4.88	5.94
19	3	1	16	1	51	19.75	-32.94	-0.75	4.88	5.94
19	3	1	16	2	51	19.75	-32.88	-0.06	4.88	5.94
19	3	1	16	3	51	19.69	-32.88	-0.31	4.88	5.94
19	3	1	16	4	51	19.69	-32.94	-0.75	4.88	5.94
19	3	1	16	5	52	19.75	-32.81	0.06	4.88	5.94
19	3	1	16	6	52	19.75	-32.81	-0.25	4.88	6
19	3	1	16	7	52	19.69	-32.75	0.37	4.88	6
19	3	1	16	8	52	20	-32.75	0.06	4.94	6
19	3	1	16	9	53	19.69	-32.69	0.44	4.88	6
19	3	1	16	10	53	19.69	-32.88	-0.06	4.88	6
19	3	1	16	11	53	19.69	-32.75	0	4.88	6
19	3	1	16	12	54	19.69	-32.63	0.5	4.88	6
19	3	1	16	13	54	19.69	-32.56	0.56	4.88	6.06
19	3	1	16	14	54	19.69	-32.5	1	4.88	6.06
19	3	1	16	15	54	19.69	-32.44	0.94	4.88	6.06
19	3	1	16	16	55	19.69	-32.38	0.88	4.88	6.13
19	3	1	16	17	55	19.75	-32.31	1.31	4.94	6.13
19	3	1	16	18	55	19.75	-32.25	1.25	4.94	6.13
19	3	1	16	19	55	19.75	-32.19	1.63	4.94	6.13
19	3	1	16	20	56	19.81	-32.19	1.56	4.94	6.19
19	3	1	16	21	56	19.81	-32.19	1.81	4.94	6.19
19	3	1	16	22	56	19.81	-32.13	2.06	4.94	6.19
19	3	1	16	23	56	19.81	-32.06	2.31	4.94	6.25
19	3	1	16	24	57	19.81	-32	2.44	4.94	6.25
19	3	1	16	25	57	19.81	-31.94	2.63	4.94	6.31
19	3	1	16	26	57	19.81	-31.87	2.75	4.94	6.31
19	3	1	16	27	57	19.75	-31.87	2.88	5	6.38
19	3	1	16	28	58	19.75	-31.81	3	5	6.38
19	3	1	16	29	58	19.75	-31.81	3.13	5	6.38
19	3	1	16	31	0	19.75	-31.75	3.25	5	6.44
19	3	1	16	32	0	19.75	-31.75	3.31	5	6.44
19	3	1	16	33	1	19.75	-31.69	3.44	5	6.5
19	3	1	16	34	1	19.75	-31.69	3.5	5	6.5
19	3	1	16	35	1	19.75	-31.69	3.56	5	6.5
19	3	1	16	36	2	19.75	-31.69	3.69	5	6.56
19	3	1	16	37	2	19.75	-31.62	3.75	5	6.56
19	3	1	16	38	2	19.75	-31.62	3.81	5	6.56
19	3	1	16	39	2	19.75	-31.62	3.88	5.06	6.56
19	3	1	16	40	3	19.75	-31.62	3.94	5.06	6.63
19	3	1	16	41	3	19.75	-31.56	4	5.06	85
19	3	1	16	42	3	19.75	-31.56	4.06	5	6.63
19	3	1	16	43	3	19.75	-31.56	4.13	5.06	6.63
19	3	1	16	44	4	19.69	-31.56	4.19	5.06	6.69
19	3	1	16	45	4	19.69	-31.56	4.25	5.06	6.69
19	3	1	16	46	4	19.69	-31.5	4.31	5.06	6.69
19	3	1	16	47	4	19.62	-31.5	4.38	5.06	6.69
19	3	1	16	48	5	19.62	-31.5	4.44	5.06	6.69
19	3	1	16	49	5	19.62	-31.5	4.5	5.06	6.69
19	3	1	16	50	5	19.62	-31.5	4.5	5.06	6.75
19	3	1	16	51	5	19.62	-31.5	4.56	5.06	6.75
19	3	1	16	52	6	19.56	-31.44	4.63	5.06	6.75
19	3	1	16	53	6	19.56	-31.44	4.69	5.06	6.75
19	3	1	16	54	6	19.5	-31.5	4.63	5.06	6.75
19	3	1	16	55	7	19.5	-31.62	4.69	5.06	6.75
19	3	1	16	56	7	19.5	-31.69	4.75	5.06	6.75
19	3	1	16	57	7	19.5	-31.62	4.81	5.06	6.81
19	3	1	16	58	7	19.44	-31.62	4.88	5.06	6.81
19	3	1	16	59	8	19.44	-31.62	4.81	5.13	6.81
19	3	1	17	0	8	19.44	-31.62	4.88	5.13	6.81
19	3	1	17	1	8	19.44	-31.62	4.94	5.13	6.81
19	3	1	17	2	8	19.37	-31.62	4.88	5.13	6.81
19	3	1	17	3	9	19.37	-31.69	4.94	5.13	6.81
19	3	1	17	4	9	19.37	-31.62	4.69	5.13	6.81
19	3	1	17	5	9	19.37	-31.69	4.81	5.13	6.81
19	3	1	17	6	9	19.37	-31.75	4.56	5.13	6.81
19	3	1	17	7	10	19.37	-31.75	4.63	5.06	6.81
19	3	1	17	8	10	19.37	-31.81	4.75	5.13	6.81
19	3	1	17	9	10	19.37	-31.81	4.5	5.13	6.81
19	3	1	17	10	10	19.37	-31.81	4.69	5.13	6.81
19	3	1	17	11	11	19.37	-31.87	4.5	5.13	6.81
19	3	1	17	12	11	19.37	-31.87	4.19	5.13	6.81
19	3	1	17	13	11	19.37	-31.87	4.44	5.13	6.75
19	3	1	17	14	11	19.37	-31.94	4	5.13	6.75
19	3	1	17	15	12	19.37	-32	4.25	5.13	6.75
19	3	1	17	16	12	19.37	-32.06	3.94	5.13	6.69
19	3	1	17	17	12	19.37	-32.06	3.63	5.13	6.69
19	3	1	17	18	12	19.31	-32.06	4	5.13	6.69
19	3	1	17	19	13	19.31	-32.06	3.56	5.13	6.69
19	3	1	17	20	13	19.31	-32.06	3.88	5.13	6.69
19	3	1	17	21	13	19.31	-32.13	3.56	5.13	6.69
19	3	1	17	22	14	19.31	-32.25	3	5.13	6.63
19	3	1	17	23	14	19.25	-32.19	3.5	5.13	6.63
19	3	1	17	24	14	19.25	-32.31	2.94	5.13	6.63
19	3	1	17	25	14	19.19	-32.25	3.31	5.06	6.56
19	3	1	17	26	15	19.19	-32.38	2.81	5.06	6.56
19	3	1	17	27	15	19.19	-32.44	2.31	5.06	6.56
19	3	1	17	28	15	19.19	-32.38	2.94	5.06	6.56
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	3	1	17	32	29	19.19	-32.06	3.94	5.06	85
19	3	1	17	33	29	19.19	-31.94	4.13	5.06	6.63
19	3	1	17	34	29	19.19	-31.94	4.13	5.13	6.63
19	3	1	17	35	30	19.19	-31.94	4.25	5.13	6.63
19	3	1	17	36	30	19.12	-31.94	4.25	5.13	6.63
19	3	1	17	37	30	19.12	-31.94	4.38	5.13	6.69
19	3	1	17	38	31	19.12	-31.94	4.44	5.13	6.69
19	3	1	17	39	31	19.12	-31.94	4.5	5.13	6.69
19	3	1	17	40	31	19.12	-31.94	4.63	5.13	6.69
19	3	1	17	41	31	19.12	-31.87	4.56	5.13	6.75
19	3	1	17	42	32	19.12	-31.94	4.69	5.13	6.75
19	3	1	17	43	32	19.06	-31.94	4.75	5.13	6.75
19	3	1	17	44	32	19.12	-31.87	4.69	5.13	6.75
19	3	1	17	45	32	19.06	-31.94	4.81	5.13	6.75
19	3	1	17	46	33	19.06	-31.94	4.81	5.13	6.75
19	3	1	17	47	33	19.06	-31.94	4.88	5.13	6.81
19	3	1	17	48	33	19	-31.94	4.88	5.13	6.81
19	3	1	17	49	33	19	-31.94	4.94	5.13	6.81
19	3	1	17	50	34	19	-31.94	5	5.13	6.81
19	3	1	17	51	34	19	-31.87	5.06	5.13	6.81
19	3	1	17	52	34	19	-31.81	5.13	5.13	6.81
19	3	1	17	53	34	19	-31.75	5.19	5.13	6.88
19	3	1	17	54	35	19	-31.69	5.25	5.19	6.88
19	3	1	17	55	35	19	-31.62	5.25	5.19	6.88
19	3	1	17	56	35	19	-31.56	5.31	5.19	6.88
19	3	1	17	57	35	19	-31.5	5.31	5.19	6.88
19	3	1	17	58	36	19	-31.5	5.38	5.19	6.88
19	3	1	17	59	36	19	-31.44	5.44	5.19	6.88
19	3	1	18	0	36	19	-31.44	5.44	5.19	6.94
19	3	1	18	1	36	19.06	-31.37	5.5	5.19	6.94
19	3	1	18	2	37	19.06	-31.37	5.5	5.19	6.94
19	3	1	18	3	37	19.12	-31.37	5.56	5.19	6.94
19	3	1	18	4	37	19.12	-31.31	5.56	5.19	6.94
19	3	1	18	5	38	19.12	-31.31	5.56	5.19	6.94
19	3	1	18	6	38	19.19	-31.31	5.63	5.19	6.94
19	3	1	18	7	38	19.19	-31.31	5.63	5.19	6.94
19	3	1	18	8	38	19.25	-31.31	5.69	5.19	6.94
19	3	1	18	9	39	19.25	-31.25	5.69	5.19	6.94
19	3	1	18	10	39	19.19	-31.25	5.69	5.19	6.94
19	3	1	18	11	39	19.19	-31.31	5.63	5.19	6.94
19	3	1	18	12	39	19.19	-31.37	5.69	5.19	6.94
19	3	1	18	13	40	19.19	-31.37	5.69	5.19	7
19	3	1	18	14	40	19.19	-31.37	5.75	5.19	6.94
19	3	1	18	15	40	19.19	-31.37	5.75	5.19	7
19	3	1	18	16	40	19.12	-31.44	5.69	5.19	7
19	3	1	18	17	41	19.12	-31.5	5.75	5.19	7
19	3	1	18	18	41	19.12	-31.5	5.81	5.19	7
19	3	1	18	19	41	19.12	-31.5	5.75	5.19	7
19	3	1	18	20	41	19.12	-31.5	5.81	5.25	7
19	3	1	18	21	42	19.12	-31.5	5.5	5.19	7
19	3	1	18	22	42	19.12	-31.62	5.56	5.19	7
19	3	1	18	23	42	19.12	-31.56	5.31	5.19	7
19	3	1	18	24	42	19.06	-31.56	5.38	5.25	6.94
19	3	1	18	25	43	19.06	-31.69	5.5	5.19	6.94
19	3	1	18	26	43	19.06	-31.69	5.19	5.19	6.94
19	3	1	18	27	43	19.06	-31.75	5.38	5.25	6.94
19	3	1	18	28	44	19.06	-31.69	5.13	5.19	6.94
19	3	1	18	29	44	19.06	-31.75	4.88	5.25	6.94
19	3	1	18	30	44	19.06	-31.75	5.13	5.19	6.94
19	3	1	18	31	44	19	-31.81	4.69	5.25	6.88
19	3	1	18	32	45	19	-31.81	4.94	5.25	6.88
19	3	1	18	33	45	19	-31.87	4.56	5.19	6.88
19	3	1	18	34	45	19.06	-31.81	4.25	5.19	6.81
19	3	1	18	35	45	19.06	-31.81	4.63	5.19	6.81
19	3	1	18	36	46	19.06	-31.87	4.13	5.19	6.81
19	3	1	18	37	46	19.06	-31.87	4.44	5.19	6.81
19	3	1	18	38	46	19.06	-32	4.06	5.19	6.75
19	3	1	18	39	46	19.06	-32.06	3.56	5.19	6.75
19	3	1	18	40	47	19.12	-32	4.06	5.19	6.75
19	3	1	18	41	47	19.12	-32.13	3.44	5.19	6.75
19	3	1	18	42	47	19.12	-32.13	3.88	5.19	6.69
19	3	1	18	43	47	19.12	-32.19	3.31	5.19	6.69
19	3	1	18	44	48	19.12	-32.25	2.81	5.19	6.69
19	3	1	18	45	48	19.12	-32.25	3.38	5.19	6.69
19	3	1	18	46	48	19.12	-32.31	2.81	5.19	6.63
19	3	1	18	47	48	19.12	-32.25	3.31	5.19	6.63
19	3	1	18	48	49	19.19	-32.31	2.81	5.19	6.63
19	3	1	18	49	49	19.19	-32.25	3.19	5.19	6.63
19	3	1	18	50	49	19.19	-32.38	2.69	5.19	6.56
19	3	1	18	51	50	19.25	-32.5	2.13	5.19	6.56
19	3	1	18	52	50	19.25	-32.44	2.69	5.13	6.56
19	3	1	18	53	50	19.25	-32.5	2	5.19	6.56
19	3	1	18	54	50	19.25	-32.44	2.5	5.19	6.5
19	3	1	18	55	51	19.25	-32.5	1.94	5.13	6.5
19	3	1	18	56	51	19.25	-32.56	1.44	5.13	6.5
19	3	1	18	57	51	19.25	-32.44	2.13	5.13	6.5
19	3	1	18	58	51	19.25	-32.5	1.5	5.13	6.44
19	3	1	18	59	52	19.25	-32.44	2.06	5.13	6.44
19	3	1	19	0	52	19.25	-32.56	1.56	5.13	6.44
19	3	1	19	1	52	19.25	-32.63	0.69	5.13	6.44
19	3	1	19	2	52	19.25	-32.5	1.56	5.13	6.44
19	3	1	19	3	53	19.25	-32.63	0.75	5.13	6.38
19	3	1	19	4	53	19.31	-32.5	1.44	5.13	6.38
19	3	1	19	5	53	19.31	-32.63	0.81	5.13	6.38
19	3	1	19	6	53	19.31	-32.5	1.37	5.13	6.31
19	3	1	19	7	54	19.31	-32.5	1.12	5.13	6.31
19	3	1	19	8	54	19.25	-32.63	0.31	5.13	6.31
19	3	1	19	9	54	19.25	-32.5	1.06	5.13	6.31
19	3	1	19	10	55	19.25	-32.56	0.25	5.13	6.31
19	3	1	19	11	55	19.25	-32.44	0.94	5.13	6.25
19	3	1	19	12	55	19.25	-32.5	0.19	5.06	6.25
19	3	1	19	13	55	19.19	-32.63	-0.69	5.06	6.25
19	3	1	19	14	56	19.19	-32.5	0.31	5.06	6.19
19	3	1	19	15	56	19.19	-32.63	-0.5	5.06	6.19
19	3	1	19	16	56	19.19	-32.5	0.31	5.06	6.19
19	3	1	19	17	56	19.19	-32.56	-0.37	5.06	6.19
19	3	1	19	18	57	19.25	-32.44	0.31	5.06	6.13
19	3	1	19	19	57	19.25	-32.56	-0.19	5.06	6.13
19	3	1	19	20	57	19.25	-32.69	-1.12	5.06	6.13
19	3	1	19	21	57	19.25	-32.5	-0.19	5.06	6.13
19	3	1	19	22	58	19.31	-32.63	-1.06	5.06	6.13
19	3	1	19	23	58	19.31	-32.5	-0.25	5	6.06
19	3	1	19	24	58	19.31	-32.63	-1	5	6.06
19	3	1	19	25	58	19.37	-32.69	-1.69	5	6
19	3	1	19	26	59	19.37	-32.56	-0.81	5	6
19	3	1	19	27	59	19.37	-32.69	-1.75	5	6
19	3	1	19	28	59	19.37	-32.5	-0.75	5	6
19	3	1	19	29	59	19.37	-32.63	-1.56	5	6
19	3	1	19	31	0	19.37	-32.44	-0.75	5	5.94
19	3	1	19	32	0	19.37	-32.56	-1.56	4.94	5.94
19	3	1	19	33	0	19.37	-32.69	-2.63	5	5.94
19	3	1	19	34	1	19.37	-32.5	-1.37	5	5.88
19	3	1	19	35	1	19.37	-32.56	-2.38	4.94	5.88
19	3	1	19	36	1	19.37	-32.44	-1.37	5	5.88
19	3	1	19	37	1	19.31	-32.56	-2.19	4.94	5.81
19	3	1	19	38	2	19.31	-32.38	-1.31	4.94	5.81
19	3	1	19	39	2	19.31	-32.5	-1.94	4.94	5.81
19	3	1	19	40	2	19.31	-32.63	-2.81	4.94	5.81
19	3	1	19	41	2	19.31	-32.44	-1.69	4.94	5.81
19	3	1	19	42	3	19.25	-32.63	-2.44	4.94	5.75
19	3	1	19	43	3	19.25	-32.44	-1.5	4.94	5.75
19	3	1	19	44	3	19.25	-32.56	-2.19	4.94	5.75
19	3	1	19	45	3	19.25	-32.5	-2.19	4.94	5.75
19	3	1	19	46	4	19.19	-32.5	-2.25	4.94	5.75
19	3	1	19	47	4	19.25	-32.63	-3.13	4.88	5.75
19	3	1	19	48	4	19.19	-32.44	-2	4.94	5.69
19	3	1	19	49	4	19.19	-32.5	-2.38	4.88	5.69
19	3	1	19	50	5	19.19	-32.38	-1.56	4.88	5.69
19	3	1	19	51	5	19.19	-32.44	-2	4.88	5.69
19	3	1	19	52	5	19.12	-32.63	-2.81	4.88	5.69
19	3	1	19	53	5	19.12	-32.44	-1.69	4.88	5.75
19	3	1	19	54	6	19.12	-32.56	-2.25	4.88	5.75
19	3	1	19	55	6	19.12	-32.44	-1.44	4.88	5.75
19	3	1	19	56	6	19.12	-32.5	-1.94	4.88	5.75
19	3	1	19	57	6	19.06	-32.38	-1.37	4.88	5.75
19	3	1	19	58	7	19.06	-32.5	-1.75	4.88	5.75
19	3	1	19	59	7	19.06	-32.56	-1.94	4.88	5.75
19	3	1	20	0	7	19	-32.44	-1.12	4.88	5.75
19	3	1	20	1	8	19	-32.44	-1.37	4.88	5.81
19	3	1	20	2	8	19	-32.31	-0.75	4.88	5.81
19	3	1	20	3	8	19	-32.38	-0.94	4.88	5.81
19	3	1	20	4	8	19	-32.44	-1.31	4.88	5.81
19	3	1	20	5	9	19.06	-32.31	-0.56	4.88	5.88
19	3	1	20	6	9	19.06	-32.38	-0.88	4.88	85
19	3	1	20	7	9	19.06	-32.25	-0.31	4.88	85
19	3	1	20	8	9	19.06	-32.31	-0.56	4.88	85
19	3	1	20	9	10	19.06	-32.25	-0.63	4.88	85
19	3	1	20	10	10	19.12	-32.19	0	4.94	85
19	3	1	20	11	10	19.12	-32.13	0.06	4.88	85
19	3	1	20	12	10	19.12	-32.06	0.5	4.94	85
19	3	1	20	13	11	19.12	-32.06	0.5	4.94	85
19	3	1	20	14	11	19.12	-32	0.31	4.94	85
19	3	1	20	15	11	19.12	-31.94	0.81	4.94	85
19	3	1	20	16	11	19.12	-31.87	0.75	4.94	85
19	3	1	20	17	12	19.12	-31.87	1.12	4.94	85
19	3	1	20	18	12	19.12	-31.81	1.06	4.94	85
19	3	1	20	19	12	19.06	-31.81	1.37	4.94	85
19	3	1	20	20	13	19.06	-31.75	1.63	4.94	85
19	3	1	20	21	13	19.06	-31.69	1.88	5	85
19	3	1	20	22	13	19.06	-31.62	2.06	5	85
19	3	1	20	23	13	19.06	-31.56	2.19	5	85
19	3	1	20	24	14	19	-31.5	2.38	5	85
19	3	1	20	25	14	19	-31.44	2.5	5	85
19	3	1	20	26	14	19	-31.44	2.63	5	85
19	3	1	20	27	14	19	-31.37	2.75	5	85
19	3	1	20	28	15	19	-31.37	2.88	5	85
19	3	1	20	29	15	19	-31.31	2.94	5	85
19	3	1	20	30	15	18.94	-31.31	3.06	5	85
19	3	1	20	31	15	18.94	-31.31	3.13	5.06	85
19	3	1	20	32	16	18.94	-31.31	3.25	5.06	85
19	3	1	20	33	16	18.94	-31.25	3.31	5.06	85
19	3	1	20	34	16	18.94	-31.25	3.44	5.06	85
19	3	1	20	35	16	18.94	-31.25	3.5	5.06	85
19	3	1	20	36	17	18.94	-31.19	3.56	5.06	85
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	3	1	20	41	11	18.06	-31.37	4	5.25	7.56
19	3	1	20	42	11	19.06	-31.19	4.13	5.44	7.81
19	3	1	20	43	12	19.12	-31.19	4.19	5.5	8.06
19	3	1	20	44	12	19.12	-31.25	4.38	5.56	8.13
19	3	1	20	45	12	19.12	-31.31	4.5	5.63	8.19
19	3	1	20	46	12	19.12	-31.44	4.63	5.63	8.13
19	3	1	20	47	13	19.12	-31.44	4.75	5.63	8.13
19	3	1	20	48	13	19.12	-31.44	4.75	5.63	8.13
19	3	1	20	49	13	19.12	-31.5	4.88	5.63	8.06
19	3	1	20	50	13	19.06	-31.5	4.94	5.63	8.06
19	3	1	20	51	14	19.06	-31.56	5	5.63	8.06
19	3	1	20	52	14	19.06	-31.5	5.06	5.63	8
19	3	1	20	53	14	19.06	-31.44	5.06	5.63	8.06
19	3	1	20	54	14	19.06	-31.5	5.13	5.63	8
19	3	1	20	55	15	19.06	-31.5	5.13	5.63	8
19	3	1	20	56	15	19.06	-31.5	5.19	5.63	8
19	3	1	20	57	15	19.12	-31.56	5.19	5.63	8
19	3	1	20	58	15	19.12	-31.56	5.19	5.63	7.94
19	3	1	20	59	16	19.12	-31.56	5.25	5.63	7.94
19	3	1	21	0	16	19.12	-31.56	5.25	5.63	7.94
19	3	1	21	1	16	19.12	-31.56	5.31	5.63	7.94
19	3	1	21	2	17	19.12	-31.5	5.31	5.63	7.94
19	3	1	21	3	17	19.06	-31.5	5.31	5.63	7.94
19	3	1	21	4	17	19.06	-31.5	5.38	5.63	7.88
19	3	1	21	5	17	19.06	-31.44	5.38	5.63	7.88
19	3	1	21	6	18	19.06	-31.5	5.44	5.63	7.88
19	3	1	21	7	18	19.06	-31.5	5.5	5.63	7.88
19	3	1	21	8	18	19	-31.5	5.44	5.63	7.88
19	3	1	21	9	18	19	-31.5	5.56	5.63	7.88
19	3	1	21	10	19	19	-31.5	5.5	5.63	7.88
19	3	1	21	11	19	19	-31.5	5.56	5.63	7.88
19	3	1	21	12	19	19	-31.44	5.56	5.63	7.88
19	3	1	21	13	19	19	-31.44	5.56	5.63	7.88
19	3	1	21	14	20	19	-31.5	5.63	5.63	7.88
19	3	1	21	15	20	19	-31.5	5.63	5.63	7.88
19	3	1	21	16	20	19	-31.5	5.69	5.63	7.88
19	3	1	21	17	20	19	-31.5	5.69	5.63	7.88
19	3	1	21	18	21	18.94	-31.44	5.69	5.63	7.88
19	3	1	21	19	21	18.94	-31.5	5.75	5.63	7.88
19	3	1	21	20	21	18.94	-31.5	5.44	5.63	7.88
19	3	1	21	21	21	18.94	-31.5	5.56	5.63	7.88
19	3	1	21	22	22	18.87	-31.5	5.31	5.63	7.88
19	3	1	21	23	22	18.87	-31.44	5.38	5.63	7.81
19	3	1	21	24	22	18.87	-31.56	5.5	5.63	7.81
19	3	1	21	25	22	18.87	-31.5	5.19	5.63	7.81
19	3	1	21	26	23	18.94	-31.44	5.38	5.63	7.81
19	3	1	21	27	23	18.94	-31.44	5.13	5.63	7.81
19	3	1	21	28	23	18.94	-31.5	4.88	5.56	7.81
19	3	1	21	29	24	18.94	-31.5	5.13	5.63	7.81
19	3	1	21	30	24	18.87	-31.56	4.63	5.56	7.81
19	3	1	21	31	24	18.94	-31.62	4.88	5.56	7.75
19	3	1	21	32	24	18.87	-31.69	4.44	5.56	7.75
19	3	1	21	33	25	18.87	-31.69	4.19	5.56	7.75
19	3	1	21	34	25	18.87	-31.69	4.5	5.56	7.75
19	3	1	21	35	25	18.87	-31.69	4.06	5.56	7.69
19	3	1	21	36	25	18.87	-31.69	4.38	5.56	7.69
19	3	1	21	37	26	18.87	-31.81	4	5.56	7.69
19	3	1	21	38	26	18.87	-31.87	3.5	5.5	7.63
19	3	1	21	39	26	18.87	-31.87	4	5.56	7.63
19	3	1	21	40	26	18.87	-32	3.38	5.5	7.63
19	3	1	21	41	27	18.87	-32	3.75	5.5	7.56
19	3	1	21	42	27	18.94	-32.13	3.25	5.5	7.56
19	3	1	21	43	27	18.94	-32.13	3.31	5.5	7.5
19	3	1	21	44	27	18.94	-32.13	3.31	5.5	7.5
19	3	1	21	45	28	18.94	-32.19	2.75	5.5	7.5
19	3	1	21	46	28	18.94	-32.13	3.19	5.5	7.44
19	3	1	21	47	28	18.94	-32.19	2.75	5.5	7.44
19	3	1	21	48	29	18.94	-32.13	3.13	5.5	7.44
19	3	1	21	49	29	18.94	-32.31	2.5	5.5	7.38
19	3	1	21	50	29	18.94	-32.38	1.94	5.5	7.38
19	3	1	21	51	29	19	-32.31	2.56	5.44	7.38
19	3	1	21	52	30	19	-32.38	1.88	5.44	7.31
19	3	1	21	53	30	19	-32.31	2.44	5.44	7.31
19	3	1	21	54	30	19	-32.44	1.88	5.44	7.31
19	3	1	21	55	30	19	-32.5	1.31	5.44	7.25
19	3	1	21	56	31	18.94	-32.38	2	5.44	7.25
19	3	1	21	57	31	18.94	-32.5	1.19	5.44	7.25
19	3	1	21	58	31	18.94	-32.44	1.81	5.44	7.19
19	3	1	21	59	31	18.94	-32.5	1.12	5.44	7.19
19	3	1	22	0	32	18.87	-32.44	1.19	5.38	7.13
19	3	1	22	1	32	18.87	-32.5	0.81	5.38	7.13
19	3	1	22	2	32	18.87	-32.63	0.06	5.38	7.06
19	3	1	22	3	32	18.87	-32.5	0.88	5.38	7.06
19	3	1	22	4	33	18.87	-32.56	0.25	5.38	7
19	3	1	22	5	33	18.81	-32.44	0.94	5.38	7
19	3	1	22	6	33	18.81	-32.5	0.37	5.31	6.94
19	3	1	22	7	33	18.81	-32.63	-0.37	5.31	6.94
19	3	1	22	8	34	18.81	-32.44	0.5	5.31	6.94
19	3	1	22	9	34	18.81	-32.56	-0.31	5.31	6.88
19	3	1	22	10	34	18.81	-32.44	0.44	5.31	6.88
19	3	1	22	11	35	18.81	-32.5	-0.25	5.31	6.88
19	3	1	22	12	35	18.87	-32.44	-0.44	5.31	6.88
19	3	1	22	13	35	18.81	-32.44	-0.06	5.31	6.81
19	3	1	22	14	35	18.87	-32.56	-0.88	5.31	6.81
19	3	1	22	15	36	18.81	-32.38	0	5.31	6.75
19	3	1	22	16	36	18.81	-32.5	-0.69	5.25	6.75
19	3	1	22	17	36	18.81	-32.38	0.06	5.25	6.75
19	3	1	22	18	36	18.81	-32.44	-0.75	5.25	6.69
19	3	1	22	19	37	18.81	-32.56	-1.63	5.25	6.69
19	3	1	22	20	37	18.81	-32.38	-0.56	5.25	6.69
19	3	1	22	21	37	18.75	-32.5	-1.37	5.25	6.63
19	3	1	22	22	37	18.75	-32.31	-0.56	5.25	6.63
19	3	1	22	23	38	18.75	-32.44	-1.25	5.25	6.63
19	3	1	22	24	38	18.75	-32.25	-0.5	5.25	6.56
19	3	1	22	25	38	18.75	-32.38	-1.06	5.25	6.56
19	3	1	22	26	38	18.75	-32.5	-1.81	5.19	6.56
19	3	1	22	27	39	18.75	-32.31	-0.88	5.19	6.5
19	3	1	22	28	39	18.81	-32.38	-1.75	5.19	6.5
19	3	1	22	29	39	18.81	-32.25	-0.94	5.19	6.5
19	3	1	22	30	39	18.81	-32.31	-1.69	5.19	6.5
19	3	1	22	31	40	18.81	-32.38	-2.44	5.19	6.44
19	3	1	22	32	40	18.87	-32.25	-1.44	5.19	6.44
19	3	1	22	33	40	18.87	-32.38	-2.31	5.19	6.44
19	3	1	22	34	40	18.87	-32.19	-1.37	5.19	6.44
19	3	1	22	35	41	18.87	-32.31	-2.19	5.13	6.38
19	3	1	22	36	41	18.87	-32.13	-1.31	5.13	6.38
19	3	1	22	37	41	18.87	-32.25	-2	5.13	6.38
19	3	1	22	38	42	18.87	-32.44	-3.13	5.13	6.38
19	3	1	22	39	42	18.87	-32.19	-1.88	5.13	6.31
19	3	1	22	40	42	18.87	-32.38	-2.88	5.13	6.31
19	3	1	22	41	42	18.81	-32.19	-1.81	5.13	6.25
19	3	1	22	42	43	18.81	-32.25	-2.69	5.06	6.25
19	3	1	22	43	43	18.81	-32.13	-1.81	5.06	6.25
19	3	1	22	44	43	18.81	-32.25	-2.56	5.06	6.19
19	3	1	22	45	43	18.81	-32.38	-3.56	5.06	6.19
19	3	1	22	46	44	18.75	-32.19	-2.31	5.06	6.19
19	3	1	22	47	44	18.75	-32.25	-3.38	5.06	6.13
19	3	1	22	48	44	18.75	-32.06	-2.31	5.06	6.13
19	3	1	22	49	44	18.75	-32.25	-3.25	5.06	6.13
19	3	1	22	50	45	18.69	-32.06	-2.31	5.06	6.06
19	3	1	22	51	45	18.69	-32.19	-3.13	5	6.06
19	3	1	22	52	45	18.69	-32.31	-4.19	5	6.06
19	3	1	22	53	45	18.75	-32.06	-2.88	5	6
19	3	1	22	54	46	18.75	-32.25	-3.88	5	6
19	3	1	22	55	46	18.75	-32.06	-2.81	5	5.94
19	3	1	22	56	46	18.75	-32.13	-3.63	4.94	5.94
19	3	1	22	57	47	18.75	-31.94	-2.69	5	5.94
19	3	1	22	58	47	18.75	-32.06	-3.31	5	5.88
19	3	1	22	59	47	18.75	-32.25	-4.44	4.94	5.94
19	3	1	23	0	47	18.75	-32	-3.06	4.94	5.88
19	3	1	23	1	48	18.75	-32.13	-3.94	4.94	5.88
19	3	1	23	2	48	18.69	-31.94	-2.88	4.94	5.88
19	3	1	23	3	48	18.69	-32.13	-3.63	4.94	5.88
19	3	1	23	4	48	18.69	-31.94	-2.75	4.94	5.88
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	3	1	23	9	35	18.69	-31.5	-1.06	4.94	6
19	3	1	23	10	36	18.69	-31.44	-0.63	5	85
19	3	1	23	11	36	18.69	-31.31	-0.44	5	6.13
19	3	1	23	12	36	18.69	-31.25	-0.06	5	6.19
19	3	1	23	13	36	18.69	-31.25	0.12	5	6.31
19	3	1	23	14	37	18.69	-31.25	0.37	5.06	6.38
19	3	1	23	15	37	18.69	-31.25	0.56	5.06	6.44
19	3	1	23	16	37	18.69	-31.25	0.69	5.06	6.5
19	3	1	23	17	37	18.69	-31.25	0.94	5.06	6.56
19	3	1	23	18	38	18.69	-31.25	1.06	5.06	6.63
19	3	1	23	19	38	18.69	-31.25	1.25	5.06	6.69
19	3	1	23	20	38	18.69	-31.25	1.37	5.13	6.75
19	3	1	23	21	38	18.75	-31.25	1.44	5.13	6.81
19	3	1	23	22	39	18.75	-31.25	1.63	5.13	6.88
19	3	1	23	23	39	18.75	-31.25	1.69	5.13	6.88
19	3	1	23	24	39	18.75	-31.31	1.88	5.13	6.94
19	3	1	23	25	40	18.75	-31.31	1.94	5.13	7
19	3	1	23	26	40	18.75	-31.31	2.13	5.19	7.06
19	3	1	23	27	40	18.75	-31.25	2.25	5.13	7.06
19	3	1	23	28	40	18.75	-31.19	2.31	5.19	7.13
19	3	1	23	29	41	18.75	-31.25	2.44	5.19	7.19
19	3	1	23	30	41	18.75	-31.25	2.5	5.19	7.19
19	3	1	23	31	41	18.69	-31.31	2.63	5.19	7.25
19	3	1	23	32	41	18.69	-31.25	2.75	5.19	7.25
19	3	1	23	33	42	18.62	-31.25	2.75	5.25	7.31
19	3	1	23	34	42	18.62	-31.25	2.88	5.19	7.31
19	3	1	23	35	42	18.62	-31.25	2.94	5.25	7.38
19	3	1	23	36	42	18.62	-31.31	3	5.25	7.38
19	3	1	23	37	43	18.56	-31.25	3.13	5.25	7.44
19	3	1	23	38	43	18.56	-31.25	3.13	5.25	7.44
19	3	1	23	39	43	18.5	-31.25	3.25	5.25	7.44
19	3	1	23	40	43	18.5	-31.25	3.31	5.25	7.5
19	3	1	23	41	44	18.5	-31.25	3.31	5.25	7.5
19	3	1	23	42	44	18.5	-31.31	3.44	5.25	7.5
19	3	1	23	43	44	18.5	-31.31	3.44	5.25	7.56
19	3	1	23	44	44	18.5	-31.31	3.56	5.25	7.56
19	3	1	23	45	45	18.44	-31.37	3.63	5.25	7.56
19	3	1	23	46	45	18.44	-31.31	3.56	5.25	7.56
19	3	1	23	47	45	18.44	-31.37	3.75	5.31	7.63
19	3	1	23	48	45	18.44	-31.37	3.75	5.31	7.63
19	3	1	23	49	46	18.44	-31.37	3.88	5.31	7.63
19	3	1	23	50	46	18.44	-31.31	3.88	5.31	7.63
19	3	1	23	51	46	18.5	-31.31	3.88	5.31	7.69
19	3	1	23	52	46	18.5	-31.31	4	5.31	7.69
19	3	1	23	53	47	18.5	-31.31	4	5.31	7.69
19	3	1	23	54	47	18.5	-31.37	4.13	5.31	7.69
19	3	1	23	55	47	18.5	-31.37	4.13	5.31	7.69
19	3	1	23	56	48	18.5	-31.31	3.81	5.31	7.69
19	3	1	23	57	48	18.5	-31.31	4	5.38	7.69
19	3	1	23	58	48	18.5	-31.31	3.81	5.38	7.69
19	3	1	23	59	48	18.44	-31.31	4	5.31	7.69
19	3	2	0	0	49	18.44	-31.31	3.81	5.38	7.69
19	3	2	0	1	49	18.44	-31.37	3.56	5.38	7.69
19	3	2	0	2	49	18.37	-31.37	3.88	5.31	7.69
19	3	2	0	3	49	18.37	-31.37	3.69	5.38	7.69
19	3	2	0	4	50	18.37	-31.44	3.88	5.38	7.63
19	3	2	0	5	50	18.37	-31.37	3.69	5.38	7.69
19	3	2	0	6	50	18.31	-31.44	3.25	5.31	7.69
19	3	2	0	7	50	18.31	-31.37	3.56	5.38	7.63
19	3	2	0	8	51	18.31	-31.5	3.19	5.38	7.63
19	3	2	0	9	51	18.31	-31.5	3.5	5.31	7.63
19	3	2	0	10	51	18.31	-31.56	3.19	5.31	7.56
19	3	2	0	11	51	18.31	-31.69	2.69	5.31	7.56
19	3	2	0	12	52	18.31	-31.62	3.19	5.31	7.5
19	3	2	0	13	52	18.31	-31.69	2.81	5.31	7.5
19	3	2	0	14	52	18.31	-31.69	3.19	5.31	7.5
19	3	2	0	15	53	18.31	-31.75	2.88	5.31	7.5
19	3	2	0	16	53	18.25	-31.81	2.13	5.31	7.44
19	3	2	0	17	53	18.25	-31.81	2.69	5.31	7.44
19	3	2	0	18	53	18.25	-31.94	2.19	5.31	7.44
19	3	2	0	19	54	18.19	-31.87	2.63	5.31	7.38
19	3	2	0	20	54	18.19	-31.94	2.19	5.25	7.38
19	3	2	0	21	54	18.19	-31.87	2.56	5.31	7.31
19	3	2	0	22	54	18.19	-31.81	2.75	5.25	7.31
19	3	2	0	23	55	18.12	-31.87	2.13	5.25	7.38
19	3	2	0	24	55	18.19	-31.81	2.63	5.25	7.31
19	3	2	0	25	55	18.12	-32	1.94	5.25	7.31
19	3	2	0	26	55	18.12	-31.94	2.44	5.25	7.31
19	3	2	0	27	56	18.19	-32.13	1.81	5.25	7.25
19	3	2	0	28	56	18.19	-32.25	1.06	5.25	7.25
19	3	2	0	29	56	18.19	-32.13	1.75	5.25	7.19
19	3	2	0	30	56	18.19	-32.25	1.12	5.25	7.19
19	3	2	0	31	57	18.25	-32.13	1.69	5.25	7.19
19	3	2	0	32	57	18.25	-32.25	1.19	5.25	7.13
19	3	2	0	33	57	18.25	-32.38	0.44	5.19	7.13
19	3	2	0	34	57	18.25	-32.31	1.25	5.19	7.13
19	3	2	0	35	58	18.25	-32.44	0.44	5.19	7.06
19	3	2	0	36	58	18.25	-32.31	1.12	5.19	7.06
19	3	2	0	37	58	18.25	-32.44	0.44	5.19	7
19	3	2	0	38	58	18.19	-32.31	1	5.19	7
19	3	2	0	39	59	18.19	-32.38	0.63	5.19	7
19	3	2	0	40	59	18.19	-32.5	-0.19	5.19	6.94
19	3	2	0	41	59	18.12	-32.38	0.63	5.19	6.94
19	3	2	0	42	59	18.12	-32.44	-0.06	5.19	6.94
19	3	2	0	44	0	18.12	-32.31	0.63	5.19	6.88
19	3	2	0	45	0	18.12	-32.44	0.06	5.13	6.88
19	3	2	0	46	0	18.12	-32.56	-1.06	5.13	6.88
19	3	2	0	47	1	18.12	-32.38	0	5.13	6.81
19	3	2	0	48	1	18.12	-32.5	-0.88	5.13	6.81
19	3	2	0	49	1	18.12	-32.31	-0.06	5.13	6.75
19	3	2	0	50	1	18.12	-32.44	-0.69	5.13	6.75
19	3	2	0	51	2	18.12	-32.25	0	5.13	6.75
19	3	2	0	52	2	18.12	-32.38	-0.44	5.13	6.69
19	3	2	0	53	2	18.06	-32.5	-1.31	5.13	6.69
19	3	2	0	54	2	18.06	-32.31	-0.37	5.06	6.63
19	3	2	0	55	3	18.06	-32.44	-1.06	5.06	6.63
19	3	2	0	56	3	18.06	-32.25	-0.31	5.06	6.63
19	3	2	0	57	3	18	-32.38	-1.06	5.06	6.56
19	3	2	0	58	3	18	-32.5	-1.94	5.06	6.56
19	3	2	0	59	4	18	-32.31	-1	5.06	6.56
19	3	2	1	0	4	18	-32.5	-1.94	5.06	6.5
19	3	2	1	1	4	18	-32.25	-1	5.06	6.5
19	3	2	1	2	4	18	-32.44	-1.69	5.06	6.5
19	3	2	1	3	5	18	-32.25	-0.94	5.06	6.44
19	3	2	1	4	5	18	-32.38	-1.56	5	6.44
19	3	2	1	5	5	18	-32.5	-2.56	5	6.44
19	3	2	1	6	5	18	-32.31	-1.37	5	6.38
19	3	2	1	7	6	18	-32.44	-2.25	5	6.38
19	3	2	1	8	6	18	-32.25	-1.37	5	6.31
19	3	2	1	9	6	18	-32.38	-2.06	5	6.31
19	3	2	1	10	7	18	-32.19	-1.31	5	6.31
19	3	2	1	11	7	18	-32.25	-1.75	5	6.31
19	3	2	1	12	7	17.94	-32.44	-2.69	5	6.31
19	3	2	1	13	7	17.94	-32.19	-1.63	5	6.25
19	3	2	1	14	8	17.94	-32.31	-2.5	5	6.25
19	3	2	1	15	8	17.94	-32.13	-1.56	4.94	6.25
19	3	2	1	16	8	17.94	-32.31	-2.31	4.94	6.25
19	3	2	1	17	8	17.94	-32.13	-1.94	4.94	6.25
19	3	2	1	18	9	18	-32.25	-2.06	4.94	6.19
19	3	2	1	19	9	18	-32.38	-2.81	4.94	6.19
19	3	2	1	20	9	18	-32.19	-1.81	4.94	6.19
19	3	2	1	21	9	18	-32.31	-2.63	4.94	6.19
19	3	2	1	22	10	18	-32.13	-1.75	4.94	6.19
19	3	2	1	23	10	18.06	-32.19	-2.44	4.94	6.13
19	3	2	1	24	10	18.06	-32.38	-3.31	4.94	6.13
19	3	2	1	25	10	18.06	-32.19	-2.19	4.88	6.13
19	3	2	1	26	11	18.12	-32.25	-2.81	4.88	6.13
19	3	2	1	27	11	18.12	-32.06	-1.94	4.94	6.13
19	3	2	1	28	11	18.12	-32.19	-2.5	4.88	6.13
19	3	2	1	29	11	18.12	-32.06	-1.69	4.88	6.13
19	3	2	1	30	12	18.19	-32.13	-2.25	4.88	6.06
19	3	2	1	31	12	18.19	-32.31	-2.94	4.88	6.13
19	3	2	1	32	12	18.19	-32.13	-1.94	4.88	6.06
19	3	2	1	33	12	18.25	-32.25	-2.63	4.88	6.13
19	3	2	1	34	13	18.25	-32.06	-1.75	4.88	6.13
19	3	2	1	35	17	18.25	-32.19	-2.31	4.88	6.13
19	3	2	1	36	17	18.25	-32.25	-2.75	4.88	6.13
19	3	2	1	37	17	18.19	-32.13	-1.88	4.88	6.06
19	3	2	1	38	17	18.19	-32.25	-2.5	4.88	6.13
19	3	2	1	39	18	18.19	-32.13	-1.63	4.88	6.13
19	3	2	1	40	18	18.19	-32.25	-2.19	4.88	6.13
19	3	2	1	41	18	18.12	-32.06	-1.44	4.88	6.13
19	3	2	1	42	18	18.12	-32.25	-1.94	4.88	6.13
19	3	2	1	43	19	18.12	-32.38	-2.5	4.88	6.13
19	3	2	1	44	19	18.12	-32.19	-1.56	4.88	6.13
19	3	2	1	45	19	18.12	-32.31	-2.19	4.88	6.13
19	3	2	1	46	19	18.12	-32.13	-1.44	4.88	6.13
19	3	2	1	47	20	18.12	-32.25	-1.81	4.88	6.13
19	3	2	1	48	20	18.06	-32.31	-2.25	4.88	6.19
19	3	2	1	49	20	18.06	-32.25	-1.44	4.88	6.19
19	3	2	1	50	20	18.12	-32.31	-1.94	4.88	6.19
19	3	2	1	51	21	18.06	-32.19	-1.19	4.88	6.19
19	3	2	1	52	21	18.06	-32.31	-1.69	4.88	6.19
19	3	2	1	53	21	18.06	-32.19	-1	4.88	6.19
19	3	2	1	54	21	18.06	-32.31	-1.44	4.88	6.19
19	3	2	1	55	22	18.06	-32.38	-1.88	4.88	6.25
19	3	2	1	56	22	18.06	-32.25	-1.06	4.88	6.25
19	3	2	1	57	22	18.06	-32.31	-1.44	4.88	6.25
19	3	2	1	58	22	18	-32.19	-0.81	4.88	6.25
19	3	2	1	59	23	18.06	-32.25	-1.12	4.88	6.25
19	3	2	2	0	23	18	-32.25	-1.5	4.88	6.31
19	3	2	2	1	23	17.94	-32.13	-0.69	4.88	6.31
19	3	2	2	2	24	17.94	-32.25	-1.12	4.88	6.31
19	3	2	2	3	24	17.94	-32.13	-0.5	4.88	6.31
19	3	2	2	4	24	17.87	-32.25	-0.88	4.88	6.31
19	3	2	2	5	24	17.87	-32.13	-0.63	4.88	6.38
19	3	2	2	6	25	17.87	-32.13	-0.44	4.88	6.38
19	3	2	2	7	25	17.81	-32.19	-0.69	4.88	6.38
19	3	2	2	8	25	17.81	-32.06	-0.12	4.88	6.44
19	3	2	2	9	25	17.81	-32.13	-0.44	4.88	6.44
19	3	2	2	10	26	17.81	-32	0.06	4.88	6.44
19	3	2	2	11	26	17.81	-32.13	-0.19	4.88	6.44
19	3	2	2	12	26	17.81	-32.13	-0.37	4.88	6.5
19	3	2	2	13	26	17.81	-32	0.19	4.88	6.5
19	3	2	2	14	27	17.81	-32.06	-0.12	4.88	6.5
19	3	2	2	15	27	17.81	-32	0.31	4.88	6.5
19	3	2	2	16	27	17.81	-32	0.19	4.94	6.56
19	3	2	2	17	27	17.81	-31.94	0.19	4.88	6.56
19	3	2	2	18	28	17.81	-31.87	0.69	4.94	6.56
19	3	2	2	19	28	17.81	-31.87	0.5	4.94	6.63
19	3	2	2	20	28	17.75	-31.81	0.88	4.94	6.63
19	3	2	2	21	28	17.75	-31.87	0.69	4.94	6.63
19	3	2	2	22	29	17.75	-31.87	0.56	4.94	6.69
19	3	2	2	23	29	17.75	-31.81	1.06	4.94	6.69
19	3	2	2	24	29	17.75	-31.87	0.81	4.94	6.75
19	3	2	2	25	30	17.75	-31.81	1.19	4.94	6.75
19	3	2	2	26	30	17.75	-31.75	1.19	4.94	6.75
19	3	2	2	27	30	17.75	-31.75	1.31	4.94	6.81
19	3	2	2	28	30	17.75	-31.69	1.63	4.94	6.81
19	3	2	2	29	31	17.81	-31.69	1.5	4.94	6.81
19	3	2	2	30	31	17.81	-31.69	1.81	4.94	6.88
19	3	2	2	31	31	17.81	-31.75	1.75	4.94	6.88
19	3	2	2	32	31	17.81	-31.75	1.81	4.94	6.88
19	3	2	2	33	32	17.81	-31.75	2.13	5	6.94
19	3	2	2	34	32	17.81	-31.75	1.94	5	6.94
19	3	2	2	35	32	17.81	-31.75	2.19	5	6.94
19	3	2	2	36	32	17.81	-31.69	2.31	5	7
19	3	2	2	37	33	17.81	-31.62	2.5	5	7
19	3	2	2	38	33	17.81	-31.62	2.63	5	7.06
19	3	2	2	39	33	17.81	-31.5	2.69	5	7.06
19	3	2	2	40	33	17.81	-31.56	2.81	5	7.13
19	3	2	2	41	34	17.75	-31.5	2.88	5	7.13
19	3	2	2	42	34	17.75	-31.5	3	5	7.13
19	3	2	2	43	34	17.75	-31.44	3.13	5.06	7.19
19	3	2	2	44	34	17.75	-31.37	3.13	5.06	7.19
19	3	2	2	45	35	17.75	-31.44	3.25	5.06	7.25
19	3	2	2	46	35	17.75	-31.44	3.25	5.06	7.25
19	3	2	2	47	35	17.75	-31.44	3.19	5.06	7.25
19	3	2	2	48	35	17.75	-31.44	3.38	5.06	7.31
19	3	2	2	49	36	17.75	-31.44	3.38	5.06	7.31
19	3	2	2	50	36	17.69	-31.5	3.5	5.06	7.31
19	3	2	2	51	36	17.69	-31.5	3.5	5.06	7.31
19	3	2	2	52	36	17.69	-31.5	3.5	5.06	7.38
19	3	2	2	53	37	17.69	-31.5	3.69	5.06	7.38
19	3	2	2	54	37	17.62	-31.5	3.63	5.06	7.38
19	3	2	2	55	37	17.62	-31.5	3.81	5.13	7.38
19	3	2	2	56	38	17.62	-31.5	3.56	5.06	7.44
19	3	2	2	57	38	17.69	-31.5	3.38	5.13	7.44
19	3	2	2	58	38	17.69	-31.56	3.5	5.13	7.44
19	3	2	2	59	38	17.69	-31.56	3.31	5.13	7.44
19	3	2	3	0	39	17.62	-31.62	3.56	5.13	7.44
19	3	2	3	1	39	17.62	-31.62	3.38	5.13	7.38
19	3	2	3	2	39	17.62	-31.62	3.56	5.13	7.38
19	3	2	3	3	39	17.62	-31.81	3	5.13	7.38
19	3	2	3	4	40	17.62	-31.81	2.81	5.13	7.38
19	3	2	3	5	40	17.62	-31.75	3.13	5.13	7.38
19	3	2	3	6	40	17.62	-31.81	2.81	5.13	7.31
19	3	2	3	7	40	17.62	-31.81	3.13	5.06	7.31
19	3	2	3	8	41	17.62	-31.81	2.81	5.06	7.31
19	3	2	3	9	41	17.62	-31.81	2.44	5.06	7.31
19	3	2	3	10	41	17.62	-31.75	2.81	5.06	7.25
19	3	2	3	11	41	17.69	-31.81	2.44	5.06	7.25
19	3	2	3	12	42	17.69	-31.81	2.81	5.06	7.25
19	3	2	3	13	42	17.69	-31.87	2.5	5.06	7.25
19	3	2	3	14	42	17.69	-31.94	2.13	5.06	7.25
19	3	2	3	15	42	17.69	-31.87	2.56	5.06	7.19
19	3	2	3	16	43	17.69	-32	2.06	5.06	7.19
19	3	2	3	17	43	17.69	-31.94	2.5	5.06	7.19
19	3	2	3	18	43	17.69	-32.06	2.06	5	7.19
19	3	2	3	19	44	17.62	-32.19	1.44	5.06	7.13
19	3	2	3	20	44	17.69	-32.06	2.06	5	7.13
19	3	2	3	21	44	17.62	-32.19	1.63	5	7.06
19	3	2	3	22	44	17.62	-32.13	2.06	5	7.06
19	3	2	3	23	45	17.62	-32.19	1.69	5	7.06
19	3	2	3	24	45	17.62	-32.13	1.63	5	7.06
19	3	2	3	25	45	17.69	-32.19	1.75	5	7
19	3	2	3	26	45	17.69	-32.31	1.06	5	7
19	3	2	3	27	46	17.69	-32.25	1.63	5	7
19	3	2	3	28	46	17.69	-32.31	1.06	5	7
19	3	2	3	29	46	17.69	-32.25	1.56	5	6.94
19	3	2	3	30	46	17.69	-32.31	1.06	5	6.94
19	3	2	3	31	47	17.69	-32.44	0.44	5	6.88
19	3	2	3	32	47	17.62	-32.31	1.12	5	6.88
19	3	2	3	33	47	17.62	-32.44	0.56	5	6.88
19	3	2	3	34	47	17.62	-32.38	1.19	5	6.81
19	3	2	3	35	48	17.62	-32.38	0.69	4.94	6.81
19	3	2	3	36	48	17.56	-32.56	-0.19	5	6.81
19	3	2	3	37	48	17.56	-32.38	0.69	4.94	6.81
19	3	2	3	38	48	17.56	-32.5	-0.06	4.94	6.81
19	3	2	3	39	49	17.56	-32.38	0.63	4.94	6.75
19	3	2	3	40	49	17.5	-32.5	0.06	4.94	6.75
19	3	2	3	41	49	17.5	-32.38	0.63	4.94	6.69
19	3	2	3	42	49	17.5	-32.31	0.56	4.94	6.75
19	3	2	3	43	50	17.5	-32.44	-0.19	4.94	6.75
19	3	2	3	44	50	17.5	-32.31	0.56	4.94	6.69
19	3	2	3	45	50	17.5	-32.44	-0.06	4.94	6.69
19	3	2	3	46	51	17.5	-32.31	0.5	4.94	6.69
19	3	2	3	47	51	17.5	-32.44	0	4.94	6.63
19	3	2	3	48	51	17.5	-32.56	-0.69	4.94	6.63
19	3	2	3	49	51	17.5	-32.38	0.12	4.88	6.63
19	3	2	3	50	52	17.5	-32.5	-0.5	4.94	6.63
19	3	2	3	51	52	17.5	-32.38	0.19	4.88	6.63
19	3	2	3	52	52	17.5	-32.5	-0.37	4.88	6.56
19	3	2	3	53	52	17.5	-32.38	-0.06	4.88	6.56
19	3	2	3	54	53	17.5	-32.5	-0.5	4.88	6.56
19	3	2	3	55	53	17.5	-32.5	-0.75	4.88	6.56
19	3	2	3	56	53	17.5	-32.31	0	4.88	6.56
19	3	2	3	57	53	17.5	-32.38	-0.44	4.88	6.5
19	3	2	3	58	54	17.5	-32.25	0.19	4.88	6.5
19	3	2	3	59	54	17.5	-32.31	-0.19	4.88	6.5
19	3	2	4	0	54	17.5	-32.44	-0.56	4.88	6.5
19	3	2	4	1	54	17.56	-32.25	0.19	4.88	6.5
19	3	2	4	2	55	17.56	-32.38	-0.25	4.88	6.5
19	3	2	4	3	55	17.56	-32.25	0.31	4.88	6.5
19	3	2	4	4	55	17.62	-32.38	-0.12	4.88	6.5
19	3	2	4	5	55	17.62	-32.38	-0.37	4.88	6.5
19	3	2	4	6	56	17.62	-32.25	0.31	4.88	6.56
19	3	2	4	7	56	17.62	-32.25	0	4.88	6.56
19	3	2	4	8	56	17.69	-32.19	0.56	4.88	6.56
19	3	2	4	9	56	17.69	-32.19	0.25	4.88	6.56
19	3	2	4	10	57	17.69	-32.06	0.44	4.88	6.56
19	3	2	4	11	57	17.69	-32.13	0.56	4.88	6.56
19	3	2	4	12	57	17.75	-32.19	0.25	4.88	6.63
19	3	2	4	13	57	17.75	-32.06	0.75	4.88	6.63
19	3	2	4	14	58	17.75	-32.13	0.44	4.88	6.63
19	3	2	4	15	58	17.75	-32	0.88	4.88	6.63
19	3	2	4	16	58	17.81	-32.06	0.75	4.88	6.63
19	3	2	4	17	59	17.81	-32	0.75	4.88	6.69
19	3	2	4	18	59	17.81	-31.94	1.19	4.88	6.69
19	3	2	4	19	59	17.81	-31.94	0.94	4.88	6.69
19	3	2	4	20	59	17.81	-31.87	1.31	4.88	6.75
19	3	2	4	22	0	17.81	-31.94	1.19	4.88	6.75
19	3	2	4	23	0	17.81	-31.94	1.06	4.88	6.75
19	3	2	4	24	0	17.75	-31.87	1.44	4.88	6.75
19	3	2	4	25	0	17.75	-31.87	1.25	4.88	6.81
19	3	2	4	26	1	17.75	-31.87	1.63	4.88	6.81
19	3	2	4	27	1	17.69	-31.81	1.56	4.88	6.81
19	3	2	4	28	1	17.69	-31.75	1.69	4.88	6.88
19	3	2	4	29	1	17.69	-31.69	2	4.88	6.88
19	3	2	4	30	2	17.62	-31.75	1.88	4.94	6.88
19	3	2	4	31	2	17.62	-31.75	2.13	4.94	6.94
19	3	2	4	32	2	17.62	-31.81	2.06	4.94	6.94
19	3	2	4	33	2	17.56	-31.75	2.13	4.94	6.94
19	3	2	4	34	3	17.56	-31.75	2.38	4.94	7
19	3	2	4	35	3	17.56	-31.81	2.25	4.94	7
19	3	2	4	36	3	17.5	-31.81	2.5	4.94	7
19	3	2	4	37	3	17.5	-31.75	2.56	4.94	7.06
19	3	2	4	38	4	17.5	-31.75	2.75	4.94	7.06
19	3	2	4	39	4	17.44	-31.69	2.88	4.94	7.13
19	3	2	4	40	4	17.44	-31.62	2.88	4.94	7.13
19	3	2	4	41	5	17.37	-31.56	3.06	4.94	7.13
19	3	2	4	42	5	17.37	-31.56	3.06	5	7.19
19	3	2	4	43	5	17.37	-31.5	3.25	5	7.19
19	3	2	4	44	5	17.37	-31.44	3.31	5	7.25
19	3	2	4	45	6	17.37	-31.44	3.31	5	7.25
19	3	2	4	46	6	17.37	-31.44	3.44	5	7.25
19	3	2	4	47	6	17.37	-31.37	3.56	5	7.31
19	3	2	4	48	6	17.44	-31.31	3.63	5	7.31
19	3	2	4	49	7	17.44	-31.25	3.69	5	7.31
19	3	2	4	50	7	17.5	-31.19	3.75	5	7.38
19	3	2	4	51	7	17.5	-31.12	3.88	5	7.38
19	3	2	4	52	7	17.5	-31.06	3.88	5	7.38
19	3	2	4	53	8	17.56	-31.06	4	5	7.38
19	3	2	4	54	8	17.56	-31	4	5	7.44
19	3	2	4	55	8	17.62	-31	4.06	5.06	7.44
19	3	2	4	56	8	17.62	-30.94	4.13	5.06	7.44
19	3	2	4	57	9	17.62	-31	4.13	5.06	7.44
19	3	2	4	58	9	17.62	-31.06	4.19	5.06	7.5
19	3	2	4	59	9	17.69	-31.06	4.25	5.06	7.5
19	3	2	5	0	9	17.69	-31.06	4.19	5.06	7.5
19	3	2	5	1	10	17.69	-31.12	4.25	5.06	7.5
19	3	2	5	2	10	17.69	-31.19	4.25	5.06	7.5
19	3	2	5	3	10	17.75	-31.19	4.31	5.06	7.5
19	3	2	5	4	10	17.75	-31.12	4.38	5.06	7.56
19	3	2	5	5	15	17.75	-31.19	4.38	5.06	7.56
19	3	2	5	6	15	17.81	-31.19	4.44	5.06	7.56
19	3	2	5	7	15	17.81	-31.25	4.25	5.06	7.56
19	3	2	5	8	16	17.81	-31.25	4.19	5.06	7.56
19	3	2	5	9	16	17.81	-31.31	4.38	5.06	7.56
19	3	2	5	10	16	17.81	-31.25	4.13	5.06	7.56
19	3	2	5	11	16	17.87	-31.25	4.25	5.06	7.56
19	3	2	5	12	17	17.87	-31.31	4.06	5.06	7.56
19	3	2	5	13	17	17.87	-31.37	4.13	5.06	7.56
19	3	2	5	14	17	17.87	-31.37	4.31	5.13	7.56
19	3	2	5	15	17	17.87	-31.37	4.06	5.06	7.56
19	3	2	5	16	18	17.94	-31.44	4.25	5.13	7.56
19	3	2	5	17	18	17.94	-31.5	3.88	5.06	7.56
19	3	2	5	18	18	17.94	-31.5	3.63	5.06	7.5
19	3	2	5	19	18	17.94	-31.56	3.94	5.06	7.5
19	3	2	5	20	19	17.94	-31.62	3.56	5.06	7.5
19	3	2	5	21	19	18	-31.62	3.81	5.06	7.5
19	3	2	5	22	19	18	-31.62	3.44	5.06	7.44
19	3	2	5	23	19	18.06	-31.62	3.19	5.06	7.44
19	3	2	5	24	20	18.06	-31.69	3.56	5.06	7.44
19	3	2	5	25	20	18.06	-31.75	3.25	5.06	7.44
19	3	2	5	26	20	18.06	-31.69	3.5	5.06	7.44
19	3	2	5	27	20	18.06	-31.75	3.06	5.06	7.38
19	3	2	5	28	21	18.06	-31.81	2.69	5.06	7.38
19	3	2	5	29	21	18.06	-31.81	3.13	5.06	7.31
19	3	2	5	30	21	18.06	-31.87	2.56	5.06	7.31
19	3	2	5	31	21	18.06	-31.87	3	5.06	7.31
19	3	2	5	32	22	18	-31.94	2.56	5	7.31
19	3	2	5	33	22	18	-31.87	2.88	5	7.25
19	3	2	5	34	22	18	-31.81	3.06	5.06	7.25
19	3	2	5	35	23	17.94	-31.81	2.69	5.06	7.25
19	3	2	5	36	23	17.94	-31.75	3.06	5	7.25
19	3	2	5	37	23	17.94	-31.94	2.38	5	7.25
19	3	2	5	38	23	17.87	-31.87	2.81	5	7.25
19	3	2	5	39	24	17.87	-32.06	2.19	5	7.19
19	3	2	5	40	24	17.81	-32.19	1.63	5	7.19
19	3	2	5	41	24	17.81	-32.06	2.19	5	7.19
19	3	2	5	42	24	17.75	-32.19	1.63	5	7.13
19	3	2	5	43	25	17.75	-32.13	2.13	5	7.13
19	3	2	5	44	25	17.75	-32.25	1.56	5	7.06
19	3	2	5	45	25	17.69	-32.38	0.81	5	7.06
19	3	2	5	46	25	17.69	-32.25	1.56	5	7.06
19	3	2	5	47	26	17.62	-32.38	0.81	5	7
19	3	2	5	48	26	17.62	-32.25	1.44	4.94	7
19	3	2	5	49	26	17.56	-32.38	0.75	4.94	6.94
19	3	2	5	50	26	17.56	-32.25	1.31	4.94	6.94
19	3	2	5	51	27	17.56	-32.31	1.06	4.94	6.94
19	3	2	5	52	27	17.5	-32.38	0.5	4.94	6.94
19	3	2	5	53	27	17.5	-32.25	1.19	4.94	6.88
19	3	2	5	54	27	17.5	-32.38	0.5	4.94	6.88
19	3	2	5	55	28	17.44	-32.25	1.06	4.94	6.88
19	3	2	5	56	28	17.44	-32.31	0.69	4.94	6.81
19	3	2	5	57	28	17.37	-32.38	0	4.94	6.81
19	3	2	5	58	29	17.37	-32.25	0.81	4.88	6.81
19	3	2	5	59	29	17.37	-32.31	0.31	4.88	6.81
19	3	2	6	0	29	17.37	-32.19	0.94	4.94	6.75
19	3	2	6	1	29	17.37	-32.25	0.5	4.94	6.75
19	3	2	6	2	30	17.44	-32.19	0.37	4.88	6.75
19	3	2	6	3	30	17.44	-32.19	0.63	4.88	6.75
19	3	2	6	4	30	17.44	-32.31	0.06	4.88	6.75
19	3	2	6	5	30	17.5	-32.19	0.75	4.88	6.75
19	3	2	6	6	31	17.5	-32.19	0.5	4.88	6.75
19	3	2	6	7	31	17.5	-32.06	1	4.88	6.75
19	3	2	6	8	31	17.5	-32.13	0.75	4.88	6.75
19	3	2	6	9	31	17.5	-32.19	0.37	4.88	6.75
19	3	2	6	10	32	17.5	-32.06	0.94	4.88	6.75
19	3	2	6	11	32	17.5	-32.13	0.63	4.88	6.75
19	3	2	6	12	32	17.5	-32	1.12	4.88	6.75
19	3	2	6	13	32	17.5	-32.06	0.81	4.88	6.75
19	3	2	6	14	33	17.5	-32.19	0.37	4.88	6.75
19	3	2	6	15	33	17.44	-32	1	4.88	6.75
19	3	2	6	16	33	17.44	-32	0.81	4.88	6.75
19	3	2	6	17	33	17.44	-31.87	1.25	4.88	6.75
19	3	2	6	18	34	17.44	-31.94	1.12	4.88	6.81
19	3	2	6	19	34	17.37	-31.94	0.75	4.88	6.81
19	3	2	6	20	34	17.44	-31.81	1.37	4.88	6.81
19	3	2	6	21	34	17.44	-31.87	1.12	4.88	6.81
19	3	2	6	22	35	17.44	-31.81	1.56	4.88	6.81
19	3	2	6	23	35	17.44	-31.81	1.31	4.88	6.81
19	3	2	6	24	35	17.37	-31.81	1.19	4.88	6.88
19	3	2	6	25	36	17.37	-31.75	1.56	4.88	6.88
19	3	2	6	26	36	17.37	-31.75	1.44	4.88	6.88
19	3	2	6	27	36	17.37	-31.75	1.81	4.88	6.88
19	3	2	6	28	36	17.37	-31.75	1.75	4.88	6.94
19	3	2	6	29	37	17.31	-31.75	2.06	4.94	6.94
19	3	2	6	30	37	17.31	-31.87	1.63	4.94	6.94
19	3	2	6	31	37	17.31	-31.87	1.56	4.94	6.94
19	3	2	6	32	37	17.31	-31.81	1.94	4.88	6.94
19	3	2	6	33	38	17.31	-31.87	1.81	4.94	7
19	3	2	6	34	38	17.31	-31.81	2.13	4.94	7
19	3	2	6	35	38	17.31	-31.81	2.06	4.94	7
19	3	2	6	36	38	17.31	-31.75	2.19	4.94	7
19	3	2	6	37	39	17.37	-31.75	2.44	4.94	7.06
19	3	2	6	38	39	17.37	-31.62	2.56	4.94	7.06
19	3	2	6	39	39	17.44	-31.62	2.75	4.94	7.13
19	3	2	6	40	39	17.44	-31.56	2.81	4.94	7.13
19	3	2	6	41	40	17.44	-31.56	2.88	4.94	7.13
19	3	2	6	42	40	17.5	-31.56	3	4.94	7.19
19	3	2	6	43	40	17.5	-31.5	3.06	4.94	7.19
19	3	2	6	44	40	17.56	-31.44	3.25	5	7.25
19	3	2	6	45	41	17.56	-31.44	3.25	5	7.25
19	3	2	6	46	41	17.62	-31.44	3.38	5	7.25
19	3	2	6	47	41	17.62	-31.37	3.5	5	7.31
19	3	2	6	48	41	17.69	-31.31	3.63	5	7.31
19	3	2	6	49	42	17.69	-31.25	3.69	5	7.31
19	3	2	6	50	42	17.69	-31.19	3.75	5	7.38
19	3	2	6	51	42	17.75	-31.12	3.81	5	7.38
19	3	2	6	52	42	17.75	-31.06	3.88	5	7.38
19	3	2	6	53	43	17.75	-31	3.94	5	7.38
19	3	2	6	54	43	17.81	-30.94	4	5	7.44
19	3	2	6	55	43	17.81	-30.87	4.06	5	7.44
19	3	2	6	56	44	17.81	-30.87	4.13	5	7.44
19	3	2	6	57	44	17.87	-30.87	4.19	5.06	7.44
19	3	2	6	58	44	17.87	-30.81	4.25	5	7.5
19	3	2	6	59	44	17.87	-30.75	4.25	5.06	7.5
19	3	2	7	0	45	17.87	-30.75	4.31	5.06	7.5
19	3	2	7	1	45	17.94	-30.75	4.38	5.06	7.5
19	3	2	7	2	45	17.94	-30.75	4.38	5.06	7.5
19	3	2	7	3	45	17.94	-30.69	4.44	5.06	7.56
19	3	2	7	4	46	17.94	-30.69	4.5	5.06	7.56
19	3	2	7	5	46	18	-30.69	4.56	5.06	7.56
19	3	2	7	6	46	18	-30.69	4.56	5.06	7.56
19	3	2	7	7	46	18	-30.69	4.63	5.06	7.56
19	3	2	7	8	47	17.94	-30.69	4.56	5.06	7.56
19	3	2	7	9	47	17.94	-30.75	4.63	5.06	7.63
19	3	2	7	10	47	17.94	-30.81	4.69	5.06	7.63
19	3	2	7	11	47	17.87	-30.75	4.69	5.06	7.63
19	3	2	7	12	48	17.87	-30.75	4.75	5.06	7.63
19	3	2	7	13	48	17.87	-30.81	4.69	5.06	7.63
19	3	2	7	14	48	17.81	-30.81	4.75	5.06	7.63
19	3	2	7	15	48	17.81	-30.81	4.81	5.06	7.63
19	3	2	7	16	49	17.81	-30.81	4.75	5.06	7.63
19	3	2	7	17	49	17.81	-30.87	4.81	5.06	7.69
19	3	2	7	18	49	17.75	-30.87	4.56	5.06	7.69
19	3	2	7	19	50	17.75	-30.94	4.69	5.13	7.69
19	3	2	7	20	50	17.75	-31	4.5	5.06	7.69
19	3	2	7	21	50	17.75	-31.12	4.5	5.13	7.63
19	3	2	7	22	50	17.75	-31.12	4.63	5.13	7.63
19	3	2	7	23	51	17.69	-31.12	4.44	5.06	7.63
19	3	2	7	24	51	17.69	-31.19	4.56	5.06	7.63
19	3	2	7	25	51	17.69	-31.25	4.38	5.06	7.63
19	3	2	7	26	51	17.69	-31.25	4.13	5.13	7.63
19	3	2	7	27	52	17.69	-31.25	4.38	5.13	7.63
19	3	2	7	28	52	17.69	-31.37	4	5.06	7.63
19	3	2	7	29	52	17.69	-31.31	4.19	5.13	7.56
19	3	2	7	30	52	17.69	-31.37	3.81	5.13	7.56
19	3	2	7	31	53	17.69	-31.37	3.63	5.13	7.56
19	3	2	7	32	53	17.62	-31.44	3.94	5.06	7.56
19	3	2	7	33	53	17.62	-31.44	3.56	5.06	7.5
19	3	2	7	34	53	17.62	-31.44	3.81	5.06	7.5
19	3	2	7	35	54	17.62	-31.5	3.5	5.06	7.5
19	3	2	7	36	54	17.62	-31.56	3	5.06	7.5
19	3	2	7	37	54	17.62	-31.5	3.44	5.06	7.44
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	3	2	7	43	23	17.5	-31.31	3.94	5.06	85
19	3	2	7	44	24	17.5	-31.19	4.13	5.06	85
19	3	2	7	45	24	17.5	-31.12	4.06	5.06	85
19	3	2	7	46	24	17.5	-31.12	4.25	5.06	85
19	3	2	7	47	24	17.5	-31.12	4.25	5.06	85
19	3	2	7	48	25	17.56	-31.19	4.38	5.06	85
19	3	2	7	49	25	17.56	-31.19	4.38	5.06	85
19	3	2	7	50	25	17.56	-31.19	4.38	5.06	85
19	3	2	7	51	25	17.62	-31.19	4.5	5.06	85
19	3	2	7	52	26	17.62	-31.19	4.44	5.06	85
19	3	2	7	53	26	17.62	-31.19	4.56	5.13	85
19	3	2	7	54	26	17.62	-31.19	4.56	5.13	85
19	3	2	7	55	26	17.69	-31.19	4.56	5.13	85
19	3	2	7	56	27	17.62	-31.19	4.69	5.13	85
19	3	2	7	57	27	17.62	-31.19	4.63	5.13	85
19	3	2	7	58	27	17.56	-31.12	4.75	5.13	85
19	3	2	7	59	27	17.62	-31.12	4.75	5.13	85
19	3	2	8	0	28	17.62	-31.12	4.81	5.13	85
19	3	2	8	1	28	17.56	-31.06	4.88	5.13	85
19	3	2	8	2	28	17.62	-31.06	4.81	5.13	85
19	3	2	8	3	28	17.56	-31.06	4.88	5.13	85
19	3	2	8	4	29	17.56	-31.06	4.88	5.13	85
19	3	2	8	5	29	17.56	-31	4.94	5.13	85
19	3	2	8	6	29	17.56	-31	5	5.13	85
19	3	2	8	7	29	17.56	-30.94	4.94	5.13	85
19	3	2	8	8	30	17.56	-30.94	5	5.13	85
19	3	2	8	9	30	17.56	-31.19	3.75	5.13	85
19	3	2	8	10	30	17.56	-31.19	3.81	5.13	85
19	3	2	8	11	31	17.56	-31.25	4.13	5.13	85
19	3	2	8	12	31	17.56	-31.25	4.19	5.13	85
19	3	2	8	13	31	17.56	-31.25	4.38	5.13	85
19	3	2	8	14	31	17.56	-31.25	4.44	5.13	85
19	3	2	8	15	32	17.56	-31.25	4.44	5.13	85
19	3	2	8	16	32	17.56	-31.25	4.56	5.13	85
19	3	2	8	17	32	17.5	-31.25	4.56	5.13	85
19	3	2	8	18	36	17.5	-31.25	4.69	5.13	85
19	3	2	8	19	37	17.5	-31.25	4.69	5.13	85
19	3	2	8	20	37	17.5	-31.19	4.69	5.13	85
19	3	2	8	21	37	17.44	-31.19	4.81	5.13	85
19	3	2	8	22	37	17.44	-31.19	4.81	5.13	85
19	3	2	8	23	38	17.44	-31.19	4.88	5.13	85
19	3	2	8	24	38	17.44	-31.19	4.88	5.13	85
19	3	2	8	25	38	17.44	-31.19	4.88	5.13	85
19	3	2	8	26	38	17.5	-31.19	5	5.13	85
19	3	2	8	27	39	17.44	-31.12	4.94	5.19	85
19	3	2	8	28	39	17.44	-31.12	5	5.19	85
19	3	2	8	29	39	17.44	-31.12	4.75	5.19	85
19	3	2	8	30	39	17.44	-31.19	4.5	5.19	85
19	3	2	8	31	40	17.44	-31.19	4.69	5.13	85
19	3	2	8	32	40	17.44	-31.19	4.44	5.19	85
19	3	2	8	33	40	17.44	-31.19	4.63	5.13	85
19	3	2	8	34	40	17.44	-31.31	4.38	5.19	85
19	3	2	8	35	41	17.5	-31.31	4.13	5.19	85
19	3	2	8	36	41	17.5	-31.31	4.38	5.19	85
19	3	2	8	37	41	17.5	-31.37	4.19	5.19	85
19	3	2	8	38	42	17.56	-31.37	4.38	5.19	85
19	3	2	8	39	42	17.56	-31.44	4.06	5.19	85
19	3	2	8	40	42	17.62	-31.5	3.56	5.19	85
19	3	2	8	41	42	17.62	-31.5	3.94	5.13	85
19	3	2	8	42	43	17.69	-31.56	3.56	5.13	85
19	3	2	8	43	43	17.69	-31.5	3.88	5.13	85
19	3	2	8	44	43	17.75	-31.5	3.56	5.13	85
19	3	2	8	45	43	17.75	-31.56	3.06	5.13	85
19	3	2	8	46	44	17.75	-31.56	3.5	5.13	85
19	3	2	8	47	44	17.69	-31.56	3.19	5.13	85
19	3	2	8	48	44	17.69	-31.5	3.5	5.13	85
19	3	2	8	49	44	17.69	-31.62	3.06	5.13	85
19	3	2	8	50	45	17.69	-31.62	2.75	5.13	85
19	3	2	8	51	45	17.69	-31.75	2.94	5.13	85
19	3	2	8	52	45	17.69	-31.81	2.44	5.13	85
19	3	2	8	53	45	17.62	-31.75	2.94	5.06	85
19	3	2	8	54	46	17.62	-31.81	2.44	5.06	85
19	3	2	8	55	46	17.62	-31.75	2.81	5.06	85
19	3	2	8	56	46	17.62	-31.81	2.44	5.06	85
19	3	2	8	57	46	17.56	-31.94	1.94	5.06	85
19	3	2	8	58	47	17.56	-31.81	2.5	5.06	85
19	3	2	8	59	47	17.56	-31.94	1.94	5.06	85
19	3	2	9	0	47	17.5	-31.87	2.38	5.06	85
19	3	2	9	1	47	17.5	-31.94	1.88	5.06	85
19	3	2	9	2	48	17.5	-32.13	1.06	5.06	85
19	3	2	9	3	48	17.5	-32	1.81	5.06	85
19	3	2	9	4	48	17.5	-32.06	1.25	5.06	85
19	3	2	9	5	49	17.5	-31.94	1.81	5	85
19	3	2	9	6	49	17.56	-32.06	1.31	5	85
19	3	2	9	7	49	17.56	-32.13	0.81	5	85
19	3	2	9	8	49	17.62	-32.06	1.37	5	85
19	3	2	9	9	50	17.62	-32.25	0.63	5	85
19	3	2	9	10	50	17.62	-32.13	1.31	5	85
19	3	2	9	11	50	17.62	-32.25	0.63	5	85
19	3	2	9	12	50	17.62	-32.06	1.25	5	85
19	3	2	9	13	51	17.69	-32.25	0.69	5	85
19	3	2	9	14	51	17.69	-32.38	-0.12	5	85
19	3	2	9	15	51	17.62	-32.19	0.69	5	85
19	3	2	9	16	51	17.62	-32.31	0.06	5	85
19	3	2	9	17	52	17.62	-32.19	0.75	4.94	85
19	3	2	9	18	52	17.62	-32.31	0.19	4.94	85
19	3	2	9	19	52	17.62	-32.38	-0.44	4.94	85
19	3	2	9	20	52	17.62	-32.19	0.44	4.94	85
19	3	2	9	21	53	17.62	-32.25	-0.25	4.94	85
19	3	2	9	22	53	17.62	-32.13	0.5	4.94	85
19	3	2	9	23	53	17.62	-32.25	-0.06	4.94	85
19	3	2	9	24	53	17.62	-32.06	0.56	4.94	85
19	3	2	9	25	54	17.62	-32.19	0.19	4.94	85
19	3	2	9	26	54	17.62	-32.31	-0.37	4.94	85
19	3	2	9	27	54	17.62	-32.13	0.37	4.94	85
19	3	2	9	28	55	17.62	-32.25	-0.19	4.94	85
19	3	2	9	29	55	17.62	-32.06	0.44	4.94	85
19	3	2	9	30	55	17.62	-32.19	0.06	4.94	85
19	3	2	9	31	55	17.62	-32.25	-0.37	4.94	85
19	3	2	9	32	56	17.56	-32.06	0.37	4.94	85
19	3	2	9	33	56	17.56	-32.13	-0.12	4.94	85
19	3	2	9	34	56	17.56	-32	0.5	4.94	85
19	3	2	9	35	56	17.56	-32.06	0.06	4.88	85
19	3	2	9	36	57	17.62	-32	0	4.94	85
19	3	2	9	37	57	17.62	-32	0.37	4.88	85
19	3	2	9	38	57	17.62	-31.87	0.44	4.94	85
19	3	2	9	39	57	17.62	-31.81	0.88	4.94	85
19	3	2	9	40	58	17.56	-31.87	0.63	4.94	85
19	3	2	9	41	58	17.56	-31.75	1.06	4.94	85
19	3	2	9	42	58	17.56	-31.81	0.75	4.88	85
19	3	2	9	43	58	17.56	-31.87	0.5	4.94	85
19	3	2	9	44	59	17.56	-31.75	1	4.88	85
19	3	2	9	45	59	17.5	-31.81	0.63	4.94	85
19	3	2	9	46	59	17.5	-31.69	1.12	4.94	85
19	3	2	9	47	59	17.5	-31.81	0.81	4.94	85
19	3	2	9	49	0	17.5	-31.75	0.75	4.88	85
19	3	2	9	50	0	17.5	-31.69	1.25	4.94	85
19	3	2	9	51	0	17.5	-31.69	1	4.94	85
19	3	2	9	52	0	17.5	-31.62	1.44	4.94	85
19	3	2	9	53	1	17.56	-31.69	1.19	4.94	85
19	3	2	9	54	1	17.56	-31.62	1.12	4.94	85
19	3	2	9	55	1	17.62	-31.56	1.56	4.94	85
19	3	2	9	56	1	17.62	-31.62	1.31	4.94	85
19	3	2	9	57	2	17.62	-31.56	1.69	4.94	85
19	3	2	9	58	2	17.62	-31.56	1.5	4.94	85
19	3	2	9	59	2	17.62	-31.5	1.63	4.94	85
19	3	2	10	0	3	17.62	-31.44	2	4.94	85
19	3	2	10	1	3	17.62	-31.37	1.88	4.94	85
19	3	2	10	2	3	17.62	-31.44	2.19	4.94	85
19	3	2	10	3	3	17.62	-31.44	2.06	4.94	85
19	3	2	10	4	4	17.62	-31.44	2.13	4.94	85
19	3	2	10	5	4	17.62	-31.37	2.44	4.94	85
19	3	2	10	6	4	17.62	-31.37	2.31	5	85
19	3	2	10	7	4	17.62	-31.37	2.56	5	85
19	3	2	10	8	5	17.69	-31.37	2.38	5	85
19	3	2	10	9	5	17.69	-31.37	2.63	5	85
19	3	2	10	10	5	17.69	-31.31	2.81	5	85
19	3	2	10	11	5	17.69	-31.19	2.81	5	85
19	3	2	10	12	6	17.75	-31.19	3	5	85
19	3	2	10	13	6	17.75	-31.19	3.06	5	85
19	3	2	10	14	6	17.75	-31.12	3.25	5	85
19	3	2	10	15	6	17.75	-31.12	3.31	5	85
19	3	2	10	16	7	17.75	-31	3.31	5	85
19	3	2	10	17	7	17.75	-31	3.44	5	85
19	3	2	10	18	7	17.75	-31	3.5	5	85
19	3	2	10	19	7	17.75	-31	3.63	5.06	85
19	3	2	10	20	8	17.75	-31	3.69	5.06	85
19	3	2	10	21	8	17.75	-30.94	3.81	5.06	85
19	3	2	10	22	8	17.75	-30.87	3.88	5.06	85
19	3	2	10	23	9	17.75	-30.81	3.94	5.06	85
19	3	2	10	24	9	17.75	-30.75	4	5.06	85
19	3	2	10	25	9	17.75	-30.69	4.06	5.06	85
19	3	2	10	26	9	17.81	-30.69	4.13	5.06	85
19	3	2	10	27	10	17.87	-30.62	4.19	5.06	85
19	3	2	10	28	10	17.87	-30.62	4.25	5.06	85
19	3	2	10	29	10	17.87	-30.56	4.31	5.06	85
19	3	2	10	30	10	17.94	-30.56	4.31	5.13	85
19	3	2	10	31	11	17.94	-30.62	4.25	5.13	85
19	3	2	10	32	11	17.94	-30.69	4.38	5.13	85
19	3	2	10	33	11	17.94	-30.75	4.31	5.06	85
19	3	2	10	34	11	18	-30.75	4.38	5.13	85
19	3	2	10	35	12	18	-30.75	4.5	5.13	85
19	3	2	10	36	12	18	-30.75	4.44	5.13	85
19	3	2	10	37	12	18.06	-30.75	4.5	5.13	85
19	3	2	10	38	12	18.06	-30.81	4.5	5.13	85
19	3	2	10	39	13	18.12	-30.81	4.44	5.13	85
19	3	2	10	40	13	18.12	-30.87	4.56	5.13	85
19	3	2	10	41	13	18.25	-30.87	4.31	5.13	85
19	3	2	10	42	13	16.37	-31.31	4.44	4.94	8
19	3	2	10	43	14	18.37	-30.94	4.25	5.13	7.69
19	3	2	10	44	14	18.44	-30.94	4.25	5.13	7.69
19	3	2	10	45	14	18.5	-30.94	4.44	5.13	7.69
19	3	2	10	46	14	18.5	-30.94	4.19	5.13	7.69
19	3	2	10	47	15	18.5	-31	4.31	5.13	7.69
19	3	2	10	48	15	18.5	-31	4.13	5.13	7.69
19	3	2	10	49	15	18.5	-31	4.13	5.13	7.69
19	3	2	10	50	16	18.5	-31	4.19	5.13	7.69
19	3	2	10	51	16	18.5	-31.12	3.75	5.13	7.69
19	3	2	10	52	16	18.5	-31.12	4	5.13	7.69
19	3	2	10	53	16	18.56	-31.25	3.63	5.13	7.63
19	3	2	10	54	17	18.5	-31.25	3.94	5.13	7.63
19	3	2	10	55	17	18.56	-31.25	4	5.13	7.63
19	3	2	10	56	17	18.56	-31.25	3.75	5.13	7.63
19	3	2	10	57	17	18.56	-31.25	4	5.13	7.63
19	3	2	10	58	18	18.56	-31.31	3.63	5.13	7.63
19	3	2	10	59	18	18.56	-31.31	3.88	5.13	7.63
19	3	2	11	0	18	18.56	-31.44	3.38	5.13	7.56
19	3	2	11	1	18	18.56	-31.5	3	5.13	7.56
19	3	2	11	2	19	18.56	-31.44	3.38	5.13	7.56
19	3	2	11	3	19	18.62	-31.56	2.88	5.06	7.5
19	3	2	11	4	19	18.62	-31.5	3.25	5.13	7.5
19	3	2	11	5	19	18.62	-31.62	2.81	5.06	7.44
19	3	2	11	6	20	18.62	-31.69	2.38	5.06	7.44
19	3	2	11	7	20	18.62	-31.56	2.88	5.06	7.44
19	3	2	11	8	20	18.62	-31.69	2.38	5.06	7.44
19	3	2	11	9	20	18.62	-31.62	2.81	5.06	7.38
19	3	2	11	10	21	18.62	-31.75	2.19	5.06	7.38
19	3	2	11	11	21	18.62	-31.87	1.56	5.06	7.31
19	3	2	11	12	21	18.62	-31.75	2.25	5.06	7.31
19	3	2	11	13	22	18.62	-31.94	1.56	5.06	7.31
19	3	2	11	14	22	18.62	-31.81	2.13	5	7.25
19	3	2	11	15	22	18.62	-31.94	1.63	5	7.25
19	3	2	11	16	22	18.62	-31.81	2.06	5.06	7.19
19	3	2	11	17	23	18.62	-31.81	1.88	5	7.19
19	3	2	11	18	23	18.69	-31.94	1.37	5	7.19
19	3	2	11	19	23	18.69	-31.81	1.94	5	7.19
19	3	2	11	20	23	18.69	-31.94	1.19	5	7.19
19	3	2	11	21	24	18.69	-31.87	1.75	5	7.13
19	3	2	11	22	24	18.69	-32.06	1.06	5	7.13
19	3	2	11	23	24	18.69	-32.19	0.37	5	7.06
19	3	2	11	24	24	18.69	-32	1.19	4.94	7.06
19	3	2	11	25	25	18.69	-32.19	0.44	5	7.06
19	3	2	11	26	25	18.69	-32	1.12	5	7
19	3	2	11	27	25	18.75	-32.13	0.56	5	7
19	3	2	11	28	25	18.69	-32	0.63	4.94	6.94
19	3	2	11	29	26	18.69	-32.13	0.37	4.94	6.94
19	3	2	11	30	26	18.69	-32.25	-0.12	4.94	6.88
19	3	2	11	31	26	18.69	-32.06	0.63	4.94	6.88
19	3	2	11	32	26	18.69	-32.19	0.19	4.94	6.88
19	3	2	11	33	27	18.69	-32	0.75	4.94	6.88
19	3	2	11	34	27	18.69	-32.06	0.44	4.94	6.81
19	3	2	11	35	27	18.69	-32.19	-0.19	4.94	6.88
19	3	2	11	36	27	18.69	-32	0.56	4.94	6.81
19	3	2	11	37	28	18.69	-32.13	0.12	4.94	6.81
19	3	2	11	38	28	18.69	-31.94	0.69	4.94	6.81
19	3	2	11	39	28	18.69	-32	0.5	4.94	6.81
19	3	2	11	40	29	18.69	-32.06	0	4.94	6.81
19	3	2	11	41	29	18.62	-31.87	0.75	4.94	6.81
19	3	2	11	42	29	18.62	-31.94	0.44	4.94	6.81
19	3	2	11	43	29	18.62	-31.81	0.94	4.94	6.81
19	3	2	11	44	30	18.62	-31.87	0.69	4.94	6.81
19	3	2	11	45	30	18.62	-31.81	0.5	4.94	6.81
19	3	2	11	46	30	18.62	-31.81	0.88	4.94	6.81
19	3	2	11	47	30	18.62	-31.87	0.5	4.94	6.88
19	3	2	11	48	31	18.62	-31.75	1.06	4.94	6.88
19	3	2	11	49	31	18.62	-31.75	0.88	4.94	6.88
19	3	2	11	50	31	18.62	-31.62	1.31	4.94	6.88
19	3	2	11	51	31	18.62	-31.62	1.31	4.94	6.88
19	3	2	11	52	32	18.62	-31.62	1.06	4.94	6.94
19	3	2	11	53	32	18.62	-31.56	1.5	4.94	6.94
19	3	2	11	54	32	18.62	-31.5	1.31	4.94	6.94
19	3	2	11	55	32	18.62	-31.44	1.69	4.94	7
19	3	2	11	56	33	18.56	-31.5	1.56	4.94	7
19	3	2	11	57	33	18.62	-31.56	1.19	4.94	7
19	3	2	11	58	35	18.62	-31.44	1.69	4.94	7
19	3	2	11	59	36	18.62	-31.44	1.63	4.94	7.06
19	3	2	12	0	36	18.62	-31.37	1.94	4.94	7.06
19	3	2	12	1	36	18.62	-31.44	1.94	4.94	7.06
19	3	2	12	2	36	18.69	-31.44	1.75	4.94	7.13
19	3	2	12	3	37	18.69	-31.44	2.13	4.94	7.13
19	3	2	12	4	37	18.75	-31.44	2.06	4.94	7.13
19	3	2	12	5	37	18.75	-31.37	2.31	5	7.19
19	3	2	12	6	37	18.75	-31.37	2.25	4.94	7.19
19	3	2	12	7	38	18.75	-31.31	2.06	5	7.19
19	3	2	12	8	38	18.75	-31.31	2.44	5	7.25
19	3	2	12	9	38	18.75	-31.25	2.56	5	7.25
19	3	2	12	10	38	18.81	-31.19	2.75	5	7.25
19	3	2	12	11	39	18.81	-31.12	2.88	5	7.31
19	3	2	12	12	39	18.81	-31.06	2.88	5	7.31
19	3	2	12	13	39	18.75	-31.06	3.06	5	7.38
19	3	2	12	14	39	18.75	-31.06	3.13	5	7.38
19	3	2	12	15	40	18.75	-31.06	3.31	5	7.38
19	3	2	12	16	40	18.75	-31.06	3.31	5	7.44
19	3	2	12	17	40	18.75	-31	3.38	5	7.44
19	3	2	12	18	40	18.75	-31.06	3.5	5.06	7.5
19	3	2	12	19	41	18.75	-31	3.63	5.06	7.5
19	3	2	12	20	41	18.75	-30.94	3.75	5.06	7.5
19	3	2	12	21	41	18.75	-30.87	3.81	5.06	7.56
19	3	2	12	22	41	18.75	-30.81	3.88	5.06	7.56
19	3	2	12	23	42	18.75	-30.75	3.94	5.06	7.56
19	3	2	12	24	42	18.75	-30.62	4.06	5.06	7.63
19	3	2	12	25	42	18.75	-30.62	4.06	5.06	7.63
19	3	2	12	26	42	18.75	-30.56	4.19	5.06	7.63
19	3	2	12	27	43	18.75	-30.5	4.19	5.06	7.63
19	3	2	12	28	43	18.75	-30.44	4.25	5.06	7.69
19	3	2	12	29	43	18.75	-30.5	4.19	5.06	7.69
19	3	2	12	30	44	18.69	-30.62	4.31	5.13	7.69
19	3	2	12	31	44	18.69	-30.69	4.31	5.13	7.75
19	3	2	12	32	44	18.69	-30.69	4.38	5.13	7.75
19	3	2	12	33	44	18.69	-30.62	4.44	5.13	7.75
19	3	2	12	34	45	18.69	-30.62	4.38	5.13	7.75
19	3	2	12	35	45	18.62	-30.69	4.44	5.13	7.75
19	3	2	12	36	45	18.69	-30.69	4.44	5.13	7.75
19	3	2	12	37	45	18.69	-30.69	4.56	5.13	7.81
19	3	2	12	38	46	18.69	-30.62	4.63	5.13	7.81
19	3	2	12	39	46	18.69	-30.69	4.31	5.13	7.81
19	3	2	12	40	46	18.69	-30.75	4.5	5.13	7.81
19	3	2	12	41	46	18.69	-30.81	4.31	5.13	7.81
19	3	2	12	42	47	18.69	-30.87	4.31	5.13	7.81
19	3	2	12	43	47	18.69	-30.87	4.44	5.13	7.81
19	3	2	12	44	47	18.69	-30.94	4.25	5.13	7.81
19	3	2	12	45	47	18.69	-30.87	4.38	5.13	7.81
19	3	2	12	46	48	18.75	-30.94	4.25	5.13	7.81
19	3	2	12	47	48	18.75	-30.94	4.25	5.13	7.81
19	3	2	12	48	48	18.75	-30.94	4.44	5.13	7.81
19	3	2	12	49	49	18.81	-31	4	5.13	7.81
19	3	2	12	50	49	18.81	-31.06	4.25	5.13	7.81
19	3	2	12	51	49	18.87	-31.12	3.88	5.13	7.75
19	3	2	12	52	49	18.87	-31.19	3.63	5.13	7.75
19	3	2	12	53	50	18.94	-31.19	3.94	5.13	7.75
19	3	2	12	54	50	18.94	-31.25	3.56	5.13	7.75
19	3	2	12	55	50	18.94	-31.25	3.81	5.13	7.69
19	3	2	12	56	50	18.94	-31.25	3.5	5.13	7.69
19	3	2	12	57	51	18.94	-31.25	3.31	5.13	7.69
19	3	2	12	58	51	18.94	-31.25	3.69	5.13	7.69
19	3	2	12	59	51	18.94	-31.31	3.06	5.13	7.63
19	3	2	13	0	51	19	-31.31	3.44	5.13	7.63
19	3	2	13	1	52	19	-31.44	3	5.13	7.63
19	3	2	13	2	52	19	-31.37	3.31	5.13	7.56
19	3	2	13	3	52	19	-31.31	3.44	5.13	7.56
19	3	2	13	4	52	19.06	-31.37	3.06	5.13	7.56
19	3	2	13	5	53	19.06	-31.31	3.44	5.13	7.56
19	3	2	13	6	53	19.06	-31.37	2.94	5.06	7.56
19	3	2	13	7	53	19.06	-31.37	3.25	5.13	7.56
19	3	2	13	8	53	19.06	-31.5	2.81	5.06	7.5
19	3	2	13	9	54	19.06	-31.62	2.25	5.06	7.5
19	3	2	13	10	54	19.06	-31.5	2.75	5.06	7.44
19	3	2	13	11	54	19.06	-31.69	2.13	5.06	7.44
19	3	2	13	12	54	19.06	-31.56	2.56	5.06	7.44
19	3	2	13	13	55	19.12	-31.75	2.06	5.06	7.38
19	3	2	13	14	55	19.12	-31.81	1.44	5.06	7.38
19	3	2	13	15	55	19.12	-31.69	2.13	5.06	7.31
19	3	2	13	16	56	19.19	-31.87	1.5	5.06	7.31
19	3	2	13	17	56	19.19	-31.69	2.06	5.06	7.31
19	3	2	13	18	56	19.25	-31.81	1.56	5.06	7.25
19	3	2	13	19	56	19.25	-31.69	1.88	5	7.25
19	3	2	13	20	57	19.25	-31.87	1.31	5	7.19
19	3	2	13	21	57	19.25	-32	0.75	5	7.19
19	3	2	13	22	57	19.25	-31.81	1.44	5	7.13
19	3	2	13	23	57	19.25	-31.94	0.75	5	7.13
19	3	2	13	24	58	19.25	-31.81	1.31	5	7.06
19	3	2	13	25	58	19.31	-31.94	0.75	5	7.06
19	3	2	13	26	58	19.31	-32.06	0.12	5	7.06
19	3	2	13	27	58	19.31	-31.81	0.94	5	7.06
19	3	2	13	28	59	19.31	-32.06	-0.31	4.94	7
19	3	2	13	29	59	19.31	-31.87	0.5	4.94	6.94
19	3	2	13	30	59	19.31	-32	-0.12	4.94	6.94
19	3	2	13	31	59	19.31	-31.81	0.37	4.94	6.88
19	3	2	13	33	0	19.37	-31.87	0	4.94	6.88
19	3	2	13	34	0	19.37	-32.06	-0.63	4.94	6.81
19	3	2	13	35	0	19.37	-31.81	0.19	4.94	6.81
19	3	2	13	36	0	19.37	-31.94	-0.44	4.94	6.81
19	3	2	13	37	1	19.44	-31.75	0.25	4.94	6.75
19	3	2	13	38	1	19.44	-31.87	-0.31	4.88	6.75
19	3	2	13	39	1	19.44	-32	-1	4.88	6.75
19	3	2	13	40	2	19.5	-31.81	-0.06	4.88	6.75
19	3	2	13	41	2	19.5	-31.94	-0.56	4.88	6.75
19	3	2	13	42	2	19.5	-31.75	0.12	4.88	6.69
19	3	2	13	43	2	19.5	-31.87	-0.31	4.88	6.69
19	3	2	13	44	3	19.5	-31.69	0.19	4.88	6.69
19	3	2	13	45	3	19.5	-31.81	-0.19	4.88	6.69
19	3	2	13	46	3	19.5	-31.94	-0.81	4.88	6.69
19	3	2	13	47	3	19.5	-31.75	-0.06	4.88	6.69
19	3	2	13	48	4	19.5	-31.87	-0.5	4.88	6.69
19	3	2	13	49	4	19.5	-31.69	0.12	4.88	6.69
19	3	2	13	50	4	19.5	-31.75	-0.19	4.88	6.69
19	3	2	13	51	4	19.5	-31.94	-0.69	4.88	6.69
19	3	2	13	52	5	19.5	-31.69	0.06	4.88	6.69
19	3	2	13	53	5	19.5	-31.75	-0.31	4.88	6.69
19	3	2	13	54	5	19.5	-31.62	0.31	4.88	6.69
19	3	2	13	55	5	19.5	-31.69	0	4.88	6.69
19	3	2	13	56	6	19.5	-31.75	-0.5	4.88	6.69
19	3	2	13	57	6	19.5	-31.62	0.19	4.88	6.69
19	3	2	13	58	6	19.5	-31.75	-0.25	4.88	6.69
19	3	2	13	59	6	19.56	-31.62	0.37	4.88	6.69
19	3	2	14	0	7	19.56	-31.62	0.12	4.88	6.75
19	3	2	14	1	7	19.56	-31.5	0.63	4.88	6.75
19	3	2	14	2	7	19.62	-31.44	0.56	4.88	6.75
19	3	2	14	3	8	19.62	-31.56	0.19	4.88	6.81
19	3	2	14	4	8	19.62	-31.44	0.75	4.88	6.81
19	3	2	14	5	8	19.62	-31.44	0.44	4.88	6.81
19	3	2	14	6	8	19.62	-31.37	0.88	4.88	6.81
19	3	2	14	7	9	19.62	-31.37	0.63	4.88	6.88
19	3	2	14	8	9	19.62	-31.5	0.19	4.88	6.88
19	3	2	14	9	9	19.62	-31.37	0.81	4.88	6.88
19	3	2	14	10	9	19.62	-31.37	0.63	4.88	6.88
19	3	2	14	11	10	19.62	-31.25	1.06	4.88	6.94
19	3	2	14	12	10	19.62	-31.25	0.94	4.88	6.94
19	3	2	14	13	10	19.62	-31.31	0.63	4.88	6.94
19	3	2	14	14	10	19.62	-31.25	1.19	4.94	7
19	3	2	14	15	11	19.56	-31.31	1	4.88	7
19	3	2	14	16	11	19.56	-31.25	1.37	4.88	7
19	3	2	14	17	11	19.56	-31.25	1.19	4.94	7.06
19	3	2	14	18	11	19.56	-31.31	0.88	4.94	7.06
19	3	2	14	19	12	19.56	-31.25	1.44	4.88	7.06
19	3	2	14	20	12	19.56	-31.19	1.37	4.94	7.06
19	3	2	14	21	12	19.56	-31.12	1.75	4.94	7.13
19	3	2	14	22	12	19.56	-31.12	1.69	4.94	7.13
19	3	2	14	23	13	19.56	-31.06	1.75	4.94	7.19
19	3	2	14	24	13	19.56	-31.12	1.94	4.94	7.19
19	3	2	14	25	13	19.5	-31.06	1.81	4.94	7.19
19	3	2	14	26	13	19.5	-31.06	2.13	4.94	7.25
19	3	2	14	27	14	19.5	-31.12	2.06	4.94	7.25
19	3	2	14	28	14	19.5	-31.06	2.31	4.94	7.25
19	3	2	14	29	14	19.5	-30.94	2.44	5	7.31
19	3	2	14	30	15	19.5	-30.87	2.56	5	7.31
19	3	2	14	31	15	19.5	-30.81	2.75	5	7.38
19	3	2	14	32	15	19.5	-30.75	2.81	5	7.38
19	3	2	14	33	15	19.5	-30.75	3	5	7.44
19	3	2	14	34	16	19.5	-30.69	3.06	5	7.44
19	3	2	14	35	16	19.5	-30.62	3.13	5	7.5
19	3	2	14	36	16	19.5	-30.62	3.25	5	7.5
19	3	2	14	37	16	19.56	-30.62	3.31	5	7.5
19	3	2	14	38	17	19.56	-30.62	3.44	5.06	7.56
19	3	2	14	39	17	19.56	-30.62	3.5	5.06	7.56
19	3	2	14	40	17	19.56	-30.56	3.5	5.06	7.56
19	3	2	14	41	17	19.56	-30.56	3.63	5.06	7.63
19	3	2	14	42	18	19.62	-30.62	3.44	5.06	7.63
19	3	2	14	43	18	19.56	-30.69	3.56	5.06	7.69
19	3	2	14	44	18	19.56	-30.75	3.38	5.06	7.63
19	3	2	14	45	18	19.56	-30.75	3.44	5.06	7.69
19	3	2	14	46	19	19.56	-30.75	3.63	5.06	7.69
19	3	2	14	47	19	19.56	-30.75	3.44	5.06	7.69
19	3	2	14	48	19	19.56	-30.75	3.63	5.06	7.69
19	3	2	14	49	20	19.56	-30.75	3.5	5.06	7.69
19	3	2	14	50	20	19.56	-30.81	3.25	5.06	7.69
19	3	2	14	51	20	19.56	-30.81	3.5	5.06	7.69
19	3	2	14	52	20	19.5	-30.94	3.19	5.06	7.69
19	3	2	14	53	21	19.5	-30.94	3.44	5.13	7.69
19	3	2	14	54	21	19.5	-31.06	3.13	5.06	7.69
19	3	2	14	55	21	19.5	-31.06	2.94	5.06	7.69
19	3	2	14	56	21	19.5	-31	3.25	5.06	7.69
19	3	2	14	57	22	19.44	-31	2.94	5.06	7.69
19	3	2	14	58	22	19.44	-31	3.25	5.06	7.69
19	3	2	14	59	22	19.44	-30.94	3.44	5.06	7.69
19	3	2	15	0	22	19.44	-30.94	2.88	5.06	7.69
19	3	2	15	1	23	19.44	-31	3.31	5.06	7.69
19	3	2	15	2	23	19.44	-31.06	2.81	5.06	7.69
19	3	2	15	3	23	19.37	-31.06	3.19	5.06	7.63
19	3	2	15	4	23	19.37	-31.19	2.69	5.06	7.63
19	3	2	15	5	24	19.37	-31.19	2.38	5.06	7.63
19	3	2	15	6	24	19.37	-31.19	2.81	5.06	7.63
19	3	2	15	7	24	19.44	-31.37	2.31	5.06	7.56
19	3	2	15	8	24	19.44	-31.25	2.75	5.06	7.56
19	3	2	15	9	25	19.44	-31.44	2.19	5.06	7.56
19	3	2	15	10	25	19.44	-31.31	2.56	5.06	7.5
19	3	2	15	11	25	19.44	-31.44	2.25	5.06	7.5
19	3	2	15	12	26	19.44	-31.5	1.75	5.06	7.5
19	3	2	15	13	26	19.44	-31.37	2.31	5	7.44
19	3	2	15	14	26	19.44	-31.5	1.75	5	7.44
19	3	2	15	15	26	19.44	-31.37	2.19	5	7.44
19	3	2	15	16	27	19.44	-31.5	1.75	5	7.38
19	3	2	15	17	27	19.44	-31.56	1.19	5	7.38
19	3	2	15	18	27	19.44	-31.44	1.88	5	7.38
19	3	2	15	19	27	19.44	-31.56	1.19	5	7.38
19	3	2	15	20	28	19.44	-31.44	1.69	5	7.31
19	3	2	15	21	28	19.44	-31.56	1.06	5	7.31
19	3	2	15	22	28	19.37	-31.56	0.81	5	7.25
19	3	2	15	23	28	19.37	-31.5	1.19	5	7.25
19	3	2	15	24	29	19.37	-31.69	0.44	4.94	7.25
19	3	2	15	25	29	19.37	-31.5	1.19	4.94	7.19
19	3	2	15	26	29	19.37	-31.62	0.56	4.94	7.19
19	3	2	15	27	29	19.37	-31.5	1.19	4.94	7.13
19	3	2	15	28	30	19.37	-31.62	0.63	4.94	7.13
19	3	2	15	29	30	19.31	-31.75	0.06	4.94	7.13
19	3	2	15	30	30	19.31	-31.56	0.88	4.94	7.06
19	3	2	15	31	30	19.31	-31.62	0.44	4.94	7.06
19	3	2	15	32	31	19.31	-31.44	1	4.94	7.06
19	3	2	15	33	31	19.31	-31.56	0.56	4.94	7.06
19	3	2	15	34	31	19.31	-31.69	-0.06	4.94	7.06
19	3	2	15	35	31	19.37	-31.5	0.75	4.94	7.06
19	3	2	15	36	32	19.37	-31.62	0.31	4.94	7
19	3	2	15	37	32	19.37	-31.44	0.88	4.94	7
19	3	2	15	38	34	19.37	-31.5	0.56	4.94	7
19	3	2	15	39	35	19.37	-31.37	0.63	4.94	7
19	3	2	15	40	35	19.37	-31.37	0.88	4.94	7
19	3	2	15	41	35	19.37	-31.44	0.5	4.88	7
19	3	2	15	42	35	19.37	-31.31	1.06	4.94	7.06
19	3	2	15	43	36	19.37	-31.37	0.81	4.94	7
19	3	2	15	44	36	19.37	-31.25	1.31	4.94	7.06
19	3	2	15	45	36	19.37	-31.31	1	4.94	7.06
19	3	2	15	46	36	19.31	-31.37	0.63	4.94	7.06
19	3	2	15	47	37	19.31	-31.25	1.19	4.94	7.06
19	3	2	15	48	37	19.31	-31.31	0.88	4.94	7.06
19	3	2	15	49	37	19.31	-31.19	1.31	4.94	7.06
19	3	2	15	50	37	19.31	-31.19	1.19	4.94	7.06
19	3	2	15	51	38	19.31	-31.25	0.94	4.94	7.13
19	3	2	15	52	38	19.25	-31.12	1.44	4.94	7.13
19	3	2	15	53	38	19.25	-31.12	1.25	4.94	7.13
19	3	2	15	54	38	19.25	-31.06	1.69	4.94	7.13
19	3	2	15	55	39	19.25	-31.06	1.5	4.94	7.13
19	3	2	15	56	39	19.25	-31.12	1.19	4.94	7.19
19	3	2	15	57	39	19.19	-31	1.69	4.94	7.19
19	3	2	15	58	40	19.25	-31.06	1.44	4.94	7.19
19	3	2	15	59	40	19.25	-31	1.88	4.94	7.19
19	3	2	16	0	40	19.19	-31	1.81	4.94	7.25
19	3	2	16	1	40	19.19	-31	1.63	4.94	7.25
19	3	2	16	2	41	19.25	-30.94	2.13	4.94	7.25
19	3	2	16	3	41	19.25	-30.94	2	4.94	7.31
19	3	2	16	4	41	19.25	-30.94	2.31	4.94	7.31
19	3	2	16	5	41	19.31	-30.87	2.31	4.94	7.31
19	3	2	16	6	42	19.31	-30.87	2.13	5	7.38
19	3	2	16	7	42	19.31	-30.87	2.5	5	7.38
19	3	2	16	8	42	19.31	-30.87	2.31	5	7.38
19	3	2	16	9	42	19.31	-30.81	2.63	5	7.38
19	3	2	16	10	43	19.25	-30.75	2.75	5	7.44
19	3	2	16	11	43	19.31	-30.69	2.94	5	7.44
19	3	2	16	12	43	19.31	-30.62	3.06	5	7.5
19	3	2	16	13	43	19.31	-30.5	3.06	5	7.5
19	3	2	16	14	44	19.31	-30.5	3.31	5.06	7.56
19	3	2	16	15	44	19.25	-30.5	3.31	5.06	7.56
19	3	2	16	16	44	19.25	-30.5	3.44	5.06	7.56
19	3	2	16	17	44	19.25	-30.44	3.5	5.06	7.63
19	3	2	16	18	45	19.25	-30.44	3.56	5.06	7.63
19	3	2	16	19	45	19.25	-30.44	3.69	5.06	7.63
19	3	2	16	20	45	19.25	-30.5	3.5	5.06	7.69
19	3	2	16	21	45	19.25	-30.56	3.63	5.06	7.69
19	3	2	16	22	46	19.25	-30.5	3.69	5.06	7.69
19	3	2	16	23	46	19.25	-30.5	3.69	5.06	7.69
19	3	2	16	24	46	19.25	-30.5	3.88	5.13	7.75
19	3	2	16	25	47	19.25	-30.44	3.63	5.13	7.75
19	3	2	16	26	47	19.25	-30.5	3.81	5.06	7.75
19	3	2	16	27	47	19.25	-30.62	3.63	5.13	7.75
19	3	2	16	28	47	19.25	-30.62	3.81	5.13	7.75
19	3	2	16	29	48	19.25	-30.56	3.94	5.13	7.75
19	3	2	16	30	48	19.25	-30.56	3.56	5.13	7.75
19	3	2	16	31	48	19.25	-30.62	3.75	5.13	7.75
19	3	2	16	32	48	19.25	-30.62	3.5	5.13	7.75
19	3	2	16	33	49	19.25	-30.75	3.25	5.13	7.75
19	3	2	16	34	49	19.31	-30.75	3.56	5.13	7.75
19	3	2	16	35	49	19.31	-30.81	3.19	5.13	7.75
19	3	2	16	36	49	19.31	-30.81	3.5	5.13	7.75
19	3	2	16	37	50	19.31	-30.87	3.19	5.13	7.75
19	3	2	16	38	50	19.31	-31	2.75	5.13	7.69
19	3	2	16	39	50	19.31	-30.87	3.19	5.13	7.69
19	3	2	16	40	50	19.31	-31	2.69	5.13	7.69
19	3	2	16	41	51	19.31	-30.94	3.13	5.13	7.69
19	3	2	16	42	51	19.31	-31	2.69	5.13	7.63
19	3	2	16	43	51	19.31	-31.06	2.38	5.13	7.63
19	3	2	16	44	51	19.31	-31.06	2.75	5.06	7.63
19	3	2	16	45	52	19.31	-31.12	2.25	5.06	7.63
19	3	2	16	46	52	19.31	-31.06	2.75	5.06	7.56
19	3	2	16	47	52	19.37	-31.12	2.31	5.06	7.56
19	3	2	16	48	53	19.37	-31.06	2.75	5.06	7.56
19	3	2	16	49	53	19.44	-31.12	2.5	5.06	7.56
19	3	2	16	50	53	19.44	-31.19	2	5.06	7.56
19	3	2	16	51	53	19.5	-31.06	2.5	5.06	7.5
19	3	2	16	52	54	19.56	-31.19	1.94	5.06	7.5
19	3	2	16	53	54	19.56	-31.12	2.38	5.06	7.5
19	3	2	16	54	54	19.62	-31.25	1.88	5.06	7.44
19	3	2	16	55	54	19.69	-31.31	1.37	5.06	7.44
19	3	2	16	56	55	19.69	-31.19	2	5	7.44
19	3	2	16	57	55	19.75	-31.31	1.5	5	7.38
19	3	2	16	58	55	19.75	-31.19	2	5	7.38
19	3	2	16	59	55	19.81	-31.31	1.37	5	7.38
19	3	2	17	0	56	19.81	-31.44	0.81	5	7.31
19	3	2	17	1	56	19.87	-31.31	1.44	5	7.31
19	3	2	17	2	56	19.87	-31.5	0.69	5	7.31
19	3	2	17	3	56	19.94	-31.31	1.37	5	7.25
19	3	2	17	4	57	19.94	-31.44	0.75	4.94	7.25
19	3	2	17	5	57	20	-31.25	1.31	4.94	7.25
19	3	2	17	6	57	20	-31.37	0.81	4.94	7.19
19	3	2	17	7	57	20	-31.5	0.25	4.94	7.19
19	3	2	17	8	58	20.06	-31.25	1	4.94	7.19
19	3	2	17	9	58	20.06	-31.37	0.56	4.94	7.19
19	3	2	17	10	58	20.12	-31.25	1.12	4.94	7.13
19	3	2	17	11	59	20.12	-31.31	0.75	4.94	7.13
19	3	2	17	12	59	20.12	-31.44	0.06	4.94	7.13
19	3	2	17	13	59	20.12	-31.25	0.88	4.94	7.13
19	3	2	17	14	59	20.19	-31.37	0.44	4.94	7.13
19	3	2	17	16	0	20.19	-31.19	1	4.94	7.13
19	3	2	17	17	0	20.19	-31.25	0.63	4.94	7.06
19	3	2	17	18	0	20.19	-31.12	1	4.94	7.06
19	3	2	17	19	0	20.25	-31.25	0.56	4.94	7.06
19	3	2	17	20	1	20.25	-31.31	0.25	4.94	7.06
19	3	2	17	21	1	20.31	-31.19	0.88	4.94	7.06
19	3	2	17	22	1	20.31	-31.19	0.63	4.94	7.06
19	3	2	17	23	1	20.37	-31.12	1.12	4.94	7.06
19	3	2	17	24	2	20.37	-31.19	0.88	4.88	7.06
19	3	2	17	25	2	20.37	-31.25	0.5	4.94	7.13
19	3	2	17	26	2	20.44	-31.12	1.12	4.94	7.06
19	3	2	17	27	2	20.44	-31.12	0.81	4.94	7.13
19	3	2	17	28	3	20.5	-31.06	1.25	4.94	7.13
19	3	2	17	29	3	20.5	-31.12	1	4.94	7.13
19	3	2	17	30	3	20.56	-31.12	0.69	4.94	7.13
19	3	2	17	31	4	20.56	-31	1.31	4.94	7.13
19	3	2	17	32	4	20.56	-31	1.12	4.94	7.19
19	3	2	17	33	4	20.62	-30.87	1.56	4.94	7.19
19	3	2	17	34	4	20.56	-30.87	1.44	4.94	7.19
19	3	2	17	35	5	20.56	-30.87	1.06	4.94	7.19
19	3	2	17	36	5	20.56	-30.81	1.63	4.94	7.25
19	3	2	17	37	5	20.56	-30.87	1.44	4.94	7.25
19	3	2	17	38	5	20.56	-30.75	1.81	4.94	7.25
19	3	2	17	39	6	20.5	-30.81	1.63	4.94	7.31
19	3	2	17	40	6	20.5	-30.75	1.88	5	7.31
19	3	2	17	41	6	20.5	-30.75	2	5	7.31
19	3	2	17	42	6	20.5	-30.75	1.88	5	7.38
19	3	2	17	43	7	20.44	-30.75	2.25	5	7.38
19	3	2	17	44	7	20.44	-30.69	2.19	5	7.38
19	3	2	17	45	7	20.44	-30.69	2.5	5	7.44
19	3	2	17	46	7	20.37	-30.62	2.69	5	7.44
19	3	2	17	47	8	20.37	-30.56	2.5	5	7.5
19	3	2	17	48	8	20.37	-30.62	2.81	5.06	7.5
19	3	2	17	49	8	20.37	-30.56	3	5.06	7.5
19	3	2	17	50	8	20.31	-30.44	3.13	5.06	7.56
19	3	2	17	51	9	20.31	-30.5	2.88	5.06	7.56
19	3	2	17	52	9	20.31	-30.56	2.69	5.06	7.56
19	3	2	17	53	9	20.31	-30.62	3	5.06	7.63
19	3	2	17	54	10	20.31	-30.56	2.69	5.06	7.63
19	3	2	17	55	10	20.25	-30.56	2.94	5.06	7.63
19	3	2	17	56	10	20.25	-30.69	2.75	5.06	7.63
19	3	2	17	57	10	20.25	-30.69	2.56	5.06	7.63
19	3	2	17	58	11	20.25	-30.69	2.88	5.06	7.63
19	3	2	17	59	11	20.25	-30.81	2.38	5.06	7.63
19	3	2	18	0	11	20.25	-30.75	2.75	5.06	7.63
19	3	2	18	1	11	20.31	-30.87	2.31	5.06	7.63
19	3	2	18	2	12	20.31	-31	1.94	5.06	7.63
19	3	2	18	3	12	20.31	-30.87	2.44	5.06	7.63
19	3	2	18	4	12	20.37	-31	1.94	5.06	7.63
19	3	2	18	5	12	20.37	-30.87	2.44	5.06	7.56
19	3	2	18	6	13	20.37	-31	2	5.06	7.56
19	3	2	18	7	13	20.37	-30.94	2	5.06	7.56
19	3	2	18	8	13	20.37	-30.94	2.19	5.06	7.56
19	3	2	18	9	13	20.37	-31.06	1.56	5.06	7.56
19	3	2	18	10	14	20.37	-31	2.13	5.06	7.56
19	3	2	18	11	14	20.31	-31.06	1.56	5.06	7.5
19	3	2	18	12	14	20.31	-31	2.06	5	7.5
19	3	2	18	13	14	20.31	-31.06	1.56	5	7.5
19	3	2	18	14	15	20.31	-31.19	1.06	5	7.5
19	3	2	18	15	15	20.31	-31.06	1.69	5	7.44
19	3	2	18	16	15	20.31	-31.19	1.19	5	7.44
19	3	2	18	17	16	20.31	-31.06	1.69	5	7.44
19	3	2	18	18	16	20.31	-31.19	1.19	5	7.38
19	3	2	18	19	16	20.31	-31.31	0.56	5	7.38
19	3	2	18	20	16	20.31	-31.12	1.31	5	7.38
19	3	2	18	21	17	20.37	-31.25	0.63	5	7.38
19	3	2	18	22	17	20.37	-31.12	1.25	5	7.31
19	3	2	18	23	17	20.37	-31.25	0.63	4.94	7.31
19	3	2	18	24	17	20.44	-31.12	1.06	4.94	7.25
19	3	2	18	25	18	20.44	-31.25	0.63	4.94	7.25
19	3	2	18	26	18	20.5	-31.31	0.12	4.94	7.25
19	3	2	18	27	18	20.5	-31.12	0.81	4.94	7.25
19	3	2	18	28	18	20.5	-31.25	0.25	4.94	7.19
19	3	2	18	29	19	20.5	-31.06	0.88	4.94	7.19
19	3	2	18	30	19	20.5	-31.12	0.5	4.94	7.19
19	3	2	18	31	19	20.5	-31.25	-0.06	4.94	7.19
19	3	2	18	32	19	20.44	-31.12	0.75	4.94	7.19
19	3	2	18	33	20	20.5	-31.19	0.31	4.88	7.19
19	3	2	18	34	20	20.44	-31.06	0.88	4.94	7.13
19	3	2	18	35	20	20.44	-31.12	0.5	4.94	7.13
19	3	2	18	36	20	20.37	-31.19	-0.12	4.94	7.13
19	3	2	18	37	21	20.37	-31.06	0.63	4.94	7.13
19	3	2	18	38	21	20.37	-31.19	0.12	4.88	7.13
19	3	2	18	39	21	20.31	-31	0.81	4.88	7.13
19	3	2	18	40	22	20.31	-31.06	0.5	4.94	7.13
19	3	2	18	41	22	20.31	-30.94	1.06	4.94	7.13
19	3	2	18	42	22	20.25	-30.87	1.12	4.94	7.19
19	3	2	18	43	22	20.25	-31	0.69	4.94	7.19
19	3	2	18	44	23	20.25	-30.81	1.25	4.94	7.19
19	3	2	18	45	23	20.25	-30.87	0.94	4.94	7.19
19	3	2	18	46	23	20.25	-30.81	1.37	4.94	7.19
19	3	2	18	47	23	20.31	-30.87	1.12	4.94	7.25
19	3	2	18	48	24	20.31	-30.94	0.69	4.94	7.25
19	3	2	18	49	24	20.31	-30.81	1.25	4.94	7.25
19	3	2	18	50	24	20.31	-30.81	1.06	4.94	7.25
19	3	2	18	51	24	20.31	-30.75	1.5	4.94	7.25
19	3	2	18	52	25	20.25	-30.75	1.37	4.94	7.25
19	3	2	18	53	25	20.25	-30.81	1.06	5	7.31
19	3	2	18	54	25	20.25	-30.69	1.63	5	7.31
19	3	2	18	55	25	20.25	-30.75	1.44	5	7.31
19	3	2	18	56	26	20.19	-30.62	1.81	5	7.31
19	3	2	18	57	26	20.19	-30.69	1.63	5	7.38
19	3	2	18	58	26	20.19	-30.69	1.25	5	7.38
19	3	2	18	59	27	20.12	-30.62	1.81	5	7.38
19	3	2	19	0	27	20.12	-30.56	1.75	5	7.44
19	3	2	19	1	27	20.12	-30.56	2.13	5	7.44
19	3	2	19	2	27	20.06	-30.56	2.13	5	7.44
19	3	2	19	3	28	20.12	-30.5	2.06	5	7.5
19	3	2	19	4	28	20.12	-30.5	2.38	5.06	7.5
19	3	2	19	5	28	20.12	-30.5	2.25	5.06	7.5
19	3	2	19	6	28	20.12	-30.5	2.56	5.06	7.56
19	3	2	19	7	29	20.12	-30.5	2.5	5.06	7.56
19	3	2	19	8	33	20.12	-30.69	1.56	5.06	7.56
19	3	2	19	9	33	20.12	-30.81	1.37	5.06	7.5
19	3	2	19	10	33	20.19	-30.87	1.12	5	7.5
19	3	2	19	11	34	20.19	-30.81	1.69	5	7.5
19	3	2	19	12	34	20.19	-30.87	1.44	5	7.5
19	3	2	19	13	34	20.19	-30.75	1.88	5	7.5
19	3	2	19	14	34	20.25	-30.81	1.69	5	7.5
19	3	2	19	15	35	20.19	-30.69	1.63	5	7.5
19	3	2	19	16	35	20.19	-30.69	2.06	5	7.5
19	3	2	19	17	35	20.19	-30.75	1.88	5	7.5
19	3	2	19	18	35	20.19	-30.69	2.25	5.06	7.5
19	3	2	19	19	36	20.12	-30.81	1.88	5.06	7.5
19	3	2	19	20	36	20.12	-30.81	1.5	5.06	7.5
19	3	2	19	21	36	20.12	-30.75	2.06	5.06	7.5
19	3	2	19	22	37	20.06	-30.81	1.63	5.06	7.56
19	3	2	19	23	37	20.06	-30.75	2.06	5.06	7.5
19	3	2	19	24	37	20.06	-30.81	1.69	5.06	7.5
19	3	2	19	25	37	20	-30.87	1.31	5.06	7.5
19	3	2	19	26	38	20	-30.75	1.94	5.06	7.5
19	3	2	19	27	38	20	-30.87	1.5	5.06	7.5
19	3	2	19	28	38	19.94	-30.75	1.94	5.06	7.5
19	3	2	19	29	38	19.94	-30.94	1.44	5.06	7.5
19	3	2	19	30	39	19.87	-31	1	5	7.5
19	3	2	19	31	39	19.87	-30.94	1.56	5	7.44
19	3	2	19	32	39	19.87	-31.12	0.94	5	7.44
19	3	2	19	33	39	19.87	-31	1.5	5	7.44
19	3	2	19	34	40	19.94	-31.12	1.06	5	7.44
19	3	2	19	35	40	19.94	-31	1.56	5	7.38
19	3	2	19	36	40	19.94	-31.19	1.12	5	7.38
19	3	2	19	37	40	19.94	-31.25	0.75	5	7.38
19	3	2	19	38	41	20	-31.06	1.37	5	7.38
19	3	2	19	39	41	20	-31.25	0.69	5	7.38
19	3	2	19	40	41	20.06	-31.06	1.25	5	7.31
19	3	2	19	41	41	20.06	-31.25	0.69	5	7.31
19	3	2	19	42	42	20.12	-31.37	0.06	5	7.31
19	3	2	19	43	42	20.12	-31.19	0.88	4.94	7.25
19	3	2	19	44	42	20.12	-31.37	0.25	4.94	7.25
19	3	2	19	45	43	20.12	-31.19	0.88	4.94	7.25
19	3	2	19	46	43	20.12	-31.31	0.31	4.94	7.19
19	3	2	19	47	43	20.12	-31.19	0.75	4.94	7.19
19	3	2	19	48	43	20.12	-31.31	0.25	4.94	7.19
19	3	2	19	49	44	20.06	-31.44	-0.19	4.94	7.13
19	3	2	19	50	44	20.06	-31.25	0.5	4.94	7.13
19	3	2	19	51	44	20.06	-31.31	0.12	4.94	7.13
19	3	2	19	52	44	20.06	-31.12	0.69	4.94	7.13
19	3	2	19	53	45	20	-31.25	0.37	4.94	7.13
19	3	2	19	54	45	20	-31.37	-0.25	4.94	7.13
19	3	2	19	55	45	20	-31.19	0.56	4.94	7.13
19	3	2	19	56	45	19.94	-31.25	0.12	4.94	7.13
19	3	2	19	57	46	19.94	-31.06	0.69	4.94	7.13
19	3	2	19	58	46	19.94	-31.19	0.31	4.94	7.13
19	3	2	19	59	46	19.87	-31.25	-0.12	4.94	7.13
19	3	2	20	0	46	19.87	-31.12	0.63	4.88	7.06
19	3	2	20	1	47	19.87	-31.12	0.31	4.94	7.13
19	3	2	20	2	47	19.81	-31	0.88	4.94	7.13
19	3	2	20	3	47	19.81	-31.06	0.63	4.94	7.13
19	3	2	20	4	47	19.81	-30.87	1.12	4.94	7.13
19	3	2	20	5	48	19.81	-31	0.75	4.94	7.13
19	3	2	20	6	48	19.81	-31.12	0.37	4.94	7.13
19	3	2	20	7	48	19.81	-30.94	1	4.94	7.13
19	3	2	20	8	49	19.81	-31	0.69	4.94	7.13
19	3	2	20	9	49	19.81	-30.87	1.19	4.94	7.13
19	3	2	20	10	49	19.87	-30.87	1.06	4.94	7.19
19	3	2	20	11	49	19.87	-30.87	0.81	4.94	7.19
19	3	2	20	12	50	19.87	-30.75	1.37	4.94	7.19
19	3	2	20	13	50	19.94	-30.75	1.19	4.94	7.25
19	3	2	20	14	50	19.94	-30.62	1.63	4.94	7.25
19	3	2	20	15	50	19.94	-30.69	1.44	4.94	7.31
19	3	2	20	16	51	19.94	-30.69	1.12	5	7.31
19	3	2	20	17	51	19.94	-30.69	1.69	5	7.31
19	3	2	20	18	51	19.94	-30.69	1.44	5	7.31
19	3	2	20	19	51	19.94	-30.56	1.81	5	7.31
19	3	2	20	20	52	19.94	-30.56	1.81	5	7.38
19	3	2	20	21	52	19.94	-30.5	1.69	5	7.38
19	3	2	20	22	52	19.94	-30.5	2.13	5	7.44
19	3	2	20	23	52	19.87	-30.5	2.06	5	7.44
19	3	2	20	24	53	19.87	-30.5	2.38	5	7.44
19	3	2	20	25	53	19.87	-30.5	2.31	5	7.5
19	3	2	20	26	53	19.87	-30.5	2.06	5.06	7.5
19	3	2	20	27	53	19.87	-30.5	2.5	5.06	7.5
19	3	2	20	28	54	19.87	-30.81	1.19	5.06	7.56
19	3	2	20	29	54	19.94	-30.75	1.75	5	7.5
19	3	2	20	30	54	19.94	-30.75	1.56	5.06	7.5
19	3	2	20	31	55	19.94	-30.69	1.94	5.06	7.5
19	3	2	20	32	55	19.94	-30.5	2.31	5.06	7.5
19	3	2	20	33	55	19.94	-30.44	2.13	5.06	7.5
19	3	2	20	34	55	19.94	-30.37	2.5	5.06	7.5
19	3	2	20	35	56	19.94	-30.37	2.19	5.06	7.56
19	3	2	20	36	56	19.94	-30.37	2.5	5.06	7.56
19	3	2	20	37	56	19.94	-30.44	2.31	5.06	7.56
19	3	2	20	38	56	19.94	-30.5	2	5.06	7.56
19	3	2	20	39	57	19.94	-30.5	2.38	5.06	7.56
19	3	2	20	40	57	19.94	-30.62	2	5.06	7.56
19	3	2	20	41	57	19.94	-30.56	2.38	5.06	7.56
19	3	2	20	42	57	19.94	-30.69	2	5.06	7.56
19	3	2	20	43	58	19.94	-30.75	1.69	5.06	7.56
19	3	2	20	44	58	19.94	-30.62	2.19	5.06	7.56
19	3	2	20	45	58	19.94	-30.69	1.75	5.06	7.56
19	3	2	20	46	58	19.94	-30.62	2.19	5.06	7.56
19	3	2	20	47	59	20	-30.75	1.88	5.06	7.56
19	3	2	20	48	59	20.06	-30.81	1.31	5.06	7.5
19	3	2	20	49	59	20.06	-30.75	1.94	5.06	7.5
19	3	2	20	50	59	20.06	-30.94	1.37	5.06	7.5
19	3	2	20	52	0	20.12	-30.81	1.88	5.06	7.5
19	3	2	20	53	0	20.12	-30.94	1.37	5.06	7.5
19	3	2	20	54	0	20.12	-30.81	1.88	5.06	7.44
19	3	2	20	55	1	20.12	-30.94	1.5	5.06	7.44
19	3	2	20	56	1	20.12	-31.06	1.06	5.06	7.44
19	3	2	20	57	1	20.12	-30.87	1.63	5.06	7.44
19	3	2	20	58	1	20.12	-31	1.12	5.06	7.38
19	3	2	20	59	2	20.12	-30.87	1.63	5	7.38
19	3	2	21	0	2	20.12	-31.06	1.06	5	7.38
19	3	2	21	1	2	20.06	-31.12	0.44	5	7.31
19	3	2	21	2	2	20.06	-31	1.19	5	7.31
19	3	2	21	3	3	20.06	-31.12	0.56	5	7.31
19	3	2	21	4	3	20.06	-31	1.19	5	7.25
19	3	2	21	5	3	20.06	-31.12	0.56	5	7.25
19	3	2	21	6	3	20.06	-31.19	0.12	5	7.25
19	3	2	21	7	4	20.06	-31.06	0.81	5	7.25
19	3	2	21	8	4	20.06	-31.25	0.12	5	7.19
19	3	2	21	9	4	20.06	-31.06	0.81	5	7.19
19	3	2	21	10	4	20.12	-31.19	0.12	5	7.19
19	3	2	21	11	5	20.12	-31.06	0.75	4.94	7.19
19	3	2	21	12	5	20.12	-31.19	0.12	4.94	7.13
19	3	2	21	13	5	20.12	-31.37	-0.63	4.94	7.13
19	3	2	21	14	5	20.12	-31.12	0.25	4.94	7.06
19	3	2	21	15	6	20.12	-31.31	-0.44	4.94	7.06
19	3	2	21	16	6	20.12	-31.06	0.31	4.94	7.06
19	3	2	21	17	6	20.12	-31.25	-0.31	4.94	7
19	3	2	21	18	7	20.12	-31.25	-0.75	4.94	7
19	3	2	21	19	7	20.06	-31.12	-0.06	4.88	7
19	3	2	21	20	7	20.06	-31.31	-0.69	4.94	7
19	3	2	21	21	7	20.06	-31.12	0.12	4.94	7
19	3	2	21	22	8	20.06	-31.19	-0.37	4.88	6.94
19	3	2	21	23	8	20.06	-31	0.25	4.94	6.94
19	3	2	21	24	8	20.06	-31.12	-0.12	4.88	6.94
19	3	2	21	25	8	20.06	-31.25	-0.75	4.88	7
19	3	2	21	26	9	20	-31.06	0.06	4.88	6.94
19	3	2	21	27	9	20	-31.19	-0.37	4.88	6.94
19	3	2	21	28	9	20	-31	0.25	4.88	6.94
19	3	2	21	29	9	20.06	-31.06	-0.12	4.88	6.94
19	3	2	21	30	10	20.06	-31.06	-0.69	4.88	6.94
19	3	2	21	31	10	20.06	-30.94	0.19	4.88	6.94
19	3	2	21	32	10	20.06	-31	-0.19	4.88	7
19	3	2	21	33	10	20.12	-30.87	0.44	4.88	7
19	3	2	21	34	11	20.12	-30.87	0.12	4.88	7
19	3	2	21	35	11	20.12	-30.75	0.63	4.94	7
19	3	2	21	36	11	20.12	-30.69	0.75	4.88	7
19	3	2	21	37	11	20.12	-30.81	0.25	4.88	7.06
19	3	2	21	38	12	20.06	-30.69	0.81	4.94	7.06
19	3	2	21	39	12	20.06	-30.75	0.56	4.94	7.13
19	3	2	21	40	12	20.06	-30.62	1.06	4.94	7.13
19	3	2	21	41	13	20.06	-30.69	0.81	4.94	7.13
19	3	2	21	42	13	20.06	-30.75	0.37	4.94	7.13
19	3	2	21	43	13	20.06	-30.62	1	4.94	7.13
19	3	2	21	44	13	20	-30.69	0.75	4.94	7.19
19	3	2	21	45	14	20	-30.56	1.19	4.94	7.19
19	3	2	21	46	14	20	-30.69	0.94	4.94	7.19
19	3	2	21	47	14	19.94	-30.69	0.44	4.94	7.19
19	3	2	21	48	14	20	-30.56	1.12	4.94	7.19
19	3	2	21	49	15	20	-30.56	0.94	5	7.25
19	3	2	21	50	15	20	-30.44	1.44	4.94	7.25
19	3	2	21	51	15	20	-30.5	1.31	5	7.25
19	3	2	21	52	15	20	-30.44	1.12	5	7.31
19	3	2	21	53	16	20	-30.44	1.5	5	7.31
19	3	2	21	54	16	20	-30.44	1.31	5	7.31
19	3	2	21	55	16	20.06	-30.37	1.75	5	7.38
19	3	2	21	56	16	20	-30.44	1.56	5	7.38
19	3	2	21	57	17	20	-30.37	1.94	5	7.38
19	3	2	21	58	17	20	-30.25	2.19	5	7.44
19	3	2	21	59	17	20	-30.19	2.06	5	7.44
19	3	2	22	0	18	20	-30.19	2.38	5.06	7.5
19	3	2	22	1	18	19.94	-30.19	2.31	5.06	7.5
19	3	2	22	2	18	19.94	-30.12	2.56	5.06	7.5
19	3	2	22	3	18	19.94	-30.12	2.5	5.06	7.56
19	3	2	22	4	19	19.94	-30.12	2.38	5.06	7.56
19	3	2	22	5	19	19.94	-30.06	2.69	5.06	7.56
19	3	2	22	6	19	19.94	-30.06	2.56	5.06	7.63
19	3	2	22	7	19	19.94	-30.06	2.81	5.06	7.63
19	3	2	22	8	20	19.94	-30.12	2.75	5.06	7.63
19	3	2	22	9	20	19.94	-30.12	2.63	5.06	7.69
19	3	2	22	10	20	19.94	-30.06	2.94	5.13	7.69
19	3	2	22	11	20	20	-30.12	2.63	5.13	7.69
19	3	2	22	12	21	20	-30.12	2.94	5.13	7.69
19	3	2	22	13	21	20	-30.19	2.69	5.13	7.75
19	3	2	22	14	21	20.06	-30.19	2.5	5.13	7.69
19	3	2	22	15	21	20.06	-30.12	2.88	5.13	7.75
19	3	2	22	16	22	20.06	-30.25	2.63	5.13	7.75
19	3	2	22	17	22	20.12	-30.19	2.94	5.13	7.75
19	3	2	22	18	22	20.12	-30.25	2.69	5.13	7.75
19	3	2	22	19	22	20.12	-30.31	2.31	5.13	7.75
19	3	2	22	20	23	20.19	-30.31	2.75	5.13	7.75
19	3	2	22	21	23	20.19	-30.37	2.31	5.13	7.75
19	3	2	22	22	23	20.19	-30.31	2.69	5.13	7.75
19	3	2	22	23	24	20.19	-30.44	2.38	5.13	7.75
19	3	2	22	24	24	20.25	-30.44	2.06	5.13	7.75
19	3	2	22	25	24	20.25	-30.44	2.5	5.13	7.75
19	3	2	22	26	24	20.25	-30.56	2.13	5.13	7.75
19	3	2	22	27	25	20.25	-30.44	2.5	5.13	7.75
19	3	2	22	28	25	20.25	-30.5	2.13	5.13	7.69
19	3	2	22	29	25	20.25	-30.44	2.56	5.13	7.69
19	3	2	22	30	25	20.19	-30.37	2.5	5.13	7.69
19	3	2	22	31	26	20.19	-30.5	2.06	5.13	7.69
19	3	2	22	32	26	20.19	-30.44	2.56	5.13	7.69
19	3	2	22	33	26	20.12	-30.56	2	5.13	7.69
19	3	2	22	34	26	20.12	-30.5	2.44	5.13	7.69
19	3	2	22	35	27	20.12	-30.62	2	5.13	7.69
19	3	2	22	36	27	20.06	-30.69	1.5	5.13	7.63
19	3	2	22	37	27	20.06	-30.56	2.13	5.13	7.63
19	3	2	22	38	27	20	-30.69	1.63	5.13	7.63
19	3	2	22	39	28	20	-30.56	2.06	5.13	7.56
19	3	2	22	40	28	20	-30.75	1.5	5.13	7.56
19	3	2	22	41	28	19.94	-30.87	0.88	5.13	7.56
19	3	2	22	42	28	19.94	-30.75	1.56	5.06	7.5
19	3	2	22	43	29	19.87	-30.94	0.88	5.06	7.5
19	3	2	22	44	29	19.87	-30.75	1.5	5.06	7.44
19	3	2	22	45	29	19.87	-30.87	0.94	5.06	7.44
19	3	2	22	46	30	19.81	-30.75	1.5	5.06	7.44
19	3	2	22	47	30	19.81	-30.87	1	5.06	7.38
19	3	2	22	48	32	19.81	-31	0.5	5.06	7.38
19	3	2	22	49	32	19.81	-30.81	1.19	5.06	7.31
19	3	2	22	50	32	19.81	-30.94	0.5	5.06	7.31
19	3	2	22	51	32	19.81	-30.75	1.12	5.06	7.31
19	3	2	22	52	33	19.81	-30.94	0.5	5.06	7.25
19	3	2	22	53	33	19.81	-31.06	-0.25	5	7.25
19	3	2	22	54	33	19.81	-30.87	0.63	5	7.25
19	3	2	22	55	33	19.81	-31	-0.06	5	7.25
19	3	2	22	56	34	19.81	-30.81	0.63	5	7.19
19	3	2	22	57	34	19.87	-31	0.06	5	7.19
19	3	2	22	58	34	19.87	-30.81	0.69	5	7.13
19	3	2	22	59	34	19.87	-30.94	0.25	5	7.13
19	3	2	23	0	35	19.87	-31.06	-0.5	5	7.13
19	3	2	23	1	35	19.87	-30.87	0.31	5	7.06
19	3	2	23	2	35	19.87	-31	-0.44	5	7.06
19	3	2	23	3	35	19.87	-30.81	0.25	5	7.06
19	3	2	23	4	36	19.87	-30.94	-0.37	4.94	7
19	3	2	23	5	36	19.87	-31	-1	4.94	7
19	3	2	23	6	36	19.87	-30.87	-0.19	4.94	7
19	3	2	23	7	36	19.87	-31	-1	4.94	7
19	3	2	23	8	37	19.87	-30.81	-0.19	4.94	6.94
19	3	2	23	9	37	19.81	-30.94	-0.69	4.94	6.94
19	3	2	23	10	37	19.81	-30.75	0	4.94	6.94
19	3	2	23	11	38	19.81	-30.87	-0.5	4.94	6.94
19	3	2	23	12	38	19.75	-31	-1.25	4.94	6.94
19	3	2	23	13	38	19.75	-30.81	-0.31	4.94	6.88
19	3	2	23	14	38	19.75	-30.94	-0.88	4.94	6.88
19	3	2	23	15	39	19.75	-30.75	-0.19	4.94	6.88
19	3	2	23	16	39	19.75	-30.87	-0.69	4.94	6.88
19	3	2	23	17	39	19.75	-30.75	-0.44	4.94	6.88
19	3	2	23	18	39	19.75	-30.81	-0.69	4.94	6.81
19	3	2	23	19	40	19.75	-30.94	-1.19	4.88	6.81
19	3	2	23	20	40	19.75	-30.75	-0.37	4.88	6.81
19	3	2	23	21	40	19.81	-30.87	-0.75	4.88	6.81
19	3	2	23	22	40	19.81	-30.69	-0.12	4.88	6.81
19	3	2	23	23	41	19.75	-30.81	-0.5	4.88	6.81
19	3	2	23	24	41	19.75	-30.94	-1.12	4.88	6.81
19	3	2	23	25	41	19.75	-30.75	-0.25	4.88	6.81
19	3	2	23	26	41	19.75	-30.87	-0.75	4.88	6.81
19	3	2	23	27	42	19.69	-30.69	-0.06	4.88	6.81
19	3	2	23	28	42	19.69	-30.75	-0.44	4.88	6.88
19	3	2	23	29	42	19.69	-30.62	-0.31	4.88	6.88
19	3	2	23	30	43	19.69	-30.75	-0.44	4.88	6.81
19	3	2	23	31	43	19.62	-30.81	-0.81	4.88	6.88
19	3	2	23	32	43	19.62	-30.62	-0.12	4.88	6.81
19	3	2	23	33	43	19.62	-30.75	-0.44	4.88	6.88
19	3	2	23	34	44	19.62	-30.56	0.19	4.88	6.88
19	3	2	23	35	44	19.62	-30.69	-0.12	4.88	6.88
19	3	2	23	36	44	19.62	-30.81	-0.56	4.88	6.88
19	3	2	23	37	44	19.62	-30.62	0.12	4.88	6.88
19	3	2	23	38	45	19.62	-30.75	-0.25	4.88	6.88
19	3	2	23	39	45	19.62	-30.56	0.31	4.94	6.94
19	3	2	23	40	45	19.62	-30.62	0.19	4.94	6.94
19	3	2	23	41	45	19.62	-30.69	-0.25	4.94	6.94
19	3	2	23	42	46	19.62	-30.5	0.44	4.94	7
19	3	2	23	43	46	19.56	-30.56	0.19	4.94	7
19	3	2	23	44	46	19.56	-30.44	0.69	4.94	7
19	3	2	23	45	46	19.56	-30.5	0.5	4.94	7
19	3	2	23	46	47	19.5	-30.44	0.37	4.94	7.06
19	3	2	23	47	47	19.5	-30.44	0.75	4.94	7.06
19	3	2	23	48	47	19.5	-30.31	1.12	4.94	7.06
19	3	2	23	49	47	19.5	-30.19	1.44	4.94	7.13
19	3	2	23	50	48	19.5	-30.19	1.25	4.94	7.19
19	3	2	23	51	48	19.5	-30.06	1.63	5	7.19
19	3	2	23	52	48	19.5	-30.12	1.5	5	7.19
19	3	2	23	53	49	19.5	-30.19	1.19	5	7.25
19	3	2	23	54	49	19.5	-30.12	1.63	5	7.25
19	3	2	23	55	49	19.5	-30.19	1.44	5	7.25
19	3	2	23	56	49	19.56	-30.06	1.81	5	7.31
19	3	2	23	57	50	19.56	-30.12	1.63	5	7.31
19	3	2	23	58	50	19.56	-30.12	1.56	5	7.38
19	3	2	23	59	50	19.56	-30.06	1.94	5	7.38
19	3	3	0	0	50	19.56	-30	1.94	5.06	7.38
19	3	3	0	1	51	19.56	-30	2.19	5.06	7.44
19	3	3	0	2	51	19.5	-30	2.19	5.06	7.44
19	3	3	0	3	51	19.56	-30	2	5.06	7.44
19	3	3	0	4	51	19.5	-30	2.38	5.06	7.5
19	3	3	0	5	52	19.5	-29.94	2.31	5.06	7.5
19	3	3	0	6	52	19.5	-29.94	2.56	5.06	7.5
19	3	3	0	7	52	19.5	-29.94	2.5	5.06	7.56
19	3	3	0	8	52	19.5	-29.94	2.38	5.13	7.56
19	3	3	0	9	53	19.56	-29.94	2.69	5.06	7.56
19	3	3	0	10	53	19.56	-29.94	2.44	5.06	7.63
19	3	3	0	11	53	19.56	-29.94	2.75	5.13	7.63
19	3	3	0	12	53	19.62	-30	2.5	5.13	7.63
19	3	3	0	13	54	19.62	-30	2.31	5.13	7.63
19	3	3	0	14	54	19.62	-30	2.69	5.13	7.63
19	3	3	0	15	54	19.69	-30.06	2.44	5.13	7.63
19	3	3	0	16	54	19.69	-30	2.75	5.13	7.69
19	3	3	0	17	55	19.69	-30.06	2.56	5.13	7.69
19	3	3	0	18	55	19.69	-30	2.56	5.13	7.63
19	3	3	0	19	55	19.62	-30.12	2.63	5.13	7.69
19	3	3	0	20	56	19.62	-30.25	2.13	5.13	7.69
19	3	3	0	21	56	19.62	-30.19	2.5	5.13	7.69
19	3	3	0	22	56	19.62	-30.25	2.19	5.13	7.63
19	3	3	0	23	56	19.56	-30.25	2.5	5.13	7.63
19	3	3	0	24	57	19.56	-30.31	2.19	5.13	7.63
19	3	3	0	25	57	19.56	-30.31	1.94	5.13	7.63
19	3	3	0	26	57	19.5	-30.25	2.38	5.13	7.63
19	3	3	0	27	57	19.5	-30.37	2.06	5.13	7.63
19	3	3	0	28	58	19.5	-30.25	2.44	5.13	7.56
19	3	3	0	29	58	19.5	-30.44	2	5.13	7.56
19	3	3	0	30	58	19.5	-30.5	1.5	5.13	7.56
19	3	3	0	31	58	19.5	-30.44	2.06	5.13	7.56
19	3	3	0	32	59	19.5	-30.56	1.56	5.13	7.56
19	3	3	0	33	59	19.5	-30.44	2	5.13	7.5
19	3	3	0	34	59	19.56	-30.62	1.56	5.13	7.5
19	3	3	0	35	59	19.56	-30.62	1.12	5.13	7.5
19	3	3	0	37	0	19.5	-30.56	1.75	5.06	7.5
19	3	3	0	38	0	19.5	-30.69	1.25	5.06	7.44
19	3	3	0	39	0	19.5	-30.56	1.81	5.06	7.44
19	3	3	0	40	1	19.5	-30.69	1.19	5.06	7.44
19	3	3	0	41	1	19.44	-30.56	1.69	5.06	7.38
19	3	3	0	42	1	19.44	-30.69	1.37	5.06	7.38
19	3	3	0	43	1	19.44	-30.81	0.81	5.06	7.38
19	3	3	0	44	2	19.37	-30.62	1.44	5.06	7.38
19	3	3	0	45	2	19.37	-30.81	0.81	5.06	7.38
19	3	3	0	46	2	19.37	-30.62	1.37	5.06	7.31
19	3	3	0	47	2	19.37	-30.75	0.88	5.06	7.31
19	3	3	0	48	3	19.37	-30.87	0.25	5.06	7.31
19	3	3	0	49	3	19.37	-30.69	1.06	5.06	7.25
19	3	3	0	50	3	19.37	-30.87	0.31	5	7.25
19	3	3	0	51	3	19.37	-30.69	0.94	5	7.25
19	3	3	0	52	4	19.37	-30.81	0.37	5	7.19
19	3	3	0	53	4	19.37	-30.62	0.94	5	7.19
19	3	3	0	54	4	19.44	-30.69	0.75	5	7.19
19	3	3	0	55	4	19.44	-30.81	0.12	5	7.19
19	3	3	0	56	5	19.44	-30.62	0.81	5	7.13
19	3	3	0	57	5	19.37	-30.81	0.19	5	7.13
19	3	3	0	58	5	19.37	-30.62	0.81	5	7.13
19	3	3	0	59	5	19.37	-30.75	0.25	5	7.13
19	3	3	1	0	6	19.37	-30.94	-0.69	5	7.06
19	3	3	1	1	6	19.37	-30.75	0.25	5	7.06
19	3	3	1	2	6	19.37	-30.94	-0.63	5	7.06
19	3	3	1	3	7	19.37	-30.69	0.12	4.94	7
19	3	3	1	4	7	19.37	-30.87	-0.56	4.94	6.94
19	3	3	1	5	7	19.37	-30.69	0.12	4.94	6.94
19	3	3	1	6	7	19.37	-30.81	-0.5	4.94	6.88
19	3	3	1	7	8	19.37	-31	-1.19	4.94	6.88
19	3	3	1	8	8	19.44	-30.75	-0.25	4.94	6.88
19	3	3	1	9	8	19.44	-30.87	-0.81	4.94	6.88
19	3	3	1	10	8	19.5	-30.62	-0.06	4.94	6.88
19	3	3	1	11	9	19.5	-30.81	-0.56	4.94	6.81
19	3	3	1	12	9	19.5	-30.81	-1	4.94	6.81
19	3	3	1	13	9	19.5	-30.69	-0.37	4.94	6.81
19	3	3	1	14	9	19.56	-30.81	-1.12	4.94	6.81
19	3	3	1	15	10	19.56	-30.62	-0.31	4.88	6.81
19	3	3	1	16	10	19.5	-30.75	-0.88	4.88	6.81
19	3	3	1	17	10	19.5	-30.56	-0.19	4.88	6.75
19	3	3	1	18	10	19.5	-30.69	-0.75	4.88	6.75
19	3	3	1	19	11	19.5	-30.87	-1.37	4.88	6.75
19	3	3	1	20	11	19.44	-30.69	-0.44	4.88	6.75
19	3	3	1	21	11	19.44	-30.81	-0.94	4.88	6.75
19	3	3	1	22	11	19.44	-30.62	-0.25	4.88	6.75
19	3	3	1	23	12	19.37	-30.69	-0.63	4.88	6.75
19	3	3	1	24	12	19.44	-30.62	-1	4.88	6.75
19	3	3	1	25	12	19.37	-30.62	-0.44	4.88	6.75
19	3	3	1	26	12	19.37	-30.75	-1	4.88	6.75
19	3	3	1	27	13	19.44	-30.56	-0.25	4.88	6.75
19	3	3	1	28	13	19.44	-30.69	-0.75	4.88	6.75
19	3	3	1	29	13	19.44	-30.56	-0.12	4.88	6.75
19	3	3	1	30	14	19.44	-30.62	-0.37	4.88	6.75
19	3	3	1	31	14	19.44	-30.69	-0.81	4.88	6.75
19	3	3	1	32	14	19.44	-30.5	-0.12	4.88	6.75
19	3	3	1	33	14	19.5	-30.62	-0.44	4.88	6.75
19	3	3	1	34	15	19.5	-30.44	0.12	4.88	6.81
19	3	3	1	35	15	19.44	-30.56	-0.19	4.88	6.81
19	3	3	1	36	15	19.44	-30.62	-0.81	4.94	6.81
19	3	3	1	37	15	19.44	-30.5	0	4.88	6.81
19	3	3	1	38	16	19.37	-30.62	-0.44	4.94	6.81
19	3	3	1	39	16	19.37	-30.44	0.19	4.88	6.81
19	3	3	1	40	16	19.37	-30.5	0	4.94	6.88
19	3	3	1	41	16	19.31	-30.31	0.37	4.94	6.88
19	3	3	1	42	17	19.31	-30.5	0	4.94	6.88
19	3	3	1	43	17	19.31	-30.56	-0.31	4.94	6.88
19	3	3	1	44	17	19.31	-30.44	0.31	4.94	6.88
19	3	3	1	45	17	19.31	-30.44	0.06	4.94	6.94
19	3	3	1	46	18	19.31	-30.31	0.63	4.94	6.94
19	3	3	1	47	18	19.31	-30.37	0.37	4.94	6.94
19	3	3	1	48	18	19.31	-30.5	0	4.94	7
19	3	3	1	49	19	19.31	-30.31	0.63	4.94	7
19	3	3	1	50	19	19.25	-30.31	0.5	4.94	7
19	3	3	1	51	19	19.25	-30.19	0.94	4.94	7
19	3	3	1	52	19	19.25	-30.19	0.81	4.94	7.06
19	3	3	1	53	20	19.25	-30.19	0.56	4.94	7.06
19	3	3	1	54	20	19.19	-30.12	1.12	5	7.06
19	3	3	1	55	20	19.19	-30.19	0.94	5	7.13
19	3	3	1	56	20	19.19	-30.06	1.31	5	7.13
19	3	3	1	57	21	19.19	-30.06	1.19	5	7.13
19	3	3	1	58	21	19.19	-30.12	0.88	5	7.19
19	3	3	1	59	21	19.19	-30.06	1.44	5	7.19
19	3	3	2	0	21	19.19	-30.06	1.44	5	7.19
19	3	3	2	1	22	19.25	-30	1.75	5	7.25
19	3	3	2	2	22	19.25	-30	1.75	5	7.25
19	3	3	2	3	22	19.31	-29.94	1.63	5	7.31
19	3	3	2	4	22	19.31	-29.94	2	5	7.31
19	3	3	2	5	23	19.31	-29.94	1.94	5.06	7.38
19	3	3	2	6	23	19.37	-29.87	2.25	5.06	7.38
19	3	3	2	7	23	19.37	-29.87	2.19	5.06	7.38
19	3	3	2	8	27	19.37	-29.81	2	5.06	7.44
19	3	3	2	9	28	19.37	-29.81	2.38	5.06	7.44
19	3	3	2	10	28	19.37	-29.87	2.06	5.06	7.44
19	3	3	2	11	28	19.37	-29.87	2.44	5.06	7.44
19	3	3	2	12	29	19.37	-29.87	2.19	5.06	7.44
19	3	3	2	13	29	19.31	-29.87	2.19	5.06	7.5
19	3	3	2	14	29	19.31	-29.87	2.44	5.06	7.5
19	3	3	2	15	29	19.31	-29.87	2.13	5.06	7.5
19	3	3	2	16	30	19.31	-29.87	2.5	5.06	7.5
19	3	3	2	17	30	19.25	-30	2.25	5.06	7.5
19	3	3	2	18	30	19.25	-29.94	2.56	5.06	7.56
19	3	3	2	19	30	19.25	-29.87	2.75	5.13	7.56
19	3	3	2	20	31	19.25	-29.94	2.44	5.13	7.56
19	3	3	2	21	31	19.25	-29.87	2.75	5.13	7.56
19	3	3	2	22	31	19.25	-30	2.31	5.13	7.56
19	3	3	2	23	31	19.31	-29.94	2.63	5.13	7.56
19	3	3	2	24	32	19.31	-30.06	2.31	5.13	7.56
19	3	3	2	25	32	19.31	-30.12	2	5.13	7.56
19	3	3	2	26	32	19.31	-30	2.44	5.13	7.56
19	3	3	2	27	32	19.31	-30.12	2.06	5.13	7.56
19	3	3	2	28	33	19.31	-30.06	2.44	5.13	7.56
19	3	3	2	29	33	19.31	-30.25	1.94	5.13	7.56
19	3	3	2	30	33	19.31	-30.31	1.44	5.13	7.5
19	3	3	2	31	33	19.25	-30.25	2.06	5.13	7.5
19	3	3	2	32	34	19.25	-30.44	1.5	5.13	7.5
19	3	3	2	33	34	19.19	-30.31	2	5.13	7.5
19	3	3	2	34	34	19.19	-30.44	1.56	5.06	7.5
19	3	3	2	35	34	19.19	-30.37	1.56	5.13	7.44
19	3	3	2	36	35	19.19	-30.44	1.69	5.06	7.44
19	3	3	2	37	35	19.19	-30.56	1.19	5.06	7.44
19	3	3	2	38	35	19.12	-30.44	1.75	5.06	7.44
19	3	3	2	39	36	19.19	-30.56	1.12	5.06	7.38
19	3	3	2	40	36	19.19	-30.44	1.63	5.06	7.38
19	3	3	2	41	36	19.19	-30.62	1.12	5.06	7.38
19	3	3	2	42	36	19.19	-30.75	0.56	5.06	7.38
19	3	3	2	43	37	19.19	-30.56	1.25	5.06	7.31
19	3	3	2	44	37	19.12	-30.75	0.69	5.06	7.31
19	3	3	2	45	37	19.12	-30.56	1.25	5.06	7.31
19	3	3	2	46	37	19.12	-30.75	0.75	5	7.25
19	3	3	2	47	38	19.06	-30.81	0.19	5.06	7.25
19	3	3	2	48	38	19.06	-30.69	0.94	5	7.19
19	3	3	2	49	38	19.06	-30.81	0.19	5	7.19
19	3	3	2	50	38	19.06	-30.62	0.88	5	7.19
19	3	3	2	51	39	19.06	-30.81	0.25	5	7.13
19	3	3	2	52	39	19.06	-30.62	0.88	5	7.13
19	3	3	2	53	39	19.06	-30.75	0.31	5	7.13
19	3	3	2	54	39	19.06	-30.94	-0.31	5	7.13
19	3	3	2	55	40	19.12	-30.75	0.5	5	7.06
19	3	3	2	56	40	19.12	-30.87	-0.12	5	7.06
19	3	3	2	57	40	19.19	-30.69	0.56	5	7.06
19	3	3	2	58	40	19.19	-30.81	0	5	7.06
19	3	3	2	59	41	19.19	-30.94	-0.75	5	7.06
19	3	3	3	0	41	19.25	-30.81	0.12	5	7
19	3	3	3	1	41	19.25	-30.94	-0.75	5	7
19	3	3	3	2	42	19.31	-30.75	0.06	5	6.94
19	3	3	3	3	42	19.31	-30.94	-0.63	4.94	6.94
19	3	3	3	4	42	19.31	-30.69	0.06	4.94	6.94
19	3	3	3	5	42	19.31	-30.87	-0.56	4.94	6.88
19	3	3	3	6	43	19.31	-31	-1.31	4.94	6.88
19	3	3	3	7	43	19.31	-30.81	-0.31	4.94	6.88
19	3	3	3	8	43	19.31	-30.94	-1.06	4.94	6.88
19	3	3	3	9	43	19.31	-30.75	-0.25	4.94	6.81
19	3	3	3	10	44	19.25	-30.87	-0.75	4.94	6.81
19	3	3	3	11	44	19.25	-30.62	-0.19	4.94	6.81
19	3	3	3	12	44	19.19	-30.75	-0.5	4.94	6.75
19	3	3	3	13	44	19.19	-30.94	-1.19	4.94	6.75
19	3	3	3	14	45	19.19	-30.69	-0.37	4.94	6.75
19	3	3	3	15	45	19.19	-30.87	-0.94	4.94	6.75
19	3	3	3	16	45	19.19	-30.62	-0.25	4.94	6.75
19	3	3	3	17	45	19.19	-30.75	-0.69	4.88	6.75
19	3	3	3	18	46	19.19	-30.87	-1.31	4.88	6.75
19	3	3	3	19	46	19.19	-30.69	-0.37	4.94	6.75
19	3	3	3	20	46	19.19	-30.81	-0.81	4.88	6.75
19	3	3	3	21	46	19.19	-30.62	-0.12	4.88	6.75
19	3	3	3	22	47	19.25	-30.75	-0.56	4.94	6.75
19	3	3	3	23	47	19.25	-30.56	0	4.94	6.75
19	3	3	3	24	47	19.25	-30.69	-0.25	4.94	6.75
19	3	3	3	25	47	19.25	-30.81	-0.81	4.88	6.75
19	3	3	3	26	48	19.31	-30.62	-0.06	4.88	6.75
19	3	3	3	27	48	19.31	-30.75	-0.5	4.88	6.75
19	3	3	3	28	48	19.31	-30.56	0.06	4.94	6.81
19	3	3	3	29	49	19.25	-30.62	-0.19	4.88	6.81
19	3	3	3	30	49	19.25	-30.81	-0.75	4.94	6.81
19	3	3	3	31	49	19.25	-30.56	0.06	4.94	6.81
19	3	3	3	32	49	19.25	-30.69	-0.31	4.94	6.81
19	3	3	3	33	50	19.19	-30.5	0.31	4.94	6.81
19	3	3	3	34	50	19.19	-30.62	0	4.94	6.88
19	3	3	3	35	50	19.19	-30.62	-0.44	4.94	6.88
19	3	3	3	36	50	19.12	-30.56	0.25	4.94	6.88
19	3	3	3	37	51	19.12	-30.62	-0.19	4.94	6.88
19	3	3	3	38	51	19.06	-30.44	0.44	4.94	6.88
19	3	3	3	39	51	19.06	-30.5	0.25	4.94	6.88
19	3	3	3	40	51	19.06	-30.37	0.69	4.94	6.94
19	3	3	3	41	52	19.06	-30.37	0.63	4.94	6.94
19	3	3	3	42	52	19.06	-30.5	0.31	4.94	7
19	3	3	3	43	52	19.06	-30.37	0.88	4.94	7
19	3	3	3	44	52	19.06	-30.44	0.63	4.94	7
19	3	3	3	45	53	19.12	-30.31	1.06	4.94	7
19	3	3	3	46	53	19.12	-30.37	0.81	4.94	7.06
19	3	3	3	47	53	19.12	-30.44	0.37	4.94	7.06
19	3	3	3	48	54	19.19	-30.31	1	5	7.06
19	3	3	3	49	54	19.19	-30.31	0.88	5	7.06
19	3	3	3	50	54	19.19	-30.19	1.25	5	7.13
19	3	3	3	51	54	19.25	-30.25	1.19	5	7.13
19	3	3	3	52	55	19.25	-30.25	0.88	5	7.13
19	3	3	3	53	55	19.25	-30.19	1.44	5	7.13
19	3	3	3	54	55	19.25	-30.25	1.19	5	7.19
19	3	3	3	55	55	19.25	-30.12	1.63	5	7.19
19	3	3	3	56	56	19.25	-30.12	1.44	5	7.25
19	3	3	3	57	56	19.25	-30.12	1.19	5	7.25
19	3	3	3	58	56	19.19	-30.06	1.69	5.06	7.25
19	3	3	3	59	56	19.19	-30.06	1.63	5	7.25
19	3	3	4	0	57	19.19	-30	2	5	7.31
19	3	3	4	1	57	19.12	-30.06	1.94	5.06	7.31
19	3	3	4	2	57	19.12	-30	2	5.06	7.31
19	3	3	4	3	57	19.12	-30	2.19	5.06	7.38
19	3	3	4	4	58	19.06	-29.94	2.13	5.06	7.38
19	3	3	4	5	58	19.06	-29.94	2.44	5.06	7.44
19	3	3	4	6	58	19.06	-29.94	2.38	5.06	7.44
19	3	3	4	7	58	19.06	-29.87	2.63	5.06	7.44
19	3	3	4	8	59	19.06	-29.75	2.75	5.06	7.5
19	3	3	4	9	59	19.06	-29.75	2.56	5.06	7.5
19	3	3	4	10	59	19.12	-29.75	2.81	5.06	7.5
19	3	3	4	12	0	19.12	-29.81	2.56	5.06	7.56
19	3	3	4	13	0	19.12	-29.81	2.81	5.06	7.56
19	3	3	4	14	0	19.12	-29.94	2.56	5.06	7.56
19	3	3	4	15	0	19.12	-29.94	2.44	5.06	7.56
19	3	3	4	16	1	19.19	-29.94	2.81	5.06	7.56
19	3	3	4	17	1	19.12	-30	2.5	5.13	7.56
19	3	3	4	18	1	19.12	-29.94	2.81	5.06	7.56
19	3	3	4	19	1	19.12	-30	2.56	5.06	7.63
19	3	3	4	20	2	19.12	-30.06	2.19	5.13	7.56
19	3	3	4	21	2	19.06	-30.06	2.63	5.13	7.56
19	3	3	4	22	2	19.06	-30.19	2.25	5.06	7.56
19	3	3	4	23	2	19.06	-30.12	2.63	5.13	7.56
19	3	3	4	24	3	19	-30.25	2.25	5.13	7.56
19	3	3	4	25	3	19	-30.31	1.88	5.13	7.56
19	3	3	4	26	3	19	-30.25	2.38	5.13	7.56
19	3	3	4	27	3	19	-30.37	2	5.06	7.56
19	3	3	4	28	4	19	-30.25	2.44	5.13	7.56
19	3	3	4	29	4	19	-30.31	2.06	5.13	7.56
19	3	3	4	30	4	19	-30.31	1.75	5.13	7.56
19	3	3	4	31	4	19	-30.31	2.13	5.13	7.5
19	3	3	4	32	5	19	-30.44	1.56	5.13	7.5
19	3	3	4	33	5	19	-30.31	2.06	5.06	7.5
19	3	3	4	34	5	19	-30.44	1.63	5.06	7.5
19	3	3	4	35	5	19	-30.31	2.06	5.06	7.44
19	3	3	4	36	6	19.06	-30.44	1.63	5.06	7.44
19	3	3	4	37	6	19.06	-30.56	1.25	5.06	7.44
19	3	3	4	38	6	19.06	-30.44	1.81	5.06	7.44
19	3	3	4	39	7	19.06	-30.56	1.25	5.06	7.44
19	3	3	4	40	7	19.06	-30.44	1.75	5.06	7.38
19	3	3	4	41	7	19.06	-30.62	1.19	5.06	7.38
19	3	3	4	42	7	19.06	-30.69	0.56	5.06	7.38
19	3	3	4	43	8	19.06	-30.56	1.31	5.06	7.31
19	3	3	4	44	8	19.06	-30.69	0.75	5.06	7.31
19	3	3	4	45	8	19.06	-30.56	1.31	5.06	7.31
19	3	3	4	46	8	19.06	-30.69	0.81	5.06	7.31
19	3	3	4	47	9	19.06	-30.5	1.31	5.06	7.25
19	3	3	4	48	9	19.06	-30.62	0.94	5	7.25
19	3	3	4	49	9	19	-30.75	0.31	5	7.25
19	3	3	4	50	9	19	-30.56	1	5	7.25
19	3	3	4	51	10	19	-30.75	0.31	5	7.19
19	3	3	4	52	10	19	-30.56	0.94	5	7.19
19	3	3	4	53	10	18.94	-30.75	0.31	5	7.19
19	3	3	4	54	10	19	-30.81	-0.37	5	7.13
19	3	3	4	55	11	19	-30.69	0.5	5	7.13
19	3	3	4	56	11	19	-30.81	-0.19	5	7.13
19	3	3	4	57	11	19	-30.62	0.5	5	7.13
19	3	3	4	58	11	19	-30.75	-0.06	5	7.06
19	3	3	4	59	12	19	-30.56	0.56	5	7.06
19	3	3	5	0	12	19	-30.75	-0.06	5	7.06
19	3	3	5	1	12	19	-30.87	-0.81	5	7
19	3	3	5	2	13	19	-30.69	0.06	5	7
19	3	3	5	3	13	19	-30.81	-0.63	5	7
19	3	3	5	4	13	18.94	-30.62	0.12	5	7
19	3	3	5	5	13	18.94	-30.75	-0.56	4.94	6.94
19	3	3	5	6	14	18.87	-30.87	-1.37	4.94	6.94
19	3	3	5	7	14	18.87	-30.69	-0.37	4.94	6.88
19	3	3	5	8	14	18.87	-30.87	-1.19	4.94	6.88
19	3	3	5	9	14	18.87	-30.62	-0.31	4.94	6.88
19	3	3	5	10	15	18.87	-30.75	-0.81	4.94	6.81
19	3	3	5	11	15	18.87	-30.56	-0.12	4.94	6.81
19	3	3	5	12	15	18.87	-30.62	-0.56	4.94	6.81
19	3	3	5	13	15	18.87	-30.81	-1.25	4.94	6.81
19	3	3	5	14	16	18.87	-30.56	-0.37	4.94	6.81
19	3	3	5	15	16	18.87	-30.69	-1	4.94	6.81
19	3	3	5	16	16	18.94	-30.5	-0.25	4.94	6.81
19	3	3	5	17	16	18.94	-30.62	-0.75	4.94	6.81
19	3	3	5	18	17	18.94	-30.56	-1.06	4.94	6.81
19	3	3	5	19	17	18.94	-30.62	-0.56	4.94	6.75
19	3	3	5	20	17	18.87	-30.75	-1.06	4.88	6.81
19	3	3	5	21	17	18.87	-30.56	-0.25	4.88	6.81
19	3	3	5	22	18	18.87	-30.62	-0.69	4.94	6.75
19	3	3	5	23	18	18.87	-30.44	-0.06	4.94	6.75
19	3	3	5	24	18	18.81	-30.56	-0.37	4.88	6.81
19	3	3	5	25	18	18.81	-30.69	-1	4.94	6.81
19	3	3	5	26	19	18.81	-30.56	-0.19	4.94	6.81
19	3	3	5	27	19	18.75	-30.62	-0.63	4.94	6.81
19	3	3	5	28	19	18.81	-30.44	0	4.94	6.81
19	3	3	5	29	20	18.81	-30.56	-0.25	4.94	6.81
19	3	3	5	30	20	18.81	-30.62	-0.75	4.94	6.81
19	3	3	5	31	20	18.81	-30.44	0.06	4.94	6.81
19	3	3	5	32	20	18.87	-30.56	-0.37	4.94	6.88
19	3	3	5	33	21	18.87	-30.37	0.25	4.94	6.88
19	3	3	5	34	21	18.81	-30.5	-0.06	4.94	6.88
19	3	3	5	35	21	18.87	-30.37	0.44	4.94	6.88
19	3	3	5	36	21	18.87	-30.37	0.37	4.94	6.88
19	3	3	5	37	22	18.87	-30.44	-0.06	4.94	6.94
19	3	3	5	38	26	18.87	-30.31	0.56	4.94	6.94
19	3	3	5	39	26	18.87	-30.37	0.37	4.94	6.94
19	3	3	5	40	26	18.81	-30.19	0.81	4.94	6.94
19	3	3	5	41	27	18.81	-30.25	0.69	4.94	6.94
19	3	3	5	42	27	18.81	-30.37	0.31	4.94	7
19	3	3	5	43	27	18.81	-30.25	0.81	4.94	7
19	3	3	5	44	27	18.81	-30.31	0.63	4.94	7
19	3	3	5	45	28	18.81	-30.19	1.06	4.94	7.06
19	3	3	5	46	28	18.81	-30.25	0.81	5	7.06
19	3	3	5	47	28	18.81	-30.31	0.37	4.94	7.06
19	3	3	5	48	29	18.81	-30.19	1	5	7.06
19	3	3	5	49	29	18.87	-30.19	0.88	5	7.06
19	3	3	5	50	29	18.87	-30.12	1.31	5	7.13
19	3	3	5	51	29	18.87	-30.19	1.19	5	7.13
19	3	3	5	52	30	18.87	-30.19	0.88	5	7.13
19	3	3	5	53	30	18.87	-30.06	1.44	5	7.19
19	3	3	5	54	30	18.87	-30.12	1.25	5	7.19
19	3	3	5	55	30	18.81	-30	1.63	5	7.19
19	3	3	5	56	31	18.81	-30	1.5	5	7.19
19	3	3	5	57	31	18.81	-30	1.37	5	7.25
19	3	3	5	58	31	18.75	-29.94	1.69	5	7.25
19	3	3	5	59	31	18.75	-29.94	1.63	5	7.25
19	3	3	6	0	32	18.69	-29.87	2	5	7.31
19	3	3	6	1	32	18.69	-29.87	2	5	7.31
19	3	3	6	2	32	18.69	-29.81	2.25	5.06	7.31
19	3	3	6	3	32	18.69	-30	1.88	5.06	7.31
19	3	3	6	4	33	18.69	-30	1.81	5.06	7.31
19	3	3	6	5	33	18.69	-29.94	2.13	5.06	7.31
19	3	3	6	6	33	18.69	-29.94	2.13	5.06	7.38
19	3	3	6	7	33	18.75	-29.81	2.38	5.06	7.38
19	3	3	6	8	34	18.75	-29.81	2.38	5.06	7.38
19	3	3	6	9	34	18.75	-29.75	2.25	5.06	7.44
19	3	3	6	10	34	18.75	-29.75	2.56	5.06	7.44
19	3	3	6	11	34	18.75	-29.81	2.31	5.06	7.44
19	3	3	6	12	35	18.75	-29.75	2.63	5.06	7.44
19	3	3	6	13	35	18.75	-29.81	2.44	5.06	7.44
19	3	3	6	14	35	18.75	-29.87	2.31	5.06	7.5
19	3	3	6	15	35	18.75	-29.81	2.63	5.06	7.5
19	3	3	6	16	36	18.75	-29.87	2.44	5.06	7.5
19	3	3	6	17	36	18.75	-29.81	2.69	5.06	7.5
19	3	3	6	18	36	18.69	-29.87	2.5	5.06	7.5
19	3	3	6	19	37	18.69	-30	2.19	5.06	7.5
19	3	3	6	20	37	18.69	-29.94	2.56	5.06	7.5
19	3	3	6	21	37	18.69	-30.06	2.19	5.06	7.5
19	3	3	6	22	37	18.69	-30	2.56	5.06	7.5
19	3	3	6	23	38	18.69	-30.12	2.19	5.06	7.5
19	3	3	6	24	38	18.75	-30.12	1.81	5.06	7.5
19	3	3	6	25	38	18.75	-30.06	2.38	5.06	7.5
19	3	3	6	26	38	18.75	-30.12	1.94	5.06	7.5
19	3	3	6	27	39	18.81	-30.06	2.38	5.06	7.5
19	3	3	6	28	39	18.81	-30.19	2.06	5.06	7.5
19	3	3	6	29	39	18.87	-30.06	2.31	5.06	7.5
19	3	3	6	30	39	18.87	-30.25	1.81	5.06	7.44
19	3	3	6	31	40	18.87	-30.37	1.44	5.06	7.44
19	3	3	6	32	40	18.87	-30.25	1.94	5.06	7.44
19	3	3	6	33	40	18.94	-30.37	1.5	5.06	7.44
19	3	3	6	34	40	18.94	-30.25	2	5.06	7.38
19	3	3	6	35	41	18.94	-30.37	1.56	5.06	7.38
19	3	3	6	36	41	18.94	-30.44	1.12	5.06	7.38
19	3	3	6	37	41	18.94	-30.31	1.75	5.06	7.31
19	3	3	6	38	42	18.94	-30.37	1.31	5.06	7.38
19	3	3	6	39	42	18.87	-30.25	1.75	5.06	7.31
19	3	3	6	40	42	18.87	-30.44	1.25	5.06	7.31
19	3	3	6	41	42	18.87	-30.5	0.56	5.06	7.31
19	3	3	6	42	43	18.81	-30.37	1.37	5	7.31
19	3	3	6	43	43	18.81	-30.56	0.75	5	7.25
19	3	3	6	44	43	18.81	-30.44	1.31	5	7.25
19	3	3	6	45	43	18.81	-30.5	0.81	5	7.25
19	3	3	6	46	44	18.81	-30.37	1.37	5	7.25
19	3	3	6	47	44	18.81	-30.5	1	5	7.19
19	3	3	6	48	44	18.81	-30.62	0.37	5	7.25
19	3	3	6	49	44	18.81	-30.44	1.06	5	7.19
19	3	3	6	50	45	18.81	-30.56	0.44	5	7.19
19	3	3	6	51	45	18.81	-30.44	1.06	5	7.19
19	3	3	6	52	45	18.81	-30.56	0.44	5	7.13
19	3	3	6	53	45	18.81	-30.69	-0.31	5	7.13
19	3	3	6	54	46	18.81	-30.5	0.56	5	7.13
19	3	3	6	55	46	18.81	-30.62	-0.06	5	7.06
19	3	3	6	56	46	18.81	-30.44	0.63	5	7.06
19	3	3	6	57	46	18.81	-30.62	0.06	5	7.06
19	3	3	6	58	47	18.75	-30.44	0.69	5	7.06
19	3	3	6	59	47	18.75	-30.56	0.12	5	7
19	3	3	7	0	47	18.69	-30.69	-0.63	5	7
19	3	3	7	1	47	18.69	-30.5	0.25	4.94	7
19	3	3	7	2	48	18.62	-30.62	-0.5	4.94	7
19	3	3	7	3	48	18.62	-30.44	0.25	4.94	7
19	3	3	7	4	48	18.62	-30.56	-0.37	4.94	6.94
19	3	3	7	5	49	18.62	-30.5	-0.69	4.94	6.94
19	3	3	7	6	49	18.62	-30.5	-0.19	4.94	6.94
19	3	3	7	7	49	18.69	-30.69	-1	4.94	6.94
19	3	3	7	8	49	18.69	-30.44	-0.19	4.94	6.88
19	3	3	7	9	50	18.69	-30.56	-0.63	4.94	6.88
19	3	3	7	10	50	18.75	-30.37	0.06	4.94	6.88
19	3	3	7	11	50	18.75	-30.5	-0.44	4.94	6.88
19	3	3	7	12	50	18.75	-30.69	-1.12	4.94	6.88
19	3	3	7	13	51	18.81	-30.44	-0.19	4.94	6.88
19	3	3	7	14	51	18.81	-30.56	-0.75	4.94	6.88
19	3	3	7	15	51	18.81	-30.37	-0.06	4.94	6.81
19	3	3	7	16	51	18.81	-30.5	-0.5	4.94	6.88
19	3	3	7	17	52	18.81	-30.31	0	4.94	6.88
19	3	3	7	18	52	18.81	-30.44	-0.25	4.94	6.81
19	3	3	7	19	52	18.81	-30.56	-0.69	4.94	6.88
19	3	3	7	20	52	18.81	-30.37	0	4.88	6.81
19	3	3	7	21	53	18.75	-30.44	-0.37	4.94	6.81
19	3	3	7	22	53	18.75	-30.25	0.19	4.94	6.88
19	3	3	7	23	53	18.75	-30.37	-0.12	4.94	6.81
19	3	3	7	24	53	18.75	-30.5	-0.69	4.94	6.88
19	3	3	7	25	54	18.75	-30.31	0.06	4.94	6.88
19	3	3	7	26	54	18.75	-30.44	-0.37	4.94	6.88
19	3	3	7	27	54	18.81	-30.25	0.25	4.94	6.88
19	3	3	7	28	55	18.81	-30.37	-0.12	4.94	6.88
19	3	3	7	29	55	18.81	-30.25	0	4.94	6.88
19	3	3	7	30	55	18.81	-30.37	-0.12	4.94	6.88
19	3	3	7	31	55	18.87	-30.5	-0.5	4.94	6.88
19	3	3	7	32	56	18.87	-30.31	0.19	4.94	6.88
19	3	3	7	33	56	18.87	-30.37	-0.12	4.94	6.88
19	3	3	7	34	56	18.81	-30.19	0.44	4.94	6.88
19	3	3	7	35	56	18.81	-30.31	0.19	4.94	6.88
19	3	3	7	36	57	18.81	-30.37	-0.25	4.94	6.88
19	3	3	7	37	57	18.81	-30.25	0.44	4.94	6.88
19	3	3	7	38	57	18.81	-30.25	0.37	4.94	6.94
19	3	3	7	39	57	18.81	-30.12	0.81	4.94	6.94
19	3	3	7	40	58	18.75	-30.19	0.63	4.94	6.94
19	3	3	7	41	58	18.75	-30.19	0.25	4.94	7
19	3	3	7	42	58	18.81	-30.06	0.88	4.94	7
19	3	3	7	43	58	18.81	-30.12	0.63	4.94	7
19	3	3	7	44	59	18.81	-30	1.06	4.94	7
19	3	3	7	45	59	18.81	-30.12	0.81	4.94	7
19	3	3	7	46	59	18.87	-30.06	0.56	4.94	7
19	3	3	7	47	59	18.87	-30.06	1.06	4.94	7.06
19	3	3	7	49	0	18.87	-30.06	0.88	5	7.06
19	3	3	7	50	0	18.87	-30	1.37	4.94	7.06
19	3	3	7	51	0	18.94	-30	1.19	5	7.13
19	3	3	7	52	0	18.94	-29.94	1.56	5	7.13
19	3	3	7	53	1	19	-29.75	1.88	5	7.13
19	3	3	7	54	1	19	-29.75	1.56	5	7.19
19	3	3	7	55	1	19	-29.75	1.94	5	7.19
19	3	3	7	56	2	19.06	-29.87	1.75	5	7.19
19	3	3	7	57	2	19	-29.81	2.06	5	7.19
19	3	3	7	58	2	19	-29.87	1.88	5	7.25
19	3	3	7	59	2	19	-29.87	1.81	5	7.25
19	3	3	8	0	3	19	-29.81	2.19	5	7.25
19	3	3	8	1	3	19	-29.81	2.13	5	7.31
19	3	3	8	2	3	19	-29.75	2.38	5	7.31
19	3	3	8	3	3	19.06	-29.75	2.31	5	7.38
19	3	3	8	4	4	19.06	-29.69	2.25	5.06	7.38
19	3	3	8	5	4	19.06	-29.69	2.56	5	7.38
19	3	3	8	6	4	19.12	-29.69	2.44	5	7.44
19	3	3	8	7	4	19.12	-29.62	2.69	5.06	7.44
19	3	3	8	8	5	19.19	-29.62	2.63	5.06	7.44
19	3	3	8	9	5	19.25	-29.62	2.5	5.06	7.44
19	3	3	8	10	5	19.25	-29.56	2.81	5.06	7.5
19	3	3	8	11	5	19.31	-29.69	2.56	5.06	7.5
19	3	3	8	12	6	19.31	-29.62	2.81	5.06	7.5
19	3	3	8	13	6	19.37	-29.75	2.56	5.06	7.5
19	3	3	8	14	6	19.37	-29.69	2.44	5.06	7.5
19	3	3	8	15	6	19.37	-29.69	2.81	5.06	7.5
19	3	3	8	16	7	19.37	-29.75	2.5	5.06	7.5
19	3	3	8	17	7	19.37	-29.62	2.81	5.06	7.56
19	3	3	8	18	7	19.37	-29.75	2.63	5.06	7.56
19	3	3	8	19	8	19.37	-29.75	2.19	5.06	7.56
19	3	3	8	20	8	19.37	-29.81	2.69	5.06	7.56
19	3	3	8	21	8	19.37	-29.94	2.25	5.06	7.56
19	3	3	8	22	8	19.37	-29.81	2.63	5.06	7.56
19	3	3	8	23	9	19.31	-29.87	2.31	5.06	7.56
19	3	3	8	24	9	19.31	-29.87	2.63	5.06	7.56
19	3	3	8	25	9	19.31	-30	2.31	5.06	7.5
19	3	3	8	26	9	19.37	-30.06	2.06	5.06	7.5
19	3	3	8	27	10	19.37	-30	2.44	5.06	7.5
19	3	3	8	28	10	19.37	-30.06	2.13	5.06	7.5
19	3	3	8	29	10	19.37	-29.94	2.5	5.06	7.5
19	3	3	8	30	10	19.44	-30.12	2.06	5.06	7.5
19	3	3	8	31	11	19.44	-30.19	1.63	5.06	7.44
19	3	3	8	32	11	19.44	-30.06	2.19	5.06	7.44
19	3	3	8	33	11	19.5	-30.19	1.75	5.06	7.44
19	3	3	8	34	11	19.5	-30.06	2.19	5.06	7.44
19	3	3	8	35	12	19.5	-30.19	1.81	5.06	7.44
19	3	3	8	36	12	19.5	-30.25	1.25	5.06	7.44
19	3	3	8	37	12	19.5	-30.19	1.94	5.06	7.44
19	3	3	8	38	12	19.5	-30.31	1.44	5.06	7.44
19	3	3	8	39	13	19.5	-30.12	1.94	5.06	7.38
19	3	3	8	40	13	19.44	-30.25	1.44	5.06	7.38
19	3	3	8	41	13	19.44	-30.12	1.88	5.06	7.38
19	3	3	8	42	14	19.44	-30.19	1.5	5.06	7.38
19	3	3	8	43	14	19.44	-30.31	1	5.06	7.31
19	3	3	8	44	14	19.44	-30.19	1.63	5.06	7.31
19	3	3	8	45	14	19.44	-30.31	1.12	5.06	7.31
19	3	3	8	46	15	19.5	-30.19	1.63	5.06	7.31
19	3	3	8	47	15	19.5	-30.31	1.12	5	7.25
19	3	3	8	48	15	19.5	-30.37	0.56	5.06	7.25
19	3	3	8	49	15	19.56	-30.25	1.31	5	7.25
19	3	3	8	50	16	19.56	-30.37	0.63	5	7.25
19	3	3	8	51	16	19.62	-30.25	1.25	5	7.25
19	3	3	8	52	16	19.62	-30.37	0.63	5	7.19
19	3	3	8	53	16	19.69	-30.25	0.75	5	7.19
19	3	3	8	54	17	19.69	-30.37	0.5	5	7.19
19	3	3	8	55	17	19.75	-30.5	0	5	7.13
19	3	3	8	56	17	19.75	-30.25	0.69	5	7.13
19	3	3	8	57	17	19.75	-30.44	0.12	5	7.13
19	3	3	8	58	18	19.75	-30.25	0.75	5	7.13
19	3	3	8	59	18	19.75	-30.37	0.31	5	7.13
19	3	3	9	0	18	19.75	-30.5	-0.5	5	7.06
19	3	3	9	1	18	19.75	-30.31	0.44	5	7.06
19	3	3	9	2	19	19.75	-30.44	-0.31	5	7.06
19	3	3	9	3	19	19.69	-30.25	0.44	5	7
19	3	3	9	4	19	19.69	-30.37	-0.19	5	7
19	3	3	9	5	20	19.69	-30.19	0.44	5	7
19	3	3	9	6	20	19.62	-30.31	0	4.94	6.94
19	3	3	9	7	20	19.62	-30.37	-0.69	5	6.94
19	3	3	9	8	20	19.62	-30.19	0.19	4.94	6.94
19	3	3	9	9	21	19.62	-30.31	-0.37	4.94	6.94
19	3	3	9	10	21	19.62	-30.12	0.31	4.94	6.94
19	3	3	9	11	21	19.69	-30.31	-0.19	4.94	6.94
19	3	3	9	12	21	19.69	-30.37	-0.81	4.94	6.94
19	3	3	9	13	22	19.69	-30.25	0	4.94	6.94
19	3	3	9	14	22	19.75	-30.37	-0.56	4.94	6.94
19	3	3	9	15	22	19.75	-30.19	0.19	4.94	6.88
19	3	3	9	16	22	19.75	-30.31	-0.37	4.94	6.88
19	3	3	9	17	23	19.75	-30.12	0.31	4.94	6.88
19	3	3	9	18	26	19.75	-30.25	-0.12	4.94	6.88
19	3	3	9	19	27	19.75	-30.37	-0.56	4.94	6.88
19	3	3	9	20	27	19.75	-30.19	0.12	4.94	6.88
19	3	3	9	21	27	19.75	-30.25	-0.25	4.94	6.88
19	3	3	9	22	27	19.75	-30.06	0.37	4.94	6.88
19	3	3	9	23	28	19.75	-30.19	0	4.94	6.88
19	3	3	9	24	28	19.75	-30.31	-0.69	4.94	6.88
19	3	3	9	25	28	19.75	-30.19	0.19	4.94	6.88
19	3	3	9	26	28	19.75	-30.31	-0.25	4.94	6.94
19	3	3	9	27	29	19.75	-30.12	0.37	4.94	6.94
19	3	3	9	28	29	19.81	-30.25	0	4.94	6.94
19	3	3	9	29	29	19.81	-30.06	0.56	4.94	6.94
19	3	3	9	30	30	19.87	-30.12	0.37	4.94	6.94
19	3	3	9	31	30	19.87	-30.19	0	4.94	6.94
19	3	3	9	32	30	19.87	-30	0.63	4.94	6.94
19	3	3	9	33	30	19.87	-30.12	0.31	5	7
19	3	3	9	34	31	19.87	-29.94	0.81	5	7
19	3	3	9	35	31	19.87	-30.06	0.56	4.94	7
19	3	3	9	36	31	19.87	-30.12	0.06	5	7
19	3	3	9	37	31	19.87	-30	0.75	5	7
19	3	3	9	38	32	19.87	-30.06	0.37	5	7
19	3	3	9	39	32	19.87	-29.94	0.88	5	7
19	3	3	9	40	32	19.81	-30	0.75	5	7.06
19	3	3	9	41	32	19.81	-30	0.31	5	7.06
19	3	3	9	42	33	19.81	-29.94	1	5	7.06
19	3	3	9	43	33	19.81	-30	0.75	5	7.06
19	3	3	9	44	33	19.81	-29.87	1.19	5	7.13
19	3	3	9	45	33	19.81	-29.94	1	5	7.13
19	3	3	9	46	34	19.87	-29.81	1.37	5	7.13
19	3	3	9	47	34	19.87	-30	1	5	7.13
19	3	3	9	48	34	19.87	-30	0.69	5	7.13
19	3	3	9	49	35	19.94	-29.87	1.25	5	7.13
19	3	3	9	50	35	20	-29.94	1.12	5	7.13
19	3	3	9	51	35	20.06	-29.81	1.5	5	7.19
19	3	3	9	52	35	20.06	-29.81	1.44	5	7.19
19	3	3	9	53	36	20.12	-29.87	1.25	5	7.19
19	3	3	9	54	36	20.12	-29.81	1.69	5	7.19
19	3	3	9	55	36	20.12	-29.81	1.56	5	7.25
19	3	3	9	56	36	20.12	-29.69	1.88	5	7.25
19	3	3	9	57	37	20.12	-29.75	1.75	5	7.31
19	3	3	9	58	37	20.12	-29.69	1.69	5	7.31
19	3	3	9	59	37	20.12	-29.62	2.06	5.06	7.31
19	3	3	10	0	37	20.12	-29.56	2.06	5.06	7.38
19	3	3	10	1	38	20.12	-29.5	2.38	5.06	7.38
19	3	3	10	2	38	20.12	-29.5	2.31	5.06	7.38
19	3	3	10	3	38	20.12	-29.44	2.19	5.06	7.44
19	3	3	10	4	38	20.12	-29.44	2.56	5.06	7.44
19	3	3	10	5	39	20.12	-29.44	2.44	5.06	7.5
19	3	3	10	6	39	20.19	-29.37	2.75	5.06	7.5
19	3	3	10	7	39	20.19	-29.44	2.69	5.06	7.5
19	3	3	10	8	39	20.19	-29.44	2.5	5.06	7.56
19	3	3	10	9	40	20.25	-29.44	2.88	5.06	7.56
19	3	3	10	10	40	20.25	-29.5	2.56	5.13	7.56
19	3	3	10	11	40	20.31	-29.5	2.88	5.13	7.56
19	3	3	10	12	41	20.37	-29.56	2.63	5.13	7.56
19	3	3	10	13	41	20.37	-29.5	2.5	5.13	7.56
19	3	3	10	14	41	20.37	-29.5	2.88	5.06	7.63
19	3	3	10	15	41	20.37	-29.56	2.63	5.13	7.63
19	3	3	10	16	42	20.37	-29.5	2.94	5.13	7.63
19	3	3	10	17	42	20.37	-29.56	2.69	5.13	7.63
19	3	3	10	18	42	20.37	-29.5	2.94	5.13	7.63
19	3	3	10	19	42	20.37	-29.69	2.44	5.13	7.63
19	3	3	10	20	43	20.37	-29.75	2.25	5.13	7.63
19	3	3	10	21	43	20.37	-29.62	2.63	5.13	7.63
19	3	3	10	22	43	20.31	-29.81	2.25	5.13	7.63
19	3	3	10	23	43	20.31	-29.69	2.63	5.13	7.63
19	3	3	10	24	44	20.31	-29.75	2.31	5.13	7.63
19	3	3	10	25	44	20.31	-29.81	2.06	5.13	7.63
19	3	3	10	26	44	20.31	-29.75	2.56	5.13	7.63
19	3	3	10	27	44	20.31	-29.87	2.19	5.13	7.63
19	3	3	10	28	45	20.37	-29.75	2.56	5.13	7.63
19	3	3	10	29	45	20.44	-29.87	2.13	5.13	7.56
19	3	3	10	30	45	20.44	-29.94	1.69	5.13	7.56
19	3	3	10	31	45	20.5	-29.81	2.25	5.06	7.56
19	3	3	10	32	46	20.5	-29.94	1.75	5.13	7.56
19	3	3	10	33	46	20.56	-29.81	2.25	5.13	7.56
19	3	3	10	34	46	20.56	-29.94	1.81	5.13	7.56
19	3	3	10	35	47	20.62	-29.87	1.69	5.06	7.5
19	3	3	10	36	47	20.62	-29.87	2	5.13	7.5
19	3	3	10	37	47	20.62	-30	1.5	5.06	7.5
19	3	3	10	38	47	20.62	-29.81	2.06	5.13	7.5
19	3	3	10	39	48	20.62	-30	1.5	5.13	7.5
19	3	3	10	40	48	20.62	-29.87	2	5.06	7.5
19	3	3	10	41	48	20.62	-30	1.5	5.06	7.44
19	3	3	10	42	48	20.62	-30.06	1	5.06	7.44
19	3	3	10	43	49	20.62	-29.94	1.63	5.06	7.44
19	3	3	10	44	49	20.62	-30.06	1.12	5.06	7.38
19	3	3	10	45	49	20.56	-29.94	1.69	5.06	7.38
19	3	3	10	46	49	20.56	-30.06	1.19	5.06	7.38
19	3	3	10	47	50	20.56	-30.12	0.63	5.06	7.38
19	3	3	10	48	50	20.56	-30	1.37	5.06	7.38
19	3	3	10	49	50	20.62	-30.19	0.69	5.06	7.38
19	3	3	10	50	50	20.62	-30	1.31	5.06	7.31
19	3	3	10	51	51	20.62	-30.12	0.69	5.06	7.31
19	3	3	10	52	51	20.62	-29.94	1.25	5.06	7.31
19	3	3	10	53	51	20.69	-30.06	0.75	5.06	7.31
19	3	3	10	54	51	20.69	-30.19	0.19	5.06	7.25
19	3	3	10	55	52	20.75	-30	0.94	5.06	7.25
19	3	3	10	56	52	20.75	-30.12	0.37	5.06	7.25
19	3	3	10	57	52	20.75	-29.94	1	5.06	7.25
19	3	3	10	58	53	20.75	-30.06	0.5	5	7.25
19	3	3	10	59	53	20.75	-30.19	-0.31	5.06	7.25
19	3	3	11	0	53	20.75	-30	0.56	5	7.19
19	3	3	11	1	53	20.75	-30.19	-0.19	5.06	7.19
19	3	3	11	2	54	20.75	-30	0.56	5.06	7.19
19	3	3	11	3	54	20.75	-30.19	-0.12	5	7.19
19	3	3	11	4	54	20.69	-29.94	0.56	5	7.13
19	3	3	11	5	54	20.69	-30.06	0	5	7.13
19	3	3	11	6	55	20.69	-30.19	-0.69	5	7.13
19	3	3	11	7	55	20.62	-30	0.19	5	7.13
19	3	3	11	8	55	20.62	-30.19	-0.69	5	7.06
19	3	3	11	9	55	20.62	-29.94	0.12	5	7.06
19	3	3	11	10	56	20.62	-30.12	-0.56	5	7.06
19	3	3	11	11	56	20.62	-29.94	0.12	5	7.06
19	3	3	11	12	56	20.62	-30.06	-0.44	5	7
19	3	3	11	13	56	20.69	-30.25	-1.19	5	7
19	3	3	11	14	57	20.69	-29.94	-0.31	5	7
19	3	3	11	15	57	20.69	-30.12	-1	5	7
19	3	3	11	16	57	20.69	-29.94	-0.25	5	6.94
19	3	3	11	17	57	20.69	-30.06	-0.81	4.94	6.94
19	3	3	11	18	58	20.69	-30	-1.19	5	6.94
19	3	3	11	19	58	20.69	-30	-0.63	4.94	6.88
19	3	3	11	20	58	20.69	-30.12	-1.19	4.94	6.94
19	3	3	11	21	59	20.62	-29.87	-0.37	4.94	6.88
19	3	3	11	22	59	20.69	-30	-0.94	4.94	6.88
19	3	3	11	23	59	20.69	-29.81	-0.25	4.94	6.88
19	3	3	11	24	59	20.69	-29.94	-0.69	4.94	6.88
19	3	3	11	26	0	20.69	-30.06	-1.56	4.94	6.88
19	3	3	11	27	0	20.69	-29.87	-0.56	4.94	6.88
19	3	3	11	28	0	20.69	-30	-1.19	4.94	6.88
19	3	3	11	29	0	20.69	-29.81	-0.37	4.94	6.88
19	3	3	11	30	1	20.69	-29.87	-0.63	4.94	6.88
19	3	3	11	31	1	20.69	-29.69	-0.06	4.94	6.88
19	3	3	11	32	1	20.75	-29.75	-0.06	4.94	6.88
19	3	3	11	33	1	20.75	-29.87	-0.75	4.94	6.94
19	3	3	11	34	2	20.81	-29.69	0.06	4.94	6.94
19	3	3	11	35	2	20.81	-29.81	-0.44	4.94	6.94
19	3	3	11	36	2	20.81	-29.62	0.12	4.94	6.94
19	3	3	11	37	2	20.81	-29.81	-0.25	4.94	6.94
19	3	3	11	38	3	20.81	-29.94	-1.06	4.94	6.94
19	3	3	11	39	3	20.81	-29.75	-0.12	4.94	6.94
19	3	3	11	40	3	20.81	-29.87	-0.56	4.94	7
19	3	3	11	41	4	20.75	-29.69	0.12	4.94	7
19	3	3	11	42	4	20.75	-29.81	-0.19	4.94	7
19	3	3	11	43	4	20.75	-29.69	0.31	5	7
19	3	3	11	44	4	20.69	-29.75	0.12	5	7
19	3	3	11	45	5	20.69	-29.87	-0.37	5	7
19	3	3	11	46	5	20.69	-29.69	0.31	5	7
19	3	3	11	47	5	20.62	-29.81	-0.06	5	7
19	3	3	11	48	5	20.62	-29.62	0.5	5	7.06
19	3	3	11	49	6	20.62	-29.75	0.19	5	7.06
19	3	3	11	50	6	20.62	-29.81	-0.31	5	7.06
19	3	3	11	51	6	20.62	-29.69	0.44	5	7.06
19	3	3	11	52	6	20.62	-29.81	0.12	5	7.06
19	3	3	11	53	7	20.69	-29.62	0.69	5	7.06
19	3	3	11	54	7	20.69	-29.75	0.44	5	7.13
19	3	3	11	55	7	20.75	-29.56	0.81	5	7.13
19	3	3	11	56	7	20.75	-29.69	0.44	5	7.13
19	3	3	11	57	8	20.81	-29.81	0.12	5	7.13
19	3	3	11	58	8	20.81	-29.69	0.69	5	7.13
19	3	3	11	59	8	20.81	-29.69	0.5	5	7.13
19	3	3	12	0	9	20.87	-29.56	1	5	7.19
19	3	3	12	1	9	20.87	-29.62	0.81	5	7.19
19	3	3	12	2	9	20.94	-29.69	0.5	5	7.19
19	3	3	12	3	9	20.94	-29.56	1.06	5.06	7.19
19	3	3	12	4	10	20.94	-29.56	0.88	5	7.25
19	3	3	12	5	10	20.94	-29.44	1.31	5.06	7.25
19	3	3	12	6	10	20.87	-29.5	1.12	5	7.25
19	3	3	12	7	10	20.87	-29.62	0.69	5.06	7.25
19	3	3	12	8	11	20.87	-29.5	1.31	5.06	7.31
19	3	3	12	9	11	20.87	-29.62	0.94	5.06	7.31
19	3	3	12	10	11	20.81	-29.44	1.37	5.06	7.31
19	3	3	12	11	11	20.81	-29.62	1.06	5	7.31
19	3	3	12	12	12	20.81	-29.5	1.25	5.06	7.31
19	3	3	12	13	12	20.81	-29.62	1	5.06	7.31
19	3	3	12	14	12	20.75	-29.75	0.56	5.06	7.31
19	3	3	12	15	12	20.75	-29.62	1.19	5.06	7.31
19	3	3	12	16	13	20.75	-29.69	0.88	5.06	7.31
19	3	3	12	17	13	20.69	-29.56	1.31	5.06	7.31
19	3	3	12	18	13	20.69	-29.62	1.06	5.06	7.31
19	3	3	12	19	13	20.69	-29.75	0.63	5.06	7.31
19	3	3	12	20	14	20.62	-29.56	1.25	5.06	7.31
19	3	3	12	21	14	20.62	-29.69	0.75	5.06	7.31
19	3	3	12	22	14	20.62	-29.56	1.25	5.06	7.31
19	3	3	12	23	15	20.62	-29.75	0.81	5.06	7.31
19	3	3	12	24	15	20.62	-29.81	0.31	5.06	7.31
19	3	3	12	25	15	20.56	-29.69	1.06	5	7.31
19	3	3	12	26	15	20.56	-29.81	0.56	5.06	7.31
19	3	3	12	27	16	20.56	-29.69	1.12	5.06	7.31
19	3	3	12	28	16	20.62	-29.81	0.69	5.06	7.31
19	3	3	12	29	16	20.62	-29.62	1.19	5.06	7.31
19	3	3	12	30	16	20.62	-29.69	0.94	5.06	7.31
19	3	3	12	31	17	20.69	-29.81	0.37	5.06	7.31
19	3	3	12	32	17	20.75	-29.62	1	5.06	7.31
19	3	3	12	33	17	20.75	-29.81	0.44	5.06	7.31
19	3	3	12	34	17	20.81	-29.69	1	5	7.31
19	3	3	12	35	18	20.81	-29.81	0.5	5.06	7.25
19	3	3	12	36	18	20.87	-29.87	-0.12	5	7.25
19	3	3	12	37	18	20.87	-29.75	0.69	5	7.25
19	3	3	12	38	18	20.87	-29.87	0.12	5	7.25
19	3	3	12	39	19	20.87	-29.69	0.75	5	7.25
19	3	3	12	40	19	20.87	-29.81	0.12	5	7.25
19	3	3	12	41	19	20.87	-29.62	0.75	5	7.25
19	3	3	12	42	20	20.87	-29.75	0.19	5	7.19
19	3	3	12	43	20	20.87	-29.87	-0.44	5	7.19
19	3	3	12	44	20	20.87	-29.69	0.37	5	7.19
19	3	3	12	45	20	20.87	-29.81	-0.25	5	7.19
19	3	3	12	46	21	20.87	-29.62	0.37	5	7.19
19	3	3	12	47	21	20.87	-29.81	-0.12	5	7.13
19	3	3	12	48	25	20.87	-29.94	-1.5	5	7.13
19	3	3	12	49	25	20.87	-29.81	-0.94	5	7.13
19	3	3	12	50	26	20.87	-29.94	-1.63	5	7.06
19	3	3	12	51	26	20.87	-29.69	-0.63	5	7.06
19	3	3	12	52	26	20.87	-29.87	-1.25	4.94	7.06
19	3	3	12	53	26	20.94	-29.62	-0.44	4.94	7
19	3	3	12	54	27	20.94	-29.75	-0.94	4.94	7
19	3	3	12	55	27	21	-29.87	-1.75	4.94	7
19	3	3	12	56	27	21.06	-29.62	-0.63	4.94	6.94
19	3	3	12	57	27	21.06	-29.81	-1.31	4.94	6.94
19	3	3	12	58	28	21.12	-29.56	-0.5	4.94	6.94
19	3	3	12	59	28	21.19	-29.75	-0.94	4.94	6.94
19	3	3	13	0	28	21.25	-29.5	-0.25	4.94	6.94
19	3	3	13	1	28	21.31	-29.62	-0.69	4.94	6.94
19	3	3	13	2	29	21.31	-29.81	-1.44	4.94	6.94
19	3	3	13	3	29	21.37	-29.56	-0.5	4.94	6.94
19	3	3	13	4	29	21.37	-29.69	-1.06	4.94	6.94
19	3	3	13	5	30	21.37	-29.5	-0.31	4.94	6.94
19	3	3	13	6	30	21.37	-29.62	-0.75	4.94	6.94
19	3	3	13	7	30	21.37	-29.44	-0.37	4.94	6.94
19	3	3	13	8	30	21.37	-29.56	-0.56	4.94	6.94
19	3	3	13	9	31	21.37	-29.69	-1.06	4.94	6.94
19	3	3	13	10	31	21.37	-29.5	-0.25	4.94	6.94
19	3	3	13	11	31	21.37	-29.62	-0.69	4.94	6.94
19	3	3	13	12	31	21.37	-29.37	-0.06	4.94	6.94
19	3	3	13	13	32	21.37	-29.5	-0.44	4.94	6.94
19	3	3	13	14	32	21.37	-29.62	-1.12	4.94	6.94
19	3	3	13	15	32	21.37	-29.44	-0.19	4.94	6.94
19	3	3	13	16	32	21.37	-29.62	-0.69	4.94	6.94
19	3	3	13	17	33	21.37	-29.44	0	4.94	6.94
19	3	3	13	18	33	21.37	-29.56	-0.44	4.94	6.94
19	3	3	13	19	33	21.37	-29.37	0.12	4.94	6.94
19	3	3	13	20	33	21.37	-29.31	0.19	4.94	7
19	3	3	13	21	34	21.37	-29.5	-0.31	5	7
19	3	3	13	22	34	21.31	-29.31	0.37	4.94	7
19	3	3	13	23	34	21.31	-29.5	-0.06	5	7
19	3	3	13	24	34	21.37	-29.31	0.5	5	7.06
19	3	3	13	25	35	21.37	-29.44	0.12	5	7.06
19	3	3	13	26	35	21.37	-29.5	-0.5	5	7.06
19	3	3	13	27	35	21.37	-29.37	0.31	5	7.06
19	3	3	13	28	36	21.44	-29.5	-0.06	5	7.06
19	3	3	13	29	36	21.5	-29.31	0.5	5	7.06
19	3	3	13	30	36	21.5	-29.44	0.25	5	7.06
19	3	3	13	31	36	21.5	-29.25	0.63	5	7.13
19	3	3	13	32	37	21.44	-29.44	0.37	5	7.13
19	3	3	13	33	37	21.44	-29.5	0	5	7.13
19	3	3	13	34	37	21.44	-29.31	0.63	5	7.13
19	3	3	13	35	37	21.44	-29.44	0.37	5	7.13
19	3	3	13	36	38	21.44	-29.25	0.88	5	7.13
19	3	3	13	37	38	21.37	-29.37	0.56	5	7.19
19	3	3	13	38	38	21.37	-29.37	0.31	5	7.19
19	3	3	13	39	38	21.37	-29.25	0.94	5	7.19
19	3	3	13	40	39	21.37	-29.31	0.69	5	7.19
19	3	3	13	41	39	21.31	-29.12	1.19	5	7.19
19	3	3	13	42	39	21.31	-29.19	0.94	5	7.25
19	3	3	13	43	39	21.31	-29.25	0.56	5	7.25
19	3	3	13	44	40	21.31	-29.19	1.19	5	7.25
19	3	3	13	45	40	21.31	-29.25	0.94	5	7.25
19	3	3	13	46	40	21.31	-29.12	1.44	5.06	7.31
19	3	3	13	47	41	21.31	-29.19	1.19	5.06	7.31
19	3	3	13	48	41	21.37	-29.06	1.37	5.06	7.31
19	3	3	13	49	41	21.37	-29.25	1.06	5.06	7.38
19	3	3	13	50	41	21.37	-29.37	0.63	5.06	7.31
19	3	3	13	51	42	21.31	-29.25	1.19	5.06	7.31
19	3	3	13	52	42	21.37	-29.31	0.88	5.06	7.31
19	3	3	13	53	42	21.31	-29.19	1.37	5.06	7.31
19	3	3	13	54	42	21.37	-29.31	1.06	5.06	7.38
19	3	3	13	55	43	21.31	-29.37	0.75	5.06	7.31
19	3	3	13	56	43	21.31	-29.25	1.31	5.06	7.38
19	3	3	13	57	43	21.31	-29.31	1	5.06	7.38
19	3	3	13	58	43	21.31	-29.19	1.44	5.06	7.38
19	3	3	13	59	44	21.31	-29.31	1	5.06	7.38
19	3	3	14	0	44	21.31	-29.44	0.44	5.06	7.38
19	3	3	14	1	44	21.31	-29.31	1.19	5.06	7.38
19	3	3	14	2	44	21.31	-29.44	0.69	5.06	7.38
19	3	3	14	3	45	21.25	-29.31	1.25	5.06	7.38
19	3	3	14	4	45	21.25	-29.44	0.81	5.06	7.38
19	3	3	14	5	45	21.31	-29.25	1.25	5.06	7.38
19	3	3	14	6	45	21.31	-29.25	1.25	5.06	7.38
19	3	3	14	7	46	21.31	-29.37	0.81	5.06	7.38
19	3	3	14	8	46	21.37	-29.25	1.37	5.06	7.38
19	3	3	14	9	46	21.37	-29.37	0.75	5.06	7.38
19	3	3	14	10	47	21.44	-29.25	1.31	5.06	7.38
19	3	3	14	11	47	21.44	-29.44	0.75	5.06	7.38
19	3	3	14	12	47	21.5	-29.56	0.06	5.06	7.38
19	3	3	14	13	47	21.5	-29.37	0.88	5.06	7.38
19	3	3	14	14	48	21.5	-29.56	0.25	5	7.38
19	3	3	14	15	48	21.5	-29.37	0.88	5.06	7.31
19	3	3	14	16	48	21.5	-29.5	0.37	5.06	7.31
19	3	3	14	17	48	21.5	-29.31	0.94	5.06	7.31
19	3	3	14	18	49	21.5	-29.37	0.81	5	7.31
19	3	3	14	19	49	21.44	-29.5	0.19	5	7.31
19	3	3	14	20	49	21.44	-29.31	0.88	5	7.31
19	3	3	14	21	49	21.44	-29.5	0.19	5	7.31
19	3	3	14	22	50	21.44	-29.31	0.81	5	7.31
19	3	3	14	23	50	21.37	-29.44	0.19	5	7.31
19	3	3	14	24	50	21.37	-29.56	-0.56	5	7.25
19	3	3	14	25	50	21.37	-29.37	0.31	5	7.25
19	3	3	14	26	51	21.37	-29.5	-0.37	5	7.25
19	3	3	14	27	51	21.37	-29.31	0.37	5	7.25
19	3	3	14	28	51	21.37	-29.44	-0.25	5	7.25
19	3	3	14	29	52	21.37	-29.25	0.44	5	7.19
19	3	3	14	30	52	21.37	-29.44	-0.25	5	7.19
19	3	3	14	31	52	21.37	-29.56	-1	5	7.19
19	3	3	14	32	52	21.37	-29.37	-0.06	5	7.13
19	3	3	14	33	53	21.37	-29.56	-0.81	5	7.19
19	3	3	14	34	53	21.37	-29.31	-0.06	5	7.13
19	3	3	14	35	53	21.37	-29.44	-0.63	5	7.13
19	3	3	14	36	53	21.37	-29.25	0.06	5	7.13
19	3	3	14	37	54	21.44	-29.31	-0.06	5	7.13
19	3	3	14	38	54	21.44	-29.5	-1	5	7.13
19	3	3	14	39	54	21.5	-29.31	-0.12	5	7.13
19	3	3	14	40	54	21.5	-29.5	-0.94	5	7.13
19	3	3	14	41	55	21.56	-29.25	-0.19	5	7.06
19	3	3	14	42	55	21.56	-29.44	-0.94	5	7.06
19	3	3	14	43	55	21.62	-29.25	-0.56	5	7.06
19	3	3	14	44	55	21.62	-29.37	-0.75	4.94	7.06
19	3	3	14	45	56	21.62	-29.5	-1.5	5	7.06
19	3	3	14	46	56	21.62	-29.25	-0.56	4.94	7
19	3	3	14	47	56	21.62	-29.44	-1.31	4.94	7
19	3	3	14	48	57	21.62	-29.19	-0.5	4.94	7
19	3	3	14	49	57	21.62	-29.37	-1	4.94	7
19	3	3	14	50	57	21.62	-29.31	-1.5	4.94	7
19	3	3	14	51	57	21.62	-29.25	-0.81	4.94	6.94
19	3	3	14	52	58	21.62	-29.44	-1.5	4.94	6.94
19	3	3	14	53	58	21.56	-29.19	-0.63	4.94	7
19	3	3	14	54	58	21.56	-29.37	-1.19	4.94	6.94
19	3	3	14	55	58	21.56	-29.12	-0.5	4.94	6.94
19	3	3	14	56	59	21.56	-29.25	-1	4.94	6.94
19	3	3	14	57	59	21.5	-29.44	-1.88	4.94	6.94
19	3	3	14	58	59	21.5	-29.19	-0.81	4.94	6.94
19	3	3	14	59	59	21.5	-29.37	-1.31	4.94	6.94
19	3	3	15	1	0	21.5	-29.19	-0.56	4.94	6.94
19	3	3	15	2	0	21.44	-29.31	-1	4.94	6.94
19	3	3	15	3	0	21.44	-29.06	-0.31	4.94	6.94
19	3	3	15	4	0	21.44	-29.25	-0.75	4.94	6.94
19	3	3	15	5	1	21.44	-29.37	-1.44	4.94	6.94
19	3	3	15	6	1	21.44	-29.12	-0.5	4.94	6.94
19	3	3	15	7	1	21.44	-29.31	-1.06	4.94	6.94
19	3	3	15	8	1	21.44	-29.06	-0.31	4.94	6.94
19	3	3	15	9	2	21.44	-29.31	-1.56	4.94	6.94
19	3	3	15	10	2	21.5	-29.06	-0.75	4.94	6.94
19	3	3	15	11	2	21.5	-29.12	-0.81	4.94	6.94
19	3	3	15	12	3	21.56	-29.19	-1.31	4.94	6.94
19	3	3	15	13	3	21.56	-29	-0.44	4.94	6.94
19	3	3	15	14	3	21.5	-29.12	-0.88	4.94	6.94
19	3	3	15	15	3	21.5	-29	-0.19	4.94	6.94
19	3	3	15	16	4	21.5	-29.06	-0.56	4.94	7
19	3	3	15	17	4	21.5	-29.19	-1.19	4.94	7
19	3	3	15	18	4	21.5	-29.06	-0.31	4.94	7
19	3	3	15	19	4	21.5	-29.19	-0.63	4.94	7
19	3	3	15	20	5	21.44	-29	0	4.94	7
19	3	3	15	21	5	21.44	-29.06	-0.25	4.94	7.06
19	3	3	15	22	5	21.44	-28.94	0.25	4.94	7
19	3	3	15	23	5	21.44	-29.06	0	4.94	7.06
19	3	3	15	24	6	21.37	-29.19	-0.5	4.94	7.06
19	3	3	15	25	6	21.37	-29	0.25	5	7.06
19	3	3	15	26	6	21.37	-29.12	-0.12	5	7.06
19	3	3	15	27	6	21.37	-28.94	0.44	5	7.06
19	3	3	15	28	7	21.31	-29.06	0.06	5	7.06
19	3	3	15	29	7	21.31	-29.12	-0.31	5	7.06
19	3	3	15	30	7	21.31	-29	0.44	5	7.13
19	3	3	15	31	8	21.31	-29	0.12	5	7.13
19	3	3	15	32	8	21.31	-28.87	0.69	5	7.13
19	3	3	15	33	8	21.31	-28.94	0.44	5	7.13
19	3	3	15	34	8	21.37	-28.75	0.88	5	7.19
19	3	3	15	35	9	21.37	-28.87	0.56	5	7.19
19	3	3	15	36	9	21.37	-29	0.25	5	7.19
19	3	3	15	37	9	21.44	-28.87	0.81	5	7.19
19	3	3	15	38	9	21.44	-28.94	0.5	5	7.19
19	3	3	15	39	10	21.5	-28.81	1	5	7.25
19	3	3	15	40	10	21.5	-28.87	0.88	5	7.25
19	3	3	15	41	10	21.5	-28.94	0.56	5	7.25
19	3	3	15	42	10	21.5	-28.81	1.12	5	7.25
19	3	3	15	43	11	21.5	-28.94	0.94	5	7.31
19	3	3	15	44	11	21.5	-28.81	1.37	5.06	7.31
19	3	3	15	45	11	21.5	-28.87	1.19	5	7.31
19	3	3	15	46	11	21.5	-28.94	0.81	5.06	7.38
19	3	3	15	47	12	21.5	-28.81	1.44	5.06	7.38
19	3	3	15	48	12	21.5	-28.87	1.19	5.06	7.38
19	3	3	15	49	12	21.5	-28.75	1.63	5.06	7.44
19	3	3	15	50	13	21.5	-28.87	1.31	5.06	7.44
19	3	3	15	51	13	21.5	-28.75	1.44	5.06	7.44
19	3	3	15	52	13	21.5	-28.94	1.19	5.06	7.44
19	3	3	15	53	13	21.5	-29.06	0.81	5.06	7.44
19	3	3	15	54	14	21.5	-28.87	1.37	5.06	7.44
19	3	3	15	55	14	21.5	-29	1.06	5.06	7.44
19	3	3	15	56	14	21.5	-28.87	1.5	5.06	7.44
19	3	3	15	57	14	21.44	-28.94	1.25	5.06	7.44
19	3	3	15	58	15	21.44	-29	0.88	5.06	7.44
19	3	3	15	59	15	21.44	-28.87	1.44	5.06	7.44
19	3	3	16	0	15	21.44	-29	1	5.06	7.44
19	3	3	16	1	15	21.44	-28.87	1.5	5.06	7.44
19	3	3	16	2	16	21.5	-29	1.06	5.06	7.44
19	3	3	16	3	16	21.5	-29.06	0.56	5.06	7.44
19	3	3	16	4	16	21.5	-28.94	1.25	5.06	7.44
19	3	3	16	5	16	21.56	-29.06	0.81	5.06	7.44
19	3	3	16	6	17	21.56	-28.87	1.31	5.06	7.44
19	3	3	16	7	17	21.56	-29.06	0.88	5.06	7.44
19	3	3	16	8	17	21.56	-28.87	1.37	5.06	7.38
19	3	3	16	9	17	21.56	-29.12	0.63	5.06	7.38
19	3	3	16	10	18	21.56	-29.25	0.12	5.06	7.38
19	3	3	16	11	18	21.56	-29.06	0.88	5.06	7.38
19	3	3	16	12	18	21.5	-29.19	0.31	5.06	7.38
19	3	3	16	13	19	21.5	-29	0.94	5.06	7.31
19	3	3	16	14	19	21.5	-29.12	0.44	5.06	7.38
19	3	3	16	15	19	21.5	-29.12	-0.06	5.06	7.31
19	3	3	16	16	19	21.44	-29.06	0.63	5.06	7.31
19	3	3	16	17	20	21.44	-29.19	0.06	5.06	7.31
19	3	3	16	18	23	21.44	-29	0.75	5.06	7.31
19	3	3	16	19	24	21.44	-29.19	0.12	5.06	7.31
19	3	3	16	20	24	21.37	-29	0.75	5.06	7.31
19	3	3	16	21	24	21.37	-29.06	0.44	5.06	7.25
19	3	3	16	22	24	21.31	-29.25	-0.19	5.06	7.31
19	3	3	16	23	25	21.31	-29	0.56	5.06	7.25
19	3	3	16	24	25	21.31	-29.19	-0.06	5.06	7.25
19	3	3	16	25	25	21.31	-29	0.56	5.06	7.25
19	3	3	16	26	25	21.31	-29.19	0	5.06	7.25
19	3	3	16	27	26	21.31	-29.12	-0.44	5.06	7.25
19	3	3	16	28	26	21.37	-29.06	0.25	5.06	7.25
19	3	3	16	29	26	21.37	-29.19	-0.56	5	7.25
19	3	3	16	30	26	21.37	-29	0.19	5.06	7.25
19	3	3	16	31	27	21.37	-29.12	-0.5	5	7.19
19	3	3	16	32	27	21.37	-28.94	0.25	5	7.19
19	3	3	16	33	27	21.37	-29.12	-0.37	5	7.19
19	3	3	16	34	28	21.37	-29.25	-1.12	5	7.19
19	3	3	16	35	28	21.37	-29.06	-0.12	5	7.13
19	3	3	16	36	28	21.37	-29.19	-0.88	5	7.13
19	3	3	16	37	28	21.37	-28.94	-0.06	5	7.13
19	3	3	16	38	29	21.37	-29.12	-0.69	5	7.13
19	3	3	16	39	29	21.37	-28.87	0	5	7.13
19	3	3	16	40	29	21.31	-29.06	-0.69	5	7.06
19	3	3	16	41	29	21.31	-29.19	-1.5	5	7.06
19	3	3	16	42	30	21.31	-29	-0.5	5	7.06
19	3	3	16	43	30	21.31	-29.12	-1.31	5	7.06
19	3	3	16	44	30	21.31	-28.94	-0.5	5	7.06
19	3	3	16	45	30	21.25	-29.06	-1.12	5	7
19	3	3	16	46	31	21.25	-28.81	-0.37	5	7
19	3	3	16	47	31	21.25	-28.94	-0.69	5	7
19	3	3	16	48	31	21.31	-29.19	-2.31	5	7
19	3	3	16	49	31	21.31	-28.94	-1.06	4.94	7
19	3	3	16	50	32	21.31	-29.12	-1.94	4.94	7
19	3	3	16	51	32	21.37	-28.87	-1	4.94	6.94
19	3	3	16	52	32	21.44	-29	-1.63	4.94	6.94
19	3	3	16	53	32	21.44	-28.75	-0.88	4.94	6.94
19	3	3	16	54	33	21.5	-28.94	-1.5	4.94	6.94
19	3	3	16	55	33	21.5	-29.12	-2.38	4.94	6.94
19	3	3	16	56	33	21.5	-28.87	-1.25	4.94	6.88
19	3	3	16	57	34	21.5	-29.06	-2.06	4.94	6.88
19	3	3	16	58	34	21.5	-28.81	-1.12	4.94	6.88
19	3	3	16	59	34	21.56	-29	-1.69	4.94	6.88
19	3	3	17	0	34	21.56	-28.75	-0.94	4.94	6.88
19	3	3	17	1	35	21.5	-28.87	-1.44	4.94	6.88
19	3	3	17	2	35	21.5	-29	-2.31	4.94	6.88
19	3	3	17	3	35	21.5	-28.75	-1.25	4.94	6.81
19	3	3	17	4	35	21.5	-28.94	-1.88	4.94	6.81
19	3	3	17	5	36	21.5	-28.69	-1	4.94	6.81
19	3	3	17	6	36	21.5	-28.87	-1.56	4.94	6.81
19	3	3	17	7	36	21.5	-28.62	-0.88	4.94	6.81
19	3	3	17	8	36	21.5	-28.75	-1.06	4.94	6.81
19	3	3	17	9	37	21.44	-28.94	-1.75	4.94	6.88
19	3	3	17	10	37	21.44	-28.69	-0.81	4.94	6.88
19	3	3	17	11	37	21.44	-28.81	-1.44	4.94	6.88
19	3	3	17	12	37	21.44	-28.62	-0.69	4.94	6.88
19	3	3	17	13	38	21.44	-28.81	-1.19	4.94	6.88
19	3	3	17	14	38	21.37	-28.56	-0.88	4.94	6.88
19	3	3	17	15	38	21.37	-28.75	-0.94	4.94	6.88
19	3	3	17	16	39	21.37	-28.87	-1.56	4.94	6.88
19	3	3	17	17	39	21.37	-28.69	-0.75	4.94	6.88
19	3	3	17	18	39	21.37	-28.81	-1.31	4.94	6.88
19	3	3	17	19	39	21.31	-28.62	-0.63	4.94	6.88
19	3	3	17	20	40	21.37	-28.75	-1	4.94	6.88
19	3	3	17	21	40	21.31	-28.81	-1.63	4.94	6.94
19	3	3	17	22	40	21.37	-28.69	-0.69	4.94	6.94
19	3	3	17	23	40	21.37	-28.81	-1.25	4.94	6.94
19	3	3	17	24	41	21.37	-28.62	-0.5	4.94	6.94
19	3	3	17	25	41	21.44	-28.75	-0.94	4.94	6.94
19	3	3	17	26	41	21.44	-28.56	-0.31	5	6.94
19	3	3	17	27	41	21.44	-28.69	-0.69	4.94	6.94
19	3	3	17	28	42	21.44	-28.81	-1.25	5	7
19	3	3	17	29	42	21.44	-28.62	-0.37	4.94	7
19	3	3	17	30	42	21.44	-28.75	-0.81	4.94	7
19	3	3	17	31	42	21.44	-28.56	-0.12	4.94	7
19	3	3	17	32	43	21.37	-28.69	-0.5	5	7
19	3	3	17	33	43	21.37	-28.5	-0.25	5	7
19	3	3	17	34	43	21.37	-28.62	-0.37	5	7
19	3	3	17	35	43	21.37	-28.75	-0.94	5	7
19	3	3	17	36	44	21.37	-28.56	-0.19	5	7
19	3	3	17	37	44	21.37	-28.69	-0.56	5	7.06
19	3	3	17	38	44	21.31	-28.5	0.06	5	7
19	3	3	17	39	45	21.31	-28.62	-0.19	5	7.06
19	3	3	17	40	45	21.31	-28.75	-0.69	5	7.06
19	3	3	17	41	45	21.31	-28.56	0.12	5	7.06
19	3	3	17	42	45	21.31	-28.62	-0.25	5	7.06
19	3	3	17	43	46	21.25	-28.5	0.37	5	7.13
19	3	3	17	44	46	21.25	-28.56	0.06	5	7.13
19	3	3	17	45	46	21.44	-28.37	0.44	5	7.13
19	3	3	17	46	46	21.44	-28.5	0.19	5	7.13
19	3	3	17	47	47	21.44	-28.62	-0.25	5	7.13
19	3	3	17	48	47	21.44	-28.5	0.44	5	7.13
19	3	3	17	49	47	21.37	-28.56	0.19	5	7.19
19	3	3	17	50	47	21.37	-28.44	0.69	5	7.13
19	3	3	17	51	48	21.37	-28.5	0.5	5	7.19
19	3	3	17	52	48	21.37	-28.56	0.12	5	7.19
19	3	3	17	53	48	21.37	-28.5	0.75	5	7.19
19	3	3	17	54	48	21.37	-28.56	0.5	5.06	7.25
19	3	3	17	55	49	21.37	-28.44	1	5	7.25
19	3	3	17	56	49	21.44	-28.5	0.75	5.06	7.25
19	3	3	17	57	49	21.44	-28.5	0.31	5	7.25
19	3	3	17	58	50	21.44	-28.44	0.94	5.06	7.25
19	3	3	17	59	50	21.44	-28.5	0.75	5.06	7.31
19	3	3	18	0	50	21.37	-28.37	1.25	5.06	7.31
19	3	3	18	1	50	21.37	-28.44	1.06	5.06	7.31
19	3	3	18	2	51	21.37	-28.25	1.5	5.06	7.38
19	3	3	18	3	51	21.37	-28.12	1.75	5.06	7.38
19	3	3	18	4	51	21.31	-28.19	1.37	5.06	7.44
19	3	3	18	5	51	21.31	-28.12	1.81	5.06	7.44
19	3	3	18	6	52	21.31	-28.25	1.56	5.06	7.44
19	3	3	18	7	52	21.25	-28.19	1.94	5.06	7.44
19	3	3	18	8	52	21.25	-28.25	1.69	5.06	7.5
19	3	3	18	9	52	21.25	-28.37	1.31	5.13	7.5
19	3	3	18	10	53	21.19	-28.25	1.88	5.13	7.5
19	3	3	18	11	53	21.19	-28.44	1.5	5.13	7.56
19	3	3	18	12	53	21.19	-28.31	1.88	5.13	7.5
19	3	3	18	13	53	21.12	-28.44	1.56	5.13	7.5
19	3	3	18	14	54	21.12	-28.5	1.19	5.06	7.56
19	3	3	18	15	54	21.12	-28.44	1.75	5.13	7.56
19	3	3	18	16	54	21.06	-28.5	1.37	5.06	7.56
19	3	3	18	17	54	21.06	-28.37	1.81	5.13	7.56
19	3	3	18	18	55	21	-28.5	1.5	5.13	7.56
19	3	3	18	19	55	21	-28.31	1.88	5.13	7.56
19	3	3	18	20	55	21	-28.5	1.56	5.13	7.56
19	3	3	18	21	56	21	-28.56	1.12	5.13	7.56
19	3	3	18	22	56	21	-28.44	1.69	5.13	7.56
19	3	3	18	23	56	21.06	-28.56	1.25	5.13	7.56
19	3	3	18	24	56	21.06	-28.44	1.69	5.13	7.56
19	3	3	18	25	57	21.06	-28.56	1.31	5.13	7.5
19	3	3	18	26	57	21.06	-28.62	0.81	5.13	7.5
19	3	3	18	27	57	21.06	-28.5	1.5	5.13	7.5
19	3	3	18	28	57	21.06	-28.62	0.88	5.06	7.5
19	3	3	18	29	58	21.06	-28.5	1.44	5.13	7.5
19	3	3	18	30	58	21.06	-28.62	0.88	5.13	7.5
19	3	3	18	31	58	21.06	-28.5	1.06	5.13	7.5
19	3	3	18	32	58	21	-28.69	0.81	5.13	7.44
19	3	3	18	33	59	21	-28.81	0.31	5.06	7.44
19	3	3	18	34	59	21	-28.62	1	5.06	7.44
19	3	3	18	35	59	21	-28.75	0.5	5.06	7.44
19	3	3	18	36	59	20.94	-28.56	1.06	5.06	7.44
19	3	3	18	38	0	20.94	-28.69	0.63	5.06	7.38
19	3	3	18	39	0	20.87	-28.81	-0.12	5.06	7.38
19	3	3	18	40	0	20.87	-28.62	0.75	5.06	7.38
19	3	3	18	41	1	20.87	-28.81	0.06	5.06	7.38
19	3	3	18	42	1	20.87	-28.56	0.75	5.06	7.38
19	3	3	18	43	1	20.81	-28.69	0.19	5.06	7.38
19	3	3	18	44	1	20.81	-28.5	0.81	5.06	7.31
19	3	3	18	45	2	20.81	-28.56	0.56	5.06	7.31
19	3	3	18	46	2	20.81	-28.75	-0.06	5.06	7.31
19	3	3	18	47	2	20.81	-28.5	0.69	5.06	7.31
19	3	3	18	48	2	20.81	-28.69	0.06	5.06	7.31
19	3	3	18	49	3	20.81	-28.5	0.69	5.06	7.31
19	3	3	18	50	3	20.81	-28.62	0.06	5.06	7.31
19	3	3	18	51	3	20.81	-28.62	-0.37	5.06	7.31
19	3	3	18	52	3	20.87	-28.62	0.19	5.06	7.25
19	3	3	18	53	4	20.87	-28.81	-0.63	5.06	7.25
19	3	3	18	54	4	20.81	-28.56	0.19	5.06	7.25
19	3	3	18	55	4	20.81	-28.75	-0.5	5.06	7.25
19	3	3	18	56	4	20.81	-28.5	0.19	5.06	7.19
19	3	3	18	57	5	20.81	-28.69	-0.37	5.06	7.19
19	3	3	18	58	5	20.81	-28.75	-1.19	5.06	7.19
19	3	3	18	59	5	20.81	-28.56	-0.12	5	7.19
19	3	3	19	0	6	20.75	-28.75	-1	5.06	7.19
19	3	3	19	1	6	20.75	-28.5	-0.19	5.06	7.19
19	3	3	19	2	6	20.75	-28.69	-0.88	5.06	7.13
19	3	3	19	3	6	20.69	-28.5	-0.12	5.06	7.13
19	3	3	19	4	7	20.69	-28.62	-0.75	5	7.13
19	3	3	19	5	7	20.69	-28.75	-1.63	5.06	7.06
19	3	3	19	6	7	20.69	-28.5	-0.56	5	7.06
19	3	3	19	7	7	20.69	-28.69	-1.31	5	7.06
19	3	3	19	8	8	20.69	-28.44	-0.44	5	7.06
19	3	3	19	9	8	20.69	-28.62	-1.25	5	7.06
19	3	3	19	10	8	20.69	-28.37	-0.5	5	7
19	3	3	19	11	8	20.75	-28.56	-1.19	5	7
19	3	3	19	12	9	20.75	-28.69	-2.06	5	7
19	3	3	19	13	9	20.75	-28.44	-0.94	5	7
19	3	3	19	14	9	20.81	-28.62	-1.75	5	7
19	3	3	19	15	9	20.81	-28.37	-0.88	5	6.94
19	3	3	19	16	10	20.81	-28.5	-1.5	5	6.94
19	3	3	19	17	10	20.87	-28.25	-0.75	5	6.94
19	3	3	19	18	10	20.87	-28.44	-1.31	5	6.94
19	3	3	19	19	10	20.87	-28.56	-2.38	5	6.94
19	3	3	19	20	11	20.87	-28.31	-1.19	5	6.94
19	3	3	19	21	11	20.87	-28.5	-2	5	6.94
19	3	3	19	22	11	20.81	-28.25	-1.06	5	6.88
19	3	3	19	23	12	20.81	-28.44	-1.88	4.94	6.88
19	3	3	19	24	12	20.81	-28.19	-1	5	6.88
19	3	3	19	25	12	20.81	-28.37	-1.69	5	6.88
19	3	3	19	26	12	20.81	-28.5	-2.75	4.94	6.88
19	3	3	19	27	13	20.81	-28.25	-1.5	4.94	6.88
19	3	3	19	28	13	20.81	-28.5	-2.5	4.94	6.88
19	3	3	19	29	13	20.81	-28.19	-1.5	4.94	6.81
19	3	3	19	30	13	20.75	-28.37	-2.19	4.94	6.81
19	3	3	19	31	14	20.75	-28.12	-1.31	4.94	6.81
19	3	3	19	32	14	20.75	-28.31	-1.94	4.94	6.81
19	3	3	19	33	14	20.75	-28.44	-3	4.94	6.81
19	3	3	19	34	14	20.75	-28.19	-1.69	4.94	6.81
19	3	3	19	35	15	20.75	-28.44	-2.5	4.94	6.81
19	3	3	19	36	15	20.81	-28.12	-1.5	4.94	6.81
19	3	3	19	37	15	20.81	-28.31	-2.19	4.94	6.75
19	3	3	19	38	15	20.81	-28.06	-1.37	4.94	6.75
19	3	3	19	39	16	20.81	-28.19	-1.75	4.94	6.75
19	3	3	19	40	16	20.81	-28.37	-2.75	4.94	6.75
19	3	3	19	41	16	20.87	-28.12	-1.56	4.94	6.75
19	3	3	19	42	16	20.87	-28.31	-2.25	4.94	6.75
19	3	3	19	43	17	20.94	-28.06	-1.31	4.94	6.75
19	3	3	19	44	17	20.94	-28.25	-1.94	4.94	6.75
19	3	3	19	45	17	20.94	-28	-1.12	4.94	6.75
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	3	3	19	53	5	20.87	-26.87	1.12	5	85
19	3	3	19	54	5	20.87	-26.81	1.37	5	85
19	3	3	19	55	5	20.81	-26.69	1.5	5	85
19	3	3	19	56	5	20.81	-26.69	1.69	5	85
19	3	3	19	57	6	20.81	-26.62	1.81	5.06	85
19	3	3	19	58	6	20.81	-26.56	2	5.06	85
19	3	3	19	59	6	20.75	-26.56	2.13	5.06	85
19	3	3	20	0	7	20.75	-26.56	2.25	5.06	85
19	3	3	20	1	7	20.75	-26.56	2.38	5.06	85
19	3	3	20	2	7	20.75	-26.56	2.44	5.13	85
19	3	3	20	3	7	20.75	-26.62	2.63	5.13	85
19	3	3	20	4	8	20.75	-26.62	2.69	5.13	85
19	3	3	20	5	8	20.69	-26.62	2.75	5.13	7.56
19	3	3	20	6	8	20.69	-26.69	2.88	5.13	7.63
19	3	3	20	7	8	20.75	-26.62	2.94	5.13	7.63
19	3	3	20	8	9	20.75	-26.69	3.06	5.13	7.69
19	3	3	20	9	9	20.75	-26.62	3.13	5.13	7.69
19	3	3	20	10	9	20.75	-26.56	3.25	5.19	7.75
19	3	3	20	11	9	20.81	-26.5	3.38	5.19	7.75
19	3	3	20	12	10	20.81	-26.56	3.38	5.19	7.81
19	3	3	20	13	10	20.87	-26.56	3.5	5.19	7.81
19	3	3	20	14	10	20.87	-26.56	3.56	5.19	7.81
19	3	3	20	15	10	20.87	-26.56	3.63	5.19	7.88
19	3	3	20	16	11	20.94	-26.5	3.75	5.19	7.88
19	3	3	20	17	11	20.94	-26.5	3.75	5.19	7.88
19	3	3	20	18	11	20.94	-26.56	3.81	5.19	7.94
19	3	3	20	19	11	20.94	-26.56	3.88	5.19	7.94
19	3	3	20	20	12	20.87	-26.56	4	5.25	8
19	3	3	20	21	12	20.94	-26.5	4.06	5.25	8
19	3	3	20	22	12	20.87	-26.56	4.06	5.25	8
19	3	3	20	23	13	20.87	-26.56	4.19	5.25	8
19	3	3	20	24	13	20.87	-26.62	4.19	5.25	8.06
19	3	3	20	25	13	20.87	-26.69	4.19	5.25	8.06
19	3	3	20	26	13	20.87	-26.69	4.31	5.25	8.06
19	3	3	20	27	14	20.81	-26.75	4.31	5.25	8.06
19	3	3	20	28	14	20.81	-26.75	4.44	5.25	8.13
19	3	3	20	29	14	20.81	-26.75	4.44	5.25	8.13
19	3	3	20	30	14	20.81	-26.75	4.44	5.25	8.13
19	3	3	20	31	15	20.75	-26.81	4.56	5.25	8.13
19	3	3	20	32	15	20.75	-26.81	4.56	5.25	8.19
19	3	3	20	33	15	20.75	-26.75	4.63	5.31	8.19
19	3	3	20	34	15	20.75	-26.75	4.69	5.31	8.19
19	3	3	20	35	16	20.69	-26.75	4.63	5.31	8.19
19	3	3	20	36	18	20.69	-26.81	4.75	5.31	8.19
19	3	3	20	37	18	20.69	-26.87	4.63	5.31	8.19
19	3	3	20	38	18	20.62	-26.81	4.69	5.31	8.25
19	3	3	20	39	19	20.69	-26.87	4.56	5.31	8.25
19	3	3	20	40	19	20.62	-27	4.44	5.31	8.25
19	3	3	20	41	19	20.69	-27	4.63	5.31	8.25
19	3	3	20	42	20	20.62	-27.12	4.5	5.31	8.25
19	3	3	20	43	20	20.69	-27.12	4.63	5.31	8.25
19	3	3	20	44	20	20.69	-27.19	4.56	5.31	8.25
19	3	3	20	45	20	20.69	-27.19	4.38	5.31	8.25
19	3	3	20	46	21	20.75	-27.19	4.56	5.31	8.25
19	3	3	20	47	21	20.75	-27.25	4.31	5.31	8.25
19	3	3	20	48	21	20.75	-27.25	4.56	5.31	8.25
19	3	3	20	49	21	20.81	-27.37	4.31	5.31	8.25
19	3	3	20	50	22	20.81	-27.37	3.94	5.31	8.25
19	3	3	20	51	22	20.81	-27.44	4.31	5.31	8.19
19	3	3	20	52	22	20.75	-27.5	4	5.31	8.19
19	3	3	20	53	22	20.75	-27.44	4.31	5.31	8.19
19	3	3	20	54	23	20.75	-27.56	4.06	5.31	8.19
19	3	3	20	55	23	20.75	-27.62	3.69	5.31	8.19
19	3	3	20	56	23	20.75	-27.62	4.13	5.31	8.19
19	3	3	20	57	23	20.75	-27.75	3.69	5.31	8.19
19	3	3	20	58	24	20.69	-27.69	4	5.38	8.19
19	3	3	20	59	24	20.69	-27.69	3.69	5.31	8.13
19	3	3	21	0	24	20.69	-27.69	3.25	5.31	8.13
19	3	3	21	1	24	20.62	-27.75	3.63	5.31	8.13
19	3	3	21	2	25	20.62	-27.81	3.25	5.31	8.13
19	3	3	21	3	25	20.62	-27.69	3.63	5.31	8.13
19	3	3	21	4	25	20.56	-27.81	3.31	5.31	8.06
19	3	3	21	5	26	20.56	-27.69	3.69	5.31	8.06
19	3	3	21	6	26	20.56	-27.62	3.69	5.31	8.06
19	3	3	21	7	26	20.56	-27.75	3.06	5.31	8.13
19	3	3	21	8	26	20.5	-27.69	3.5	5.31	8.06
19	3	3	21	9	27	20.5	-27.81	3.06	5.31	8.06
19	3	3	21	10	27	20.5	-27.69	3.38	5.31	8.06
19	3	3	21	11	27	20.5	-27.81	3	5.31	8.06
19	3	3	21	12	27	20.56	-27.94	2.44	5.31	8
19	3	3	21	13	28	20.56	-27.87	3	5.31	8
19	3	3	21	14	28	20.56	-27.94	2.56	5.31	8
19	3	3	21	15	28	20.62	-27.81	3	5.31	8
19	3	3	21	16	28	20.62	-27.87	2.63	5.31	7.94
19	3	3	21	17	29	20.62	-27.94	2.06	5.31	7.94
19	3	3	21	18	29	20.62	-27.87	2.56	5.31	7.94
19	3	3	21	19	29	20.62	-28	2	5.31	7.94
19	3	3	21	20	29	20.62	-27.87	2.56	5.31	7.88
19	3	3	21	21	30	20.56	-28	2.06	5.31	7.88
19	3	3	21	22	30	20.56	-27.81	2.56	5.31	7.88
19	3	3	21	23	30	20.56	-27.94	2.19	5.31	7.88
19	3	3	21	24	30	20.56	-28.06	1.56	5.31	7.81
19	3	3	21	25	31	20.5	-27.87	2.19	5.25	7.81
19	3	3	21	26	31	20.5	-27.94	1.69	5.25	7.81
19	3	3	21	27	31	20.5	-27.81	2.25	5.25	7.75
19	3	3	21	28	32	20.5	-27.94	1.69	5.25	7.75
19	3	3	21	29	32	20.5	-28	0.94	5.25	7.75
19	3	3	21	30	32	20.44	-27.94	1.69	5.25	7.75
19	3	3	21	31	32	20.44	-28.06	1.06	5.25	7.75
19	3	3	21	32	33	20.44	-27.87	1.69	5.25	7.69
19	3	3	21	33	33	20.44	-28	1.19	5.25	7.69
19	3	3	21	34	33	20.44	-27.81	1.75	5.25	7.69
19	3	3	21	35	33	20.44	-27.94	1.31	5.25	7.63
19	3	3	21	36	34	20.5	-28.12	0.44	5.25	7.63
19	3	3	21	37	34	20.5	-27.87	1.25	5.25	7.63
19	3	3	21	38	34	20.5	-28.06	0.63	5.25	7.63
19	3	3	21	39	34	20.5	-27.87	1.25	5.25	7.56
19	3	3	21	40	35	20.56	-28	0.69	5.25	7.56
19	3	3	21	41	35	20.56	-27.81	1.06	5.25	7.56
19	3	3	21	42	35	20.56	-27.94	0.81	5.25	7.56
19	3	3	21	43	35	20.62	-28.06	0.12	5.25	7.5
19	3	3	21	44	36	20.62	-27.87	0.88	5.25	7.5
19	3	3	21	45	36	20.62	-27.94	0.31	5.25	7.5
19	3	3	21	46	36	20.62	-27.75	0.94	5.25	7.44
19	3	3	21	47	37	20.62	-27.94	0.37	5.19	7.44
19	3	3	21	48	37	20.62	-28.06	-0.56	5.19	7.44
19	3	3	21	49	37	20.62	-27.87	0.37	5.19	7.44
19	3	3	21	50	37	20.56	-28	-0.31	5.19	7.44
19	3	3	21	51	38	20.56	-27.81	0.44	5.19	7.38
19	3	3	21	52	38	20.56	-27.94	-0.19	5.19	7.38
19	3	3	21	53	38	20.5	-27.75	0.5	5.19	7.38
19	3	3	21	54	38	20.5	-27.87	0	5.19	7.38
19	3	3	21	55	39	20.5	-28.06	-0.94	5.19	7.38
19	3	3	21	56	39	20.5	-27.81	0.12	5.19	7.31
19	3	3	21	57	39	20.44	-28	-0.69	5.19	7.31
19	3	3	21	58	39	20.44	-27.75	0.12	5.19	7.31
19	3	3	21	59	40	20.44	-27.94	-0.5	5.19	7.31
19	3	3	22	0	40	20.37	-27.69	0.19	5.19	7.31
19	3	3	22	1	40	20.37	-27.81	-0.37	5.19	7.25
19	3	3	22	2	40	20.37	-28	-1.37	5.19	7.25
19	3	3	22	3	41	20.37	-27.75	-0.31	5.19	7.25
19	3	3	22	4	41	20.37	-27.87	-1.06	5.19	7.25
19	3	3	22	5	41	20.37	-27.69	-0.19	5.19	7.25
19	3	3	22	6	41	20.44	-27.81	-0.81	5.19	7.19
19	3	3	22	7	42	20.44	-27.56	-0.12	5.19	7.19
19	3	3	22	8	42	20.44	-27.75	-0.69	5.13	7.19
19	3	3	22	9	42	20.44	-27.94	-1.75	5.19	7.19
19	3	3	22	10	43	20.5	-27.69	-0.69	5.13	7.19
19	3	3	22	11	43	20.5	-27.87	-1.5	5.13	7.13
19	3	3	22	12	43	20.5	-27.62	-0.63	5.13	7.13
19	3	3	22	13	43	20.5	-27.81	-1.31	5.13	7.13
19	3	3	22	14	44	20.5	-27.56	-0.5	5.13	7.13
19	3	3	22	15	44	20.5	-27.69	-1.12	5.13	7.13
19	3	3	22	16	44	20.5	-27.94	-2.31	5.13	7.13
19	3	3	22	17	44	20.44	-27.62	-1.12	5.13	7.06
19	3	3	22	18	45	20.44	-27.75	-2	5.13	7.06
19	3	3	22	19	45	20.44	-27.5	-1.06	5.13	7.06
19	3	3	22	20	45	20.37	-27.69	-1.81	5.13	7
19	3	3	22	21	45	20.37	-27.44	-0.94	5.13	7
19	3	3	22	22	46	20.37	-27.62	-1.63	5.13	7
19	3	3	22	23	46	20.31	-27.81	-2.69	5.13	7
19	3	3	22	24	46	20.31	-27.56	-1.44	5.13	6.94
19	3	3	22	25	46	20.31	-27.75	-2.31	5.06	6.94
19	3	3	22	26	47	20.25	-27.44	-1.31	5.13	6.94
19	3	3	22	27	47	20.25	-27.62	-2.13	5.06	6.94
19	3	3	22	28	47	20.19	-27.37	-1.25	5.06	6.94
19	3	3	22	29	47	20.19	-27.56	-2	5.06	6.88
19	3	3	22	30	48	20.19	-27.62	-2.75	5.06	6.88
19	3	3	22	31	48	20.19	-27.44	-1.81	5.06	6.88
19	3	3	22	32	48	20.19	-27.62	-2.75	5.06	6.88
19	3	3	22	33	49	20.25	-27.37	-1.69	5.06	6.88
19	3	3	22	34	49	20.25	-27.56	-2.44	5.06	6.88
19	3	3	22	35	49	20.25	-27.31	-1.5	5.06	6.88
19	3	3	22	36	49	20.25	-27.5	-2.25	5.06	6.81
19	3	3	22	37	50	20.25	-27.25	-1.69	5.06	6.81
19	3	3	22	38	50	20.31	-27.37	-2.06	5.06	6.81
19	3	3	22	39	50	20.31	-27.62	-3.13	5.06	6.81
19	3	3	22	40	50	20.37	-27.37	-1.94	5.06	6.81
19	3	3	22	41	51	20.37	-27.5	-2.81	5.06	6.81
19	3	3	22	42	51	20.37	-27.31	-1.81	5.06	6.81
19	3	3	22	43	51	20.37	-27.44	-2.56	5	6.81
19	3	3	22	44	51	20.37	-27.19	-1.69	5	6.75
19	3	3	22	45	52	20.44	-27.37	-2.38	5	6.75
19	3	3	22	46	52	20.44	-27.56	-3.63	5	6.75
19	3	3	22	47	52	20.44	-27.31	-2.25	5	6.75
19	3	3	22	48	52	20.5	-27.5	-3.38	5	6.75
19	3	3	22	49	53	20.5	-27.19	-2.19	5	6.75
19	3	3	22	50	53	20.5	-27.44	-3.06	5	6.75
19	3	3	22	51	53	20.56	-27.12	-2.06	5	6.75
19	3	3	22	52	53	20.56	-27.31	-2.88	5	6.69
19	3	3	22	53	54	20.56	-27.06	-2	5	6.69
19	3	3	22	54	54	20.56	-27.25	-2.69	5	6.69
19	3	3	22	55	54	20.56	-27.44	-3.81	5	6.69
19	3	3	22	56	55	20.56	-27.12	-2.5	5	6.69
19	3	3	22	57	55	20.5	-27.37	-3.56	5	6.69
19	3	3	22	58	55	20.5	-27.06	-2.44	5	6.69
19	3	3	22	59	55	20.5	-27.31	-3.38	4.94	6.63
19	3	3	23	0	56	20.44	-27	-2.31	4.94	6.63
19	3	3	23	1	56	20.44	-27.19	-3.19	4.94	6.63
19	3	3	23	2	56	20.37	-27.06	-2.94	4.94	6.63
19	3	3	23	3	56	20.37	-27.12	-2.94	4.94	6.63
19	3	3	23	4	57	20.31	-27.37	-4.19	4.94	6.63
19	3	3	23	5	57	20.31	-27.06	-2.81	4.94	6.56
19	3	3	23	6	57	20.31	-27.25	-3.88	4.94	6.56
19	3	3	23	7	57	20.25	-26.94	-2.69	4.94	6.56
19	3	3	23	8	58	20.25	-27.19	-3.56	4.94	6.56
19	3	3	23	9	58	20.25	-26.94	-2.56	4.94	6.56
19	3	3	23	10	58	20.25	-27.12	-3.38	4.88	6.56
19	3	3	23	11	58	20.25	-27.12	-3.69	4.88	6.5
19	3	3	23	12	59	20.19	-27.06	-3.19	4.88	6.5
19	3	3	23	13	59	20.19	-27.25	-4.25	4.88	6.5
19	3	3	23	14	59	20.19	-27	-2.94	4.88	6.5
19	3	3	23	15	59	20.19	-27.19	-3.88	4.88	6.5
19	3	3	23	17	0	20.19	-26.87	-2.75	4.88	6.44
19	3	3	23	18	0	20.19	-27.12	-3.56	4.88	6.44
19	3	3	23	19	0	20.19	-26.87	-2.63	4.88	6.44
19	3	3	23	20	1	20.19	-27.06	-3.31	4.88	6.44
19	3	3	23	21	1	20.19	-27.25	-4.44	4.88	6.44
19	3	3	23	22	1	20.12	-27	-3.06	4.88	6.44
19	3	3	23	23	1	20.12	-27.19	-4	4.88	6.44
19	3	3	23	24	2	20.12	-26.94	-2.88	4.88	6.44
19	3	3	23	25	2	20.12	-27.12	-3.69	4.88	6.44
19	3	3	23	26	2	20.06	-26.81	-2.69	4.81	6.44
19	3	3	23	27	2	20.06	-26.94	-2.81	4.88	6.44
19	3	3	23	28	3	20.06	-26.75	-2.25	4.88	6.44
19	3	3	23	29	3	20.06	-26.94	-2.63	4.88	6.44
19	3	3	23	30	3	20.06	-27.12	-3.56	4.88	6.44
19	3	3	23	31	3	20.06	-26.87	-2.44	4.88	6.44
19	3	3	23	32	4	20.06	-27.12	-3.31	4.88	6.44
19	3	3	23	33	4	20.12	-26.81	-2.31	4.88	6.44
19	3	3	23	34	4	20.12	-27.06	-3	4.88	6.5
19	3	3	23	35	4	20.12	-26.81	-2.19	4.88	6.5
19	3	3	23	36	5	20.12	-27	-2.81	4.88	6.44
19	3	3	23	37	5	20.12	-27.12	-3.75	4.88	6.5
19	3	3	23	38	5	20.12	-26.87	-2.5	4.88	6.5
19	3	3	23	39	6	20.12	-27.12	-3.25	4.88	6.5
19	3	3	23	40	6	20.12	-26.87	-2.25	4.88	6.5
19	3	3	23	41	6	20.12	-27.06	-2.94	4.88	6.44
19	3	3	23	42	6	20.06	-26.81	-2.06	4.88	6.5
19	3	3	23	43	7	20.06	-27	-2.69	4.88	6.5
19	3	3	23	44	7	20.06	-27	-3	4.88	6.5
19	3	3	23	45	7	20	-26.94	-2.63	4.88	6.5
19	3	3	23	46	7	20	-27.12	-3.38	4.88	6.5
19	3	3	23	47	8	19.94	-26.87	-2.31	4.88	6.5
19	3	3	23	48	8	19.94	-27	-2.94	4.88	6.5
19	3	3	23	49	8	19.94	-26.81	-2	4.88	6.5
19	3	3	23	50	8	19.94	-26.94	-2.63	4.88	6.5
19	3	3	23	51	9	19.94	-26.75	-1.88	4.88	6.5
19	3	3	23	52	9	19.94	-26.87	-2.31	4.88	6.5
19	3	3	23	53	9	19.94	-27.06	-3	4.88	6.5
19	3	3	23	54	9	19.94	-26.81	-2	4.88	6.5
19	3	3	23	55	10	19.94	-27.06	-2.69	4.88	6.5
19	3	3	23	56	10	19.94	-26.81	-1.81	4.88	6.56
19	3	3	23	57	10	19.94	-27	-2.31	4.88	6.5
19	3	3	23	58	10	20	-26.75	-1.56	4.88	6.56
19	3	3	23	59	11	20	-26.87	-1.81	4.88	6.56
19	3	4	0	0	11	20	-27.06	-2.5	4.88	6.56
19	3	4	0	1	11	20.06	-26.87	-1.56	4.88	6.56
19	3	4	0	2	12	20.06	-27.06	-2.25	4.88	6.56
19	3	4	0	3	12	20.06	-26.81	-1.44	4.88	6.56
19	3	4	0	4	12	20.06	-27	-2.06	4.88	6.56
19	3	4	0	5	12	20.06	-26.75	-1.31	4.88	6.56
19	3	4	0	6	16	20.06	-26.87	-1.5	4.88	6.63
19	3	4	0	7	16	20.06	-27.06	-2.13	4.88	6.63
19	3	4	0	8	17	20	-26.87	-1.25	4.88	6.63
19	3	4	0	9	17	20	-27.06	-1.81	4.88	6.63
19	3	4	0	10	17	20	-26.81	-1.12	4.88	6.63
19	3	4	0	11	17	19.94	-27.06	-1.56	4.94	6.63
19	3	4	0	12	18	19.94	-27.06	-1.94	4.94	6.63
19	3	4	0	13	18	19.87	-27	-1.31	4.94	6.63
19	3	4	0	14	18	19.87	-27.19	-2	4.88	6.69
19	3	4	0	15	18	19.87	-26.94	-1.12	4.88	6.63
19	3	4	0	16	19	19.81	-27.06	-1.75	4.88	6.69
19	3	4	0	17	19	19.81	-26.87	-1	4.94	6.69
19	3	4	0	18	19	19.81	-27.06	-1.44	4.94	6.69
19	3	4	0	19	19	19.81	-27.12	-2.06	4.94	6.69
19	3	4	0	20	20	19.75	-26.94	-1.06	4.94	6.69
19	3	4	0	21	20	19.81	-27.12	-1.63	4.94	6.69
19	3	4	0	22	20	19.81	-26.87	-0.88	4.94	6.69
19	3	4	0	23	20	19.81	-27.06	-1.37	4.94	6.69
19	3	4	0	24	21	19.81	-26.81	-0.69	4.94	6.69
19	3	4	0	25	21	19.87	-27	-1.12	4.94	6.69
19	3	4	0	26	21	19.87	-27.25	-2.44	4.94	6.69
19	3	4	0	27	22	19.94	-27	-1.25	4.94	6.69
19	3	4	0	28	22	19.94	-27.12	-1.69	4.94	6.69
19	3	4	0	29	22	20	-26.94	-0.94	4.94	6.69
19	3	4	0	30	22	20	-27	-1.25	4.94	6.69
19	3	4	0	31	23	20	-26.81	-0.63	4.94	6.69
19	3	4	0	32	23	20	-27	-0.94	4.94	6.75
19	3	4	0	33	23	20	-27.12	-1.44	4.94	6.75
19	3	4	0	34	23	20	-26.94	-0.63	4.94	6.75
19	3	4	0	35	24	20	-27.06	-1.06	4.94	6.75
19	3	4	0	36	24	20	-26.87	-0.44	4.94	6.75
19	3	4	0	37	24	20	-27	-0.69	4.94	6.75
19	3	4	0	38	24	20	-27	-1	4.94	6.75
19	3	4	0	39	25	19.94	-26.94	-0.31	4.94	6.75
19	3	4	0	40	25	19.94	-27	-0.63	4.94	6.81
19	3	4	0	41	25	19.87	-26.87	-0.06	4.94	6.81
19	3	4	0	42	25	19.87	-27	-0.37	4.94	6.81
19	3	4	0	43	26	19.87	-26.81	0.12	4.94	6.81
19	3	4	0	44	26	19.81	-27	-0.19	4.94	6.81
19	3	4	0	45	26	19.81	-27.06	-0.63	4.94	6.88
19	3	4	0	46	27	19.81	-26.94	0.06	4.94	6.88
19	3	4	0	47	27	19.81	-27	-0.19	4.94	6.88
19	3	4	0	48	27	19.75	-26.87	0.31	5	6.88
19	3	4	0	49	27	19.75	-27	0.06	4.94	6.88
19	3	4	0	50	28	19.81	-27	-0.25	5	6.94
19	3	4	0	51	28	19.81	-26.94	0.37	5	6.94
19	3	4	0	52	28	19.81	-27.06	0	5	6.94
19	3	4	0	53	28	19.81	-26.87	0.56	5	7
19	3	4	0	54	29	19.81	-27	0.25	5	7
19	3	4	0	55	29	19.81	-26.87	0.69	5	7
19	3	4	0	56	29	19.87	-26.94	0.44	5	7
19	3	4	0	57	29	19.81	-27	0.25	5	7
19	3	4	0	58	30	19.81	-26.87	0.81	5	7
19	3	4	0	59	30	19.81	-27	0.56	5	7.06
19	3	4	1	0	30	19.81	-26.81	1	5	7.06
19	3	4	1	1	30	19.81	-26.94	0.81	5	7.13
19	3	4	1	2	31	19.75	-26.94	0.5	5	7.13
19	3	4	1	3	31	19.75	-26.87	1.06	5	7.13
19	3	4	1	4	31	19.75	-26.94	0.75	5	7.13
19	3	4	1	5	31	19.69	-26.81	1.25	5	7.13
19	3	4	1	6	32	19.69	-26.94	1	5	7.19
19	3	4	1	7	32	19.69	-26.81	1.06	5	7.19
19	3	4	1	8	32	19.62	-26.81	1.37	5	7.19
19	3	4	1	9	33	19.62	-26.87	1.12	5.06	7.19
19	3	4	1	10	33	19.62	-26.81	1.56	5.06	7.25
19	3	4	1	11	33	19.62	-26.87	1.37	5.06	7.25
19	3	4	1	12	33	19.62	-26.75	1.75	5.06	7.25
19	3	4	1	13	34	19.62	-26.81	1.56	5.06	7.31
19	3	4	1	14	34	19.62	-26.87	1.44	5.06	7.31
19	3	4	1	15	34	19.62	-26.75	1.88	5.06	7.31
19	3	4	1	16	34	19.62	-26.87	1.63	5.06	7.31
19	3	4	1	17	35	19.62	-26.75	2	5.06	7.31
19	3	4	1	18	35	19.56	-26.87	1.81	5.06	7.31
19	3	4	1	19	35	19.56	-26.94	1.5	5.06	7.38
19	3	4	1	20	35	19.56	-26.81	1.94	5.06	7.38
19	3	4	1	21	36	19.5	-26.94	1.75	5.06	7.38
19	3	4	1	22	36	19.5	-26.81	2.06	5.06	7.38
19	3	4	1	23	36	19.5	-26.87	1.88	5.06	7.44
19	3	4	1	24	36	19.5	-26.87	1.5	5.06	7.44
19	3	4	1	25	37	19.5	-26.87	2.06	5.06	7.44
19	3	4	1	26	37	19.5	-27.12	1.19	5.06	7.44
19	3	4	1	27	37	19.56	-27	1.69	5.06	7.44
19	3	4	1	28	37	19.56	-27.12	1.44	5.06	7.44
19	3	4	1	29	38	19.56	-26.94	1.81	5.06	7.44
19	3	4	1	30	38	19.62	-26.94	1.81	5.06	7.44
19	3	4	1	31	38	19.62	-27.06	1.44	5.06	7.44
19	3	4	1	32	38	19.69	-26.94	1.94	5.06	7.44
19	3	4	1	33	39	19.69	-27	1.56	5.06	7.44
19	3	4	1	34	39	19.69	-26.87	2	5.06	7.44
19	3	4	1	35	39	19.69	-27	1.75	5.06	7.44
19	3	4	1	36	40	19.69	-27.12	1.12	5.06	7.44
19	3	4	1	37	40	19.69	-27	1.75	5.06	7.38
19	3	4	1	38	40	19.69	-27.12	1.31	5.06	7.44
19	3	4	1	39	40	19.69	-27	1.81	5.06	7.44
19	3	4	1	40	41	19.62	-27.12	1.44	5.06	7.38
19	3	4	1	41	41	19.62	-26.94	1.5	5.06	7.38
19	3	4	1	42	41	19.62	-27.19	1.31	5.06	7.38
19	3	4	1	43	41	19.62	-27.31	0.88	5.06	7.38
19	3	4	1	44	42	19.56	-27.12	1.44	5.06	7.38
19	3	4	1	45	42	19.56	-27.25	1.06	5.06	7.38
19	3	4	1	46	42	19.56	-27.06	1.56	5.06	7.31
19	3	4	1	47	42	19.56	-27.19	1.12	5.06	7.31
19	3	4	1	48	43	19.56	-27.25	0.44	5.06	7.31
19	3	4	1	49	43	19.56	-27.12	1.19	5.06	7.31
19	3	4	1	50	43	19.62	-27.25	0.69	5.06	7.31
19	3	4	1	51	43	19.62	-27.06	1.25	5.06	7.31
19	3	4	1	52	44	19.62	-27.25	0.88	5.06	7.25
19	3	4	1	53	44	19.62	-27	1.37	5.06	7.25
19	3	4	1	54	44	19.62	-27.19	1.06	5.06	7.25
19	3	4	1	55	44	19.62	-27.25	0.44	5.06	7.25
19	3	4	1	56	45	19.62	-27.06	1.12	5.06	7.25
19	3	4	1	57	45	19.62	-27.25	0.56	5.06	7.25
19	3	4	1	58	45	19.62	-27.06	1.12	5.06	7.19
19	3	4	1	59	46	19.56	-27.19	0.63	5.06	7.19
19	3	4	2	0	46	19.56	-27.31	-0.12	5.06	7.19
19	3	4	2	1	46	19.56	-27.12	0.75	5.06	7.19
19	3	4	2	2	46	19.56	-27.25	0.12	5.06	7.19
19	3	4	2	3	47	19.5	-27.06	0.81	5.06	7.13
19	3	4	2	4	47	19.5	-27.25	0.37	5.06	7.19
19	3	4	2	5	47	19.5	-27	0.94	5.06	7.13
19	3	4	2	6	47	19.44	-27.19	0.44	5.06	7.13
19	3	4	2	7	48	19.44	-27.31	-0.37	5.06	7.13
19	3	4	2	8	48	19.44	-27.06	0.5	5.06	7.13
19	3	4	2	9	48	19.5	-27.19	-0.19	5.06	7.13
19	3	4	2	10	48	19.5	-27	0.56	5.06	7.13
19	3	4	2	11	49	19.5	-27.12	0	5.06	7.06
19	3	4	2	12	49	19.5	-26.87	0.63	5.06	7.06
19	3	4	2	13	49	19.5	-27	0.44	5.06	7.06
19	3	4	2	14	49	19.5	-27.19	-0.31	5.06	7.06
19	3	4	2	15	50	19.56	-27	0.44	5.06	7.06
19	3	4	2	16	50	19.56	-27.19	-0.25	5.06	7.06
19	3	4	2	17	50	19.56	-27	0.44	5.06	7.06
19	3	4	2	18	50	19.56	-27.12	-0.19	5	7.06
19	3	4	2	19	51	19.56	-26.94	0.12	5	7.06
19	3	4	2	20	51	19.56	-27.06	-0.06	5	7
19	3	4	2	21	51	19.56	-27.19	-0.94	5	7
19	3	4	2	22	51	19.56	-26.94	-0.06	5	7
19	3	4	2	23	52	19.56	-27.12	-0.69	5	7
19	3	4	2	24	52	19.56	-26.94	0	5	7
19	3	4	2	25	52	19.5	-27.06	-0.56	5	6.94
19	3	4	2	26	53	19.5	-27	-0.63	5	6.94
19	3	4	2	27	53	19.5	-27	-0.56	5	6.94
19	3	4	2	28	53	19.5	-27.12	-1.44	5	6.94
19	3	4	2	29	53	19.5	-26.87	-0.5	5	6.94
19	3	4	2	30	54	19.5	-27.06	-1.19	5	6.88
19	3	4	2	31	54	19.5	-26.81	-0.44	5	6.88
19	3	4	2	32	54	19.5	-27	-0.94	5	6.88
19	3	4	2	33	54	19.5	-26.81	-0.69	5	6.88
19	3	4	2	34	54	19.5	-26.94	-0.81	5	6.88
19	3	4	2	35	54	19.5	-27.06	-1.63	5	6.88
19	3	4	2	36	55	19.5	-26.81	-0.69	5	6.81
19	3	4	2	37	55	19.56	-27	-1.5	5	6.81
19	3	4	2	38	55	19.56	-26.75	-0.69	5	6.81
19	3	4	2	39	55	19.56	-26.87	-1.37	4.94	6.81
19	3	4	2	40	56	19.56	-26.62	-0.63	5	6.81
19	3	4	2	41	56	19.56	-26.81	-1.19	4.94	6.81
19	3	4	2	42	56	19.56	-27	-2.19	4.94	6.81
19	3	4	2	43	56	19.5	-26.75	-1.06	4.94	6.75
19	3	4	2	44	57	19.5	-26.94	-1.88	4.94	6.75
19	3	4	2	45	57	19.5	-26.69	-1	4.94	6.75
19	3	4	2	46	57	19.44	-26.62	-0.88	4.94	6.75
19	3	4	2	47	57	19.44	-26.44	-0.31	4.94	6.75
19	3	4	2	48	58	19.44	-26.69	-0.81	4.94	6.75
19	3	4	2	49	58	19.37	-26.87	-1.81	5	6.81
19	3	4	2	50	58	19.37	-26.69	-0.81	4.94	6.75
19	3	4	2	51	58	19.37	-26.87	-1.69	4.94	6.75
19	3	4	2	52	59	19.37	-26.62	-0.81	4.94	6.75
19	3	4	2	53	59	19.37	-26.87	-1.44	4.94	6.75
19	3	4	2	54	59	19.37	-26.62	-0.69	4.94	6.75
19	3	4	2	55	59	19.37	-26.81	-1.37	4.94	6.75
19	3	4	2	57	0	19.37	-26.94	-2.19	4.94	6.75
19	3	4	2	58	0	19.37	-26.69	-1.12	4.94	6.75
19	3	4	2	59	0	19.37	-26.87	-1.81	4.94	6.75
19	3	4	3	0	0	19.37	-26.62	-1	4.94	6.75
19	3	4	3	1	1	19.44	-26.75	-1.56	4.94	6.75
19	3	4	3	2	1	19.44	-26.5	-0.81	4.94	6.75
19	3	4	3	3	1	19.44	-26.69	-1.37	4.94	6.69
19	3	4	3	4	1	19.5	-26.87	-2.06	4.94	6.69
19	3	4	3	5	2	19.5	-26.62	-1.06	4.94	6.69
19	3	4	3	6	2	19.5	-26.75	-1.75	4.94	6.69
19	3	4	3	7	2	19.5	-26.56	-0.94	4.94	6.69
19	3	4	3	8	3	19.5	-26.69	-1.44	4.94	6.69
19	3	4	3	9	3	19.5	-26.44	-0.75	4.94	6.69
19	3	4	3	10	3	19.5	-26.62	-1.19	4.94	6.69
19	3	4	3	11	3	19.44	-26.75	-1.75	4.94	6.69
19	3	4	3	12	4	19.44	-26.56	-0.88	4.94	6.69
19	3	4	3	13	4	19.44	-26.75	-1.37	4.94	6.75
19	3	4	3	14	4	19.37	-26.5	-0.69	4.94	6.75
19	3	4	3	15	4	19.37	-26.69	-1.19	4.94	6.75
19	3	4	3	16	5	19.31	-26.5	-0.5	4.94	6.75
19	3	4	3	17	5	19.31	-26.62	-0.88	4.94	6.75
19	3	4	3	18	5	19.31	-26.81	-1.44	4.94	6.75
19	3	4	3	19	5	19.25	-26.56	-0.56	4.94	6.75
19	3	4	3	20	6	19.25	-26.75	-1.06	4.94	6.75
19	3	4	3	21	6	19.25	-26.5	-0.37	4.94	6.75
19	3	4	3	22	6	19.25	-26.69	-0.81	4.94	6.75
19	3	4	3	23	6	19.25	-26.5	-0.44	4.94	6.75
19	3	4	3	24	7	19.31	-26.62	-0.56	4.94	6.75
19	3	4	3	25	7	19.31	-26.81	-1.06	4.94	6.75
19	3	4	3	26	7	19.31	-26.56	-0.31	4.94	6.75
19	3	4	3	27	7	19.37	-26.75	-0.75	4.94	6.81
19	3	4	3	28	8	19.37	-26.56	-0.12	4.94	6.81
19	3	4	3	29	8	19.44	-26.69	-0.44	5	6.81
19	3	4	3	30	8	19.44	-26.75	-1	4.94	6.81
19	3	4	3	31	8	19.44	-26.56	-0.19	4.94	6.81
19	3	4	3	32	9	19.44	-26.75	-0.63	5	6.81
19	3	4	3	33	9	19.44	-26.56	0	4.94	6.81
19	3	4	3	34	9	19.44	-26.69	-0.44	4.94	6.88
19	3	4	3	35	9	19.37	-26.5	0.12	4.94	6.81
19	3	4	3	36	10	19.37	-26.62	-0.19	4.94	6.88
19	3	4	3	37	10	19.37	-26.75	-0.56	5	6.88
19	3	4	3	38	10	19.31	-26.56	0.12	5	6.88
19	3	4	3	39	10	19.31	-26.69	-0.25	5	6.88
19	3	4	3	40	11	19.31	-26.5	0.37	5	6.88
19	3	4	3	41	11	19.31	-26.62	0	5	6.88
19	3	4	3	42	11	19.31	-26.56	-0.19	5	6.94
19	3	4	3	43	12	19.25	-26.56	0.31	4.94	6.94
19	3	4	3	44	12	19.31	-26.69	-0.12	5	6.94
19	3	4	3	45	12	19.31	-26.5	0.5	5	6.94
19	3	4	3	46	12	19.31	-26.62	0.12	5	6.94
19	3	4	3	47	13	19.31	-26.5	0.63	5	6.94
19	3	4	3	48	13	19.31	-26.56	0.37	5	6.94
19	3	4	3	49	13	19.31	-26.62	0.06	5	6.94
19	3	4	3	50	13	19.37	-26.5	0.69	5	7
19	3	4	3	51	14	19.37	-26.62	0.37	5	7
19	3	4	3	52	14	19.37	-26.5	0.88	5	7
19	3	4	3	53	14	19.37	-26.62	0.56	5	7
19	3	4	3	54	14	19.37	-26.62	0.19	5	7.06
19	3	4	3	55	15	19.37	-26.5	0.88	5	7.06
19	3	4	3	56	17	19.37	-26.62	0.5	5	7.06
19	3	4	3	57	17	19.37	-26.5	1	5	7.06
19	3	4	3	58	17	19.31	-26.56	0.81	5	7.06
19	3	4	3	59	18	19.31	-26.44	1.25	5	7.06
19	3	4	4	0	18	19.25	-26.31	1.31	5	7.06
19	3	4	4	1	18	19.25	-26.37	1.12	5	7.13
19	3	4	4	2	18	19.25	-26.25	1.56	5.06	7.13
19	3	4	4	3	19	19.19	-26.37	1.25	5.06	7.13
19	3	4	4	4	19	19.19	-26.25	1.63	5.06	7.19
19	3	4	4	5	19	19.19	-26.44	1.37	5.06	7.19
19	3	4	4	6	19	19.19	-26.5	1.12	5.06	7.19
19	3	4	4	7	20	19.19	-26.37	1.63	5.06	7.19
19	3	4	4	8	20	19.19	-26.5	1.31	5.06	7.19
19	3	4	4	9	20	19.19	-26.37	1.75	5.06	7.19
19	3	4	4	10	20	19.19	-26.5	1.44	5.06	7.25
19	3	4	4	11	21	19.19	-26.5	1	5.06	7.25
19	3	4	4	12	21	19.19	-26.5	1.56	5.06	7.25
19	3	4	4	13	21	19.19	-26.56	1.25	5.06	7.25
19	3	4	4	14	21	19.19	-26.5	1.69	5.06	7.25
19	3	4	4	15	22	19.19	-26.56	1.44	5.06	7.25
19	3	4	4	16	22	19.19	-26.44	1.81	5.06	7.25
19	3	4	4	17	22	19.19	-26.44	1.75	5.06	7.25
19	3	4	4	18	23	19.19	-26.56	1.25	5.06	7.31
19	3	4	4	19	23	19.19	-26.44	1.75	5.06	7.31
19	3	4	4	20	23	19.19	-26.56	1.31	5.06	7.31
19	3	4	4	21	23	19.19	-26.44	1.75	5.06	7.31
19	3	4	4	22	24	19.12	-26.62	1.44	5.06	7.31
19	3	4	4	23	24	19.19	-26.75	0.81	5.06	7.31
19	3	4	4	24	24	19.19	-26.56	1.5	5.06	7.25
19	3	4	4	25	24	19.19	-26.69	1.06	5.06	7.25
19	3	4	4	26	25	19.19	-26.56	1.56	5.06	7.25
19	3	4	4	27	25	19.19	-26.69	1.19	5.06	7.25
19	3	4	4	28	25	19.25	-26.56	1.63	5.06	7.25
19	3	4	4	29	25	19.25	-26.69	1.25	5.06	7.25
19	3	4	4	30	26	19.25	-26.81	0.63	5.06	7.25
19	3	4	4	31	26	19.25	-26.62	1.25	5.06	7.25
19	3	4	4	32	26	19.31	-26.75	0.81	5.06	7.25
19	3	4	4	33	26	19.31	-26.56	1.31	5.06	7.25
19	3	4	4	34	27	19.31	-26.75	0.94	5.06	7.19
19	3	4	4	35	27	19.25	-26.81	0.25	5.06	7.19
19	3	4	4	36	27	19.25	-26.69	1.06	5.06	7.19
19	3	4	4	37	27	19.25	-26.81	0.44	5.06	7.19
19	3	4	4	38	28	19.19	-26.62	1.06	5.06	7.19
19	3	4	4	39	28	19.19	-26.81	0.56	5.06	7.19
19	3	4	4	40	28	19.12	-26.56	1.12	5.06	7.19
19	3	4	4	41	28	19.12	-26.75	0.63	5.06	7.19
19	3	4	4	42	29	19.12	-26.87	0	5	7.13
19	3	4	4	43	29	19.12	-26.69	0.75	5	7.13
19	3	4	4	44	29	19.12	-26.81	0.19	5.06	7.13
19	3	4	4	45	29	19.12	-26.56	0.81	5	7.13
19	3	4	4	46	30	19.12	-26.69	0.31	5	7.06
19	3	4	4	47	30	19.12	-26.56	0.44	5	7.06
19	3	4	4	48	30	19.12	-26.69	0.31	5	7.06
19	3	4	4	49	31	19.12	-26.81	-0.31	5.06	7.06
19	3	4	4	50	31	19.12	-26.62	0.44	5	7.06
19	3	4	4	51	31	19.12	-26.75	-0.12	5	7.06
19	3	4	4	52	31	19.19	-26.5	0.5	5	7
19	3	4	4	53	32	19.19	-26.69	0	5	7
19	3	4	4	54	32	19.19	-26.75	-0.69	5	7
19	3	4	4	55	32	19.19	-26.62	0.06	5	7
19	3	4	4	56	32	19.19	-26.81	-1.19	5	7
19	3	4	4	57	33	19.25	-26.56	-0.25	5	6.94
19	3	4	4	58	33	19.25	-26.69	-0.81	5	6.94
19	3	4	4	59	33	19.25	-26.44	-0.12	5	6.94
19	3	4	5	0	33	19.25	-26.56	-0.56	5	6.94
19	3	4	5	1	34	19.25	-26.62	-1.19	5	6.88
19	3	4	5	2	34	19.25	-26.5	-0.37	5	6.88
19	3	4	5	3	34	19.19	-26.62	-1.12	5	6.88
19	3	4	5	4	34	19.19	-26.44	-0.25	5	6.88
19	3	4	5	5	35	19.12	-26.56	-0.81	5	6.88
19	3	4	5	6	35	19.12	-26.31	-0.12	5	6.88
19	3	4	5	7	35	19.12	-26.5	-0.69	4.94	6.88
19	3	4	5	8	35	19.12	-26.62	-1.5	5	6.88
19	3	4	5	9	36	19.06	-26.44	-0.56	5	6.81
19	3	4	5	10	36	19.06	-26.62	-1.31	5	6.81
19	3	4	5	11	36	19.06	-26.37	-0.5	4.94	6.81
19	3	4	5	12	36	19.06	-26.56	-1.12	4.94	6.81
19	3	4	5	13	37	19	-26.31	-0.37	4.94	6.81
19	3	4	5	14	37	19.06	-26.44	-0.94	4.94	6.81
19	3	4	5	15	37	19.06	-26.5	-1.63	5	6.81
19	3	4	5	16	38	19.06	-26.37	-0.75	4.94	6.81
19	3	4	5	17	38	19.06	-26.56	-1.56	4.94	6.81
19	3	4	5	18	38	19.06	-26.31	-0.69	4.94	6.81
19	3	4	5	19	38	19.06	-26.5	-1.44	4.94	6.75
19	3	4	5	20	39	19.06	-26.25	-0.63	4.94	6.75
19	3	4	5	21	39	19.06	-26.44	-1.25	4.94	6.75
19	3	4	5	22	39	19.06	-26.31	-1.37	4.94	6.75
19	3	4	5	23	39	19.06	-26.31	-1.12	4.94	6.75
19	3	4	5	24	40	19.06	-26.5	-1.94	4.94	6.75
19	3	4	5	25	40	19.12	-26.25	-1	4.94	6.69
19	3	4	5	26	40	19.12	-26.44	-1.69	4.94	6.69
19	3	4	5	27	40	19.12	-26.19	-0.88	4.94	6.69
19	3	4	5	28	41	19.12	-26.31	-1.44	4.94	6.69
19	3	4	5	29	41	19.19	-26.06	-0.81	4.94	6.69
19	3	4	5	30	41	19.12	-26.25	-1.25	4.94	6.69
19	3	4	5	31	41	19.19	-26.37	-2	4.94	6.69
19	3	4	5	32	42	19.12	-26.19	-1.06	4.94	6.69
19	3	4	5	33	42	19.12	-26.31	-1.75	4.94	6.69
19	3	4	5	34	42	19.12	-26.12	-0.94	4.94	6.69
19	3	4	5	35	42	19.12	-26.25	-1.56	4.94	6.69
19	3	4	5	36	43	19.06	-26.06	-0.81	4.94	6.69
19	3	4	5	37	43	19.06	-26.19	-1	4.94	6.69
19	3	4	5	38	43	19	-26.37	-1.75	4.94	6.69
19	3	4	5	39	43	19	-26.12	-0.81	4.94	6.69
19	3	4	5	40	44	19	-26.31	-1.5	4.94	6.69
19	3	4	5	41	44	19	-26.06	-0.75	4.94	6.69
19	3	4	5	42	44	19	-26.25	-1.25	4.94	6.69
19	3	4	5	43	44	19	-26	-0.63	4.94	6.69
19	3	4	5	44	45	19	-26.19	-0.88	4.94	6.69
19	3	4	5	45	45	19.06	-26.31	-1.56	4.94	6.69
19	3	4	5	46	45	19.06	-26.12	-0.69	4.94	6.69
19	3	4	5	47	45	19.06	-26.25	-1.12	4.94	6.69
19	3	4	5	48	46	19.06	-26.06	-0.44	4.94	6.75
19	3	4	5	49	46	19.06	-26.19	-0.94	4.94	6.75
19	3	4	5	50	46	19.06	-26	-0.31	4.94	6.75
19	3	4	5	51	47	19.06	-26.19	-0.81	4.94	6.75
19	3	4	5	52	47	19.06	-26.31	-1.37	4.94	6.75
19	3	4	5	53	47	19.06	-26.12	-0.56	4.94	6.75
19	3	4	5	54	47	19.06	-26.25	-1.19	4.94	6.75
19	3	4	5	55	48	19.06	-26.06	-0.44	4.94	6.75
19	3	4	5	56	48	19	-26.19	-0.94	4.94	6.75
19	3	4	5	57	48	19	-26	-0.63	4.94	6.75
19	3	4	5	58	48	19	-26.12	-0.69	4.94	6.75
19	3	4	5	59	49	19	-26.25	-1.19	4.94	6.75
19	3	4	6	0	49	19	-26.06	-0.37	4.94	6.75
19	3	4	6	1	49	19	-26.19	-0.88	4.94	6.75
19	3	4	6	2	49	19	-26	-0.25	4.94	6.75
19	3	4	6	3	50	19	-26.12	-0.69	4.94	6.75
19	3	4	6	4	50	19	-26.19	-1.25	4.94	6.75
19	3	4	6	5	50	19	-26.06	-0.37	4.94	6.75
19	3	4	6	6	50	19	-26.25	-0.88	4.94	6.75
19	3	4	6	7	51	19	-26.06	-0.19	4.94	6.75
19	3	4	6	8	51	19	-26.19	-0.56	4.94	6.75
19	3	4	6	9	51	19	-26	0.06	4.94	6.75
19	3	4	6	10	51	19	-26.12	-0.31	4.94	6.81
19	3	4	6	11	52	19	-26.25	-0.81	4.94	6.81
19	3	4	6	12	52	19.06	-26.12	0	4.94	6.81
19	3	4	6	13	52	19.06	-26.25	-0.44	4.94	6.81
19	3	4	6	14	52	19.06	-26.06	0.12	4.94	6.81
19	3	4	6	15	53	19.06	-26.12	-0.19	4.94	6.81
19	3	4	6	16	53	19.06	-25.94	0.25	4.94	6.81
19	3	4	6	17	53	19.12	-26.06	0.06	4.94	6.81
19	3	4	6	18	53	19.12	-26.19	-0.25	4.94	6.88
19	3	4	6	19	54	19.12	-26	0.37	4.94	6.88
19	3	4	6	20	54	19.12	-26.12	-0.06	4.94	6.88
19	3	4	6	21	54	19.12	-25.94	0.5	4.94	6.88
19	3	4	6	22	54	19.12	-26.06	0.12	4.94	6.88
19	3	4	6	23	55	19.12	-26.19	-0.31	4.94	6.94
19	3	4	6	24	55	19.06	-26	0.44	4.94	6.94
19	3	4	6	25	55	19.06	-26.12	0	4.94	6.88
19	3	4	6	26	56	19.06	-25.94	0.56	5	6.88
19	3	4	6	27	56	19	-26.06	0.25	5	6.94
19	3	4	6	28	56	19	-25.87	0.75	5	6.94
19	3	4	6	29	56	18.94	-25.94	0.56	5	6.94
19	3	4	6	30	57	18.94	-26.06	0.37	5	6.94
19	3	4	6	31	57	18.94	-25.87	0.88	5	6.94
19	3	4	6	32	57	18.94	-26	0.56	5	7
19	3	4	6	33	57	19	-25.87	1	5	7
19	3	4	6	34	58	18.94	-26	0.75	5	7
19	3	4	6	35	58	19	-26.06	0.37	5	7
19	3	4	6	36	58	19	-25.94	1	5	7
19	3	4	6	37	58	19	-26	0.75	5	7.06
19	3	4	6	38	59	19.06	-25.81	1.19	5	7.06
19	3	4	6	39	59	19.06	-25.94	1	5	7.06
19	3	4	6	40	59	19.06	-25.87	0.75	5	7.06
19	3	4	6	41	59	19.12	-25.81	1.31	5	7.06
19	3	4	6	43	0	19.12	-25.94	1	5	7.06
19	3	4	6	44	0	19.19	-25.81	1.44	5	7.13
19	3	4	6	45	0	19.19	-25.87	1.25	5	7.13
19	3	4	6	46	0	19.19	-25.75	1.63	5	7.13
19	3	4	6	47	1	19.19	-25.87	1.5	5	7.13
19	3	4	6	48	1	19.19	-25.87	1.44	5.06	7.13
19	3	4	6	49	1	19.19	-25.75	1.81	5	7.19
19	3	4	6	50	1	19.19	-25.81	1.69	5	7.19
19	3	4	6	51	2	19.19	-25.69	2	5.06	7.19
19	3	4	6	52	2	19.19	-25.81	1.88	5.06	7.25
19	3	4	6	53	2	19.19	-25.87	1.75	5.06	7.25
19	3	4	6	54	2	19.19	-25.75	2.13	5.06	7.25
19	3	4	6	55	3	19.19	-25.81	1.88	5.06	7.31
19	3	4	6	56	3	19.19	-25.69	2.25	5.06	7.31
19	3	4	6	57	3	19.19	-25.94	1.56	5.06	7.31
19	3	4	6	58	4	19.12	-25.94	1.25	5.06	7.31
19	3	4	6	59	4	19.12	-25.87	1.81	5.06	7.31
19	3	4	7	0	4	19.12	-25.94	1.63	5.06	7.31
19	3	4	7	1	4	19.19	-25.81	2	5.06	7.31
19	3	4	7	2	5	19.19	-25.87	1.81	5.06	7.31
19	3	4	7	3	5	19.19	-25.75	2.19	5.06	7.31
19	3	4	7	4	5	19.25	-25.62	2.38	5.06	7.31
19	3	4	7	5	5	19.25	-25.69	2.13	5.06	7.38
19	3	4	7	6	6	19.25	-25.62	2.44	5.06	7.38
19	3	4	7	7	6	19.25	-25.81	2.13	5.06	7.38
19	3	4	7	8	6	19.25	-25.69	2.44	5.06	7.38
19	3	4	7	9	6	19.25	-25.81	2.19	5.06	7.38
19	3	4	7	10	7	19.25	-25.87	1.75	5.06	7.38
19	3	4	7	11	7	19.25	-25.81	2.25	5.06	7.38
19	3	4	7	12	7	19.25	-25.94	1.94	5.06	7.38
19	3	4	7	13	7	19.25	-25.81	2.31	5.06	7.38
19	3	4	7	14	8	19.25	-25.87	2.06	5.13	7.38
19	3	4	7	15	8	19.25	-26	1.5	5.06	7.38
19	3	4	7	16	8	19.25	-25.87	2.13	5.06	7.38
19	3	4	7	17	8	19.31	-26	1.69	5.06	7.38
19	3	4	7	18	9	19.31	-25.87	2.13	5.13	7.38
19	3	4	7	19	9	19.31	-26	1.75	5.13	7.38
19	3	4	7	20	9	19.37	-25.81	2.19	5.06	7.38
19	3	4	7	21	9	19.37	-25.94	1.94	5.13	7.38
19	3	4	7	22	10	19.37	-26	1.44	5.06	7.38
19	3	4	7	23	10	19.37	-25.87	2	5.13	7.38
19	3	4	7	24	10	19.44	-25.94	1.63	5.06	7.38
19	3	4	7	25	10	19.44	-25.81	2.06	5.13	7.38
19	3	4	7	26	11	19.44	-25.94	1.75	5.13	7.38
19	3	4	7	27	11	19.44	-26.06	1	5.06	7.38
19	3	4	7	28	11	19.44	-25.94	1.75	5.06	7.38
19	3	4	7	29	12	19.44	-26.06	1.25	5.06	7.31
19	3	4	7	30	12	19.44	-25.87	1.75	5.06	7.31
19	3	4	7	31	12	19.37	-26	1.37	5.06	7.31
19	3	4	7	32	12	19.37	-25.87	1.81	5.06	7.31
19	3	4	7	33	13	19.31	-25.94	1.56	5.06	7.31
19	3	4	7	34	13	19.31	-26.12	1	5.06	7.31
19	3	4	7	35	13	19.25	-25.87	1.63	5.06	7.31
19	3	4	7	36	13	19.25	-26	1.19	5.06	7.31
19	3	4	7	37	14	19.25	-25.87	1.69	5.06	7.31
19	3	4	7	38	14	19.19	-26	1.19	5.06	7.31
19	3	4	7	39	14	19.25	-26.12	0.56	5.06	7.25
19	3	4	7	40	14	19.19	-26	1.25	5.06	7.25
19	3	4	7	41	15	19.25	-26.12	0.69	5.06	7.25
19	3	4	7	42	15	19.25	-25.94	1.31	5.06	7.25
19	3	4	7	43	15	19.31	-26.06	0.81	5.06	7.25
19	3	4	7	44	15	19.31	-25.87	1.37	5.06	7.25
19	3	4	7	45	16	19.31	-25.94	0.94	5.06	7.25
19	3	4	7	46	19	19.31	-26.06	0.31	5.06	7.19
19	3	4	7	47	20	19.31	-25.87	1.06	5.06	7.19
19	3	4	7	48	20	19.31	-26	0.5	5.06	7.19
19	3	4	7	49	20	19.31	-25.81	1.06	5.06	7.19
19	3	4	7	50	20	19.31	-25.94	0.56	5.06	7.19
19	3	4	7	51	21	19.31	-25.75	0.81	5.06	7.19
19	3	4	7	52	21	19.25	-25.87	0.69	5.06	7.13
19	3	4	7	53	21	19.25	-26	0	5.06	7.19
19	3	4	7	54	21	19.25	-25.75	0.75	5.06	7.13
19	3	4	7	55	22	19.25	-25.94	0.19	5.06	7.13
19	3	4	7	56	22	19.25	-25.69	0.81	5.06	7.13
19	3	4	7	57	22	19.25	-25.87	0.25	5.06	7.13
19	3	4	7	58	22	19.25	-25.94	-0.31	5.06	7.13
19	3	4	7	59	23	19.31	-25.75	0.31	5.06	7.06
19	3	4	8	0	23	19.25	-25.94	-0.37	5.06	7.06
19	3	4	8	1	23	19.25	-25.69	0.37	5.06	7.06
19	3	4	8	2	24	19.25	-25.87	-0.19	5.06	7.06
19	3	4	8	3	24	19.25	-25.62	0.5	5.06	7.06
19	3	4	8	4	24	19.25	-25.75	0	5.06	7
19	3	4	8	5	24	19.19	-25.87	-0.81	5.06	7
19	3	4	8	6	25	19.19	-25.69	0.06	5.06	7
19	3	4	8	7	25	19.19	-25.87	-0.69	5.06	7
19	3	4	8	8	25	19.12	-25.62	0.12	5.06	7
19	3	4	8	9	25	19.12	-25.75	-0.56	5	7
19	3	4	8	10	26	19.12	-25.5	0.12	5.06	7
19	3	4	8	11	26	19.12	-25.69	-0.37	5	6.94
19	3	4	8	12	26	19.12	-25.81	-1.19	5	6.94
19	3	4	8	13	26	19.12	-25.62	-0.31	5	6.94
19	3	4	8	14	27	19.19	-25.75	-1	5	6.94
19	3	4	8	15	27	19.19	-25.5	-0.19	5	6.94
19	3	4	8	16	27	19.19	-25.69	-0.81	5	6.94
19	3	4	8	17	27	19.25	-25.44	-0.12	5	6.88
19	3	4	8	18	28	19.25	-25.56	-0.75	5	6.88
19	3	4	8	19	28	19.25	-25.62	-1.31	5	6.88
19	3	4	8	20	28	19.25	-25.5	-0.81	5	6.88
19	3	4	8	21	28	19.25	-25.69	-1.69	5	6.88
19	3	4	8	22	29	19.25	-25.37	-0.69	5	6.81
19	3	4	8	23	29	19.25	-25.56	-1.31	5	6.81
19	3	4	8	24	29	19.25	-25.31	-0.56	5	6.81
19	3	4	8	25	29	19.25	-25.44	-1.12	5	6.81
19	3	4	8	26	30	19.25	-25.31	-1	5	6.81
19	3	4	8	27	30	19.25	-25.37	-1	5	6.81
19	3	4	8	28	30	19.25	-25.56	-1.94	5	6.81
19	3	4	8	29	30	19.25	-25.31	-0.94	4.94	6.75
19	3	4	8	30	31	19.25	-25.5	-1.69	4.94	6.75
19	3	4	8	31	31	19.25	-25.25	-0.81	4.94	6.75
19	3	4	8	32	31	19.31	-25.37	-1.5	5	6.75
19	3	4	8	33	31	19.31	-25.19	-0.75	4.94	6.75
19	3	4	8	34	32	19.37	-25.31	-1.19	4.94	6.75
19	3	4	8	35	32	19.37	-25.5	-2.25	4.94	6.69
19	3	4	8	36	32	19.37	-25.25	-1.12	4.94	6.69
19	3	4	8	37	33	19.37	-25.37	-1.94	4.94	6.69
19	3	4	8	38	33	19.37	-25.12	-1	4.94	6.69
19	3	4	8	39	33	19.37	-25.31	-1.56	4.94	6.69
19	3	4	8	40	33	19.37	-25.06	-0.81	4.94	6.69
19	3	4	8	41	34	19.37	-25.25	-1.44	4.94	6.69
19	3	4	8	42	34	19.37	-25.44	-2.38	4.94	6.69
19	3	4	8	43	34	19.37	-25.12	-1.25	4.94	6.69
19	3	4	8	44	34	19.31	-25.37	-2	4.94	6.69
19	3	4	8	45	35	19.31	-25.06	-1.12	4.94	6.69
19	3	4	8	46	35	19.31	-25.25	-1.75	4.94	6.69
19	3	4	8	47	35	19.31	-25	-1	4.94	6.69
19	3	4	8	48	35	19.37	-25.19	-1.5	4.94	6.69
19	3	4	8	49	36	19.37	-25.19	-1.75	4.94	6.69
19	3	4	8	50	36	19.37	-25.12	-1.37	4.94	6.69
19	3	4	8	51	36	19.37	-25.25	-2	4.94	6.69
19	3	4	8	52	36	19.37	-25	-1.06	4.94	6.63
19	3	4	8	53	37	19.44	-25.19	-1.69	4.94	6.63
19	3	4	8	54	37	19.44	-25	-0.94	4.94	6.63
19	3	4	8	55	37	19.44	-25.12	-1.37	4.94	6.69
19	3	4	8	56	37	19.44	-24.94	-0.88	4.94	6.69
19	3	4	8	57	38	19.44	-25.06	-1.12	4.94	6.69
19	3	4	8	58	38	19.44	-25.25	-1.69	4.94	6.69
19	3	4	8	59	38	19.37	-25	-0.88	4.94	6.69
19	3	4	9	0	38	19.37	-25.19	-1.37	4.94	6.69
19	3	4	9	1	39	19.37	-24.94	-0.69	4.94	6.69
19	3	4	9	2	39	19.37	-25.12	-1.12	4.94	6.69
19	3	4	9	3	39	19.37	-24.94	-0.69	4.94	6.69
19	3	4	9	4	39	19.37	-25.06	-0.94	4.94	6.69
19	3	4	9	5	40	19.37	-25.19	-1.56	4.94	6.69
19	3	4	9	6	40	19.37	-25	-0.75	4.94	6.69
19	3	4	9	7	40	19.37	-25.12	-1.19	4.94	6.69
19	3	4	9	8	41	19.37	-24.94	-0.5	4.94	6.69
19	3	4	9	9	41	19.44	-25.12	-0.94	4.94	6.69
19	3	4	9	10	41	19.44	-24.94	-0.75	4.94	6.69
19	3	4	9	11	41	19.5	-25	-0.69	4.94	6.69
19	3	4	9	12	42	19.5	-25.19	-1.25	4.94	6.75
19	3	4	9	13	42	19.5	-24.94	-0.44	4.94	6.69
19	3	4	9	14	42	19.5	-25.12	-0.94	4.94	6.75
19	3	4	9	15	42	19.5	-24.94	-0.31	4.94	6.75
19	3	4	9	16	43	19.5	-25.06	-0.75	4.94	6.75
19	3	4	9	17	43	19.56	-25.06	-1.06	4.94	6.75
19	3	4	9	18	43	19.56	-25	-0.44	4.94	6.75
19	3	4	9	19	43	19.56	-25.12	-0.94	4.94	6.75
19	3	4	9	20	44	19.56	-24.94	-0.25	4.94	6.75
19	3	4	9	21	44	19.56	-25.12	-0.63	4.94	6.75
19	3	4	9	22	44	19.62	-24.94	-0.06	4.94	6.75
19	3	4	9	23	44	19.62	-25.06	-0.44	4.94	6.75
19	3	4	9	24	45	19.62	-25.19	-1.06	4.94	6.81
19	3	4	9	25	45	19.62	-25.06	-0.19	4.94	6.75
19	3	4	9	26	45	19.69	-25.12	-0.5	4.94	6.81
19	3	4	9	27	45	19.69	-24.94	0.06	4.94	6.81
19	3	4	9	28	46	19.75	-25.12	-0.25	4.94	6.81
19	3	4	9	29	46	19.75	-24.94	0.25	4.94	6.81
19	3	4	9	30	46	19.75	-25.06	-0.06	4.94	6.81
19	3	4	9	31	46	19.75	-25.25	-0.5	5	6.88
19	3	4	9	32	47	19.75	-25	0.19	5	6.88
19	3	4	9	33	47	19.75	-25.19	-0.25	4.94	6.88
19	3	4	9	34	47	19.69	-25	0.37	5	6.88
19	3	4	9	35	47	19.69	-25.12	0	5	6.88
19	3	4	9	36	48	19.69	-25.06	-0.19	5	6.88
19	3	4	9	37	48	19.62	-25.06	0.31	5	6.88
19	3	4	9	38	48	19.62	-25.19	-0.06	5	6.88
19	3	4	9	39	49	19.69	-25	0.5	5	6.88
19	3	4	9	40	49	19.69	-25.12	0.19	5	6.94
19	3	4	9	41	49	19.69	-24.94	0.69	5	6.94
19	3	4	9	42	49	19.69	-25.06	0.37	5	6.94
19	3	4	9	43	50	19.75	-25.19	0	5	6.94
19	3	4	9	44	50	19.75	-25	0.63	5	7
19	3	4	9	45	50	19.75	-25.12	0.25	5	6.94
19	3	4	9	46	50	19.81	-25	0.75	5	6.94
19	3	4	9	47	51	19.81	-25.12	0.5	5	7
19	3	4	9	48	51	19.81	-25	0.44	5	7
19	3	4	9	49	51	19.81	-25	0.88	5	7
19	3	4	9	50	51	19.81	-25.12	0.5	5	7
19	3	4	9	51	52	19.81	-25	1	5	7
19	3	4	9	52	52	19.81	-25.06	0.75	5	7.06
19	3	4	9	53	52	19.81	-24.94	1.19	5	7.06
19	3	4	9	54	52	19.81	-25.06	0.94	5	7.06
19	3	4	9	55	53	19.75	-25.19	0.63	5.06	7.06
19	3	4	9	56	53	19.75	-25	1.19	5	7.06
19	3	4	9	57	53	19.75	-25.12	0.88	5.06	7.06
19	3	4	9	58	53	19.75	-25	1.31	5.06	7.13
19	3	4	9	59	54	19.75	-25.12	1.06	5.06	7.13
19	3	4	10	0	54	19.75	-25.19	0.56	5.06	7.13
19	3	4	10	1	54	19.81	-25.06	1.19	5.06	7.13
19	3	4	10	2	54	19.81	-25.19	0.88	5.06	7.13
19	3	4	10	3	55	19.87	-25.06	1.37	5.06	7.13
19	3	4	10	4	55	19.87	-25.12	1.06	5.06	7.13
19	3	4	10	5	55	19.94	-25	1.5	5.06	7.13
19	3	4	10	6	56	19.94	-25.12	1.19	5.06	7.19
19	3	4	10	7	56	19.94	-25.25	0.63	5.06	7.13
19	3	4	10	8	56	19.94	-25.06	1.25	5.06	7.19
19	3	4	10	9	56	19.94	-25.19	0.81	5.06	7.13
19	3	4	10	10	57	19.94	-25	1.31	5.06	7.19
19	3	4	10	11	57	19.94	-25.12	1	5.06	7.19
19	3	4	10	12	57	19.94	-25.12	0.56	5.06	7.19
19	3	4	10	13	57	19.94	-25.06	1.25	5.06	7.19
19	3	4	10	14	58	19.94	-25.19	0.75	5.06	7.19
19	3	4	10	15	58	19.87	-25.06	1.31	5.06	7.19
19	3	4	10	16	58	19.87	-25	1.31	5.06	7.19
19	3	4	10	17	58	19.87	-24.87	1.75	5.06	7.19
19	3	4	10	18	59	19.87	-25.06	1.25	5.06	7.19
19	3	4	10	19	59	19.87	-25.19	0.56	5.06	7.19
19	3	4	10	20	59	19.87	-25.06	1.25	5.06	7.19
19	3	4	10	21	59	19.94	-25.25	0.69	5.06	7.19
19	3	4	10	23	0	19.94	-25.06	1.25	5.06	7.19
19	3	4	10	24	0	20	-25.19	0.75	5.06	7.19
19	3	4	10	25	0	20	-25	1.19	5.06	7.19
19	3	4	10	26	0	20	-25.12	0.94	5.06	7.19
19	3	4	10	27	1	20.06	-25.31	0.25	5.06	7.19
19	3	4	10	28	1	20.06	-25.06	0.94	5.06	7.19
19	3	4	10	29	1	20.12	-25.25	0.37	5.06	7.19
19	3	4	10	30	1	20.12	-25.06	1	5.06	7.19
19	3	4	10	31	2	20.19	-25.19	0.56	5.06	7.19
19	3	4	10	32	2	20.19	-25.25	0	5.06	7.19
19	3	4	10	33	2	20.25	-25.12	0.63	5.06	7.13
19	3	4	10	34	2	20.25	-25.31	0	5.06	7.13
19	3	4	10	35	3	20.25	-25.12	0.75	5.06	7.13
19	3	4	10	36	3	20.25	-25.25	0.25	5.06	7.13
19	3	4	10	37	3	20.25	-25	0.81	5.06	7.13
19	3	4	10	38	4	20.25	-25.12	0.31	5.06	7.13
19	3	4	10	39	4	20.19	-25.31	-0.56	5.06	7.13
19	3	4	10	40	4	20.19	-25.06	0.37	5.06	7.06
19	3	4	10	41	4	20.19	-25.25	-0.25	5.06	7.06
19	3	4	10	42	5	20.12	-25	0.44	5.06	7.06
19	3	4	10	43	5	20.12	-25.12	-0.06	5	7.06
19	3	4	10	44	5	20.12	-24.94	0.56	5.06	7.06
19	3	4	10	45	5	20.06	-25.06	0.06	5.06	7.06
19	3	4	10	46	6	20.06	-25.19	-0.81	5.06	7.06
19	3	4	10	47	6	20.06	-25	0.19	5	7.06
19	3	4	10	48	6	20	-25.19	-0.5	5	7.06
19	3	4	10	49	6	20	-24.94	0.25	5	7.06
19	3	4	10	50	7	20	-25.06	-0.37	5	7
19	3	4	10	51	7	20	-24.87	0.31	5	7
19	3	4	10	52	7	20	-25	-0.19	5	7
19	3	4	10	53	7	20.06	-25.19	-1.12	5	7
19	3	4	10	54	8	20.06	-24.87	-0.06	5	7
19	3	4	10	55	8	20.12	-25.06	-0.75	5	7
19	3	4	10	56	8	20.12	-24.81	0	5	7
19	3	4	10	57	8	20.19	-25	-0.63	5	7
19	3	4	10	58	9	20.25	-24.75	0	5	6.94
19	3	4	10	59	9	20.25	-24.94	-0.56	5	6.94
19	3	4	11	0	9	20.31	-25.06	-1.5	5	6.94
19	3	4	11	1	9	20.37	-24.81	-0.44	5	6.94
19	3	4	11	2	10	20.37	-25.06	-1.19	5	6.94
19	3	4	11	3	10	20.37	-24.81	-0.37	5	6.94
19	3	4	11	4	10	20.37	-24.94	-1	5	6.94
19	3	4	11	5	10	20.37	-24.69	-0.25	5	6.88
19	3	4	11	6	11	20.31	-24.81	-0.88	5	6.88
19	3	4	11	7	11	20.31	-24.87	-1.31	5	6.88
19	3	4	11	8	11	20.31	-24.75	-0.94	5	6.88
19	3	4	11	9	12	20.25	-24.94	-1.81	5	6.88
19	3	4	11	10	12	20.25	-24.69	-0.81	5	6.88
19	3	4	11	11	12	20.25	-24.81	-1.5	5	6.88
19	3	4	11	12	12	20.19	-24.56	-0.69	5	6.81
19	3	4	11	13	13	20.25	-24.69	-1.25	5	6.81
19	3	4	11	14	13	20.25	-24.5	-0.56	4.94	6.81
19	3	4	11	15	13	20.25	-24.62	-0.94	4.94	6.81
19	3	4	11	16	13	20.25	-24.81	-1.94	4.94	6.81
19	3	4	11	17	14	20.25	-24.56	-0.88	4.94	6.81
19	3	4	11	18	14	20.31	-24.75	-1.63	4.94	6.81
19	3	4	11	19	14	20.31	-24.5	-0.75	4.94	6.81
19	3	4	11	20	14	20.31	-24.69	-1.44	4.94	6.81
19	3	4	11	21	15	20.37	-24.44	-0.69	4.94	6.81
19	3	4	11	22	15	20.37	-24.62	-1.25	4.94	6.81
19	3	4	11	23	15	20.37	-24.81	-2.19	4.94	6.81
19	3	4	11	24	15	20.37	-24.5	-1.06	4.94	6.75
19	3	4	11	25	16	20.37	-24.75	-1.81	4.94	6.75
19	3	4	11	26	16	20.31	-24.5	-0.94	4.94	6.75
19	3	4	11	27	16	20.31	-24.62	-1.5	4.94	6.75
19	3	4	11	28	16	20.25	-24.44	-0.75	4.94	6.75
19	3	4	11	29	17	20.25	-24.56	-1.31	4.94	6.75
19	3	4	11	30	17	20.25	-24.62	-1.88	4.94	6.75
19	3	4	11	31	17	20.19	-24.44	-1	4.94	6.75
19	3	4	11	32	17	20.19	-24.69	-1.75	4.94	6.75
19	3	4	11	33	18	20.19	-24.44	-0.88	4.94	6.75
19	3	4	11	34	18	20.19	-24.56	-1.5	4.94	6.75
19	3	4	11	35	18	20.19	-24.37	-0.75	4.94	6.75
19	3	4	11	36	22	20.19	-24.5	-1.25	4.94	6.75
19	3	4	11	37	22	20.25	-24.62	-1.94	4.94	6.75
19	3	4	11	38	23	20.25	-24.44	-0.88	4.94	6.75
19	3	4	11	39	23	20.25	-24.56	-1.5	4.94	6.75
19	3	4	11	40	23	20.25	-24.31	-0.69	4.94	6.75
19	3	4	11	41	23	20.25	-24.5	-1.19	4.94	6.75
19	3	4	11	42	24	20.19	-24.31	-0.5	4.94	6.75
19	3	4	11	43	24	20.19	-24.44	-0.94	4.94	6.75
19	3	4	11	44	24	20.19	-24.56	-1.56	4.94	6.75
19	3	4	11	45	24	20.12	-24.37	-0.56	4.94	6.75
19	3	4	11	46	25	20.12	-24.5	-1.12	4.94	6.81
19	3	4	11	47	25	20.12	-24.31	-0.37	4.94	6.81
19	3	4	11	48	25	20.12	-24.5	-0.88	4.94	6.81
19	3	4	11	49	25	20.19	-24.31	-0.25	4.94	6.81
19	3	4	11	50	26	20.19	-24.5	-0.63	4.94	6.81
19	3	4	11	51	26	20.19	-24.62	-1.25	4.94	6.81
19	3	4	11	52	26	20.25	-24.37	-0.37	4.94	6.81
19	3	4	11	53	26	20.31	-24.56	-0.88	4.94	6.88
19	3	4	11	54	27	20.31	-24.37	-0.19	4.94	6.88
19	3	4	11	55	27	20.37	-24.5	-0.63	4.94	6.88
19	3	4	11	56	27	20.44	-24.31	-0.06	4.94	6.88
19	3	4	11	57	27	20.44	-24.5	-0.37	4.94	6.88
19	3	4	11	58	28	20.5	-24.62	-0.88	4.94	6.88
19	3	4	11	59	28	20.5	-24.44	-0.06	4.94	6.88
19	3	4	12	0	28	20.5	-24.56	-0.56	5	6.94
19	3	4	12	1	28	20.5	-24.37	0.06	4.94	6.94
19	3	4	12	2	29	20.5	-24.5	-0.31	5	6.94
19	3	4	12	3	29	20.5	-24.31	0.25	5	6.94
19	3	4	12	4	29	20.5	-24.37	0.12	4.94	6.94
19	3	4	12	5	29	20.5	-24.56	-0.37	5	6.94
19	3	4	12	6	30	20.44	-24.37	0.25	5	6.94
19	3	4	12	7	30	20.44	-24.56	-0.12	5	7
19	3	4	12	8	30	20.44	-24.37	0.44	5	7
19	3	4	12	9	31	20.44	-24.5	0.12	5	7
19	3	4	12	10	31	20.37	-24.5	-0.12	5	7
19	3	4	12	11	31	20.37	-24.44	0.37	5	7
19	3	4	12	12	31	20.37	-24.62	-0.06	5	7
19	3	4	12	13	32	20.44	-24.44	0.56	5	7
19	3	4	12	14	32	20.37	-24.56	0.12	5	7
19	3	4	12	15	32	20.44	-24.37	0.69	5	7.06
19	3	4	12	16	32	20.44	-24.5	0.31	5	7.06
19	3	4	12	17	33	20.5	-24.62	-0.06	5	7.06
19	3	4	12	18	33	20.5	-24.44	0.63	5	7.06
19	3	4	12	19	33	20.56	-24.56	0.31	5	7.06
19	3	4	12	20	33	20.56	-24.37	0.81	5	7.06
19	3	4	12	21	34	20.56	-24.5	0.56	5	7.06
19	3	4	12	22	34	20.56	-24.31	0.94	5.06	7.13
19	3	4	12	23	34	20.56	-24.44	0.75	5	7.13
19	3	4	12	24	34	20.5	-24.56	0.44	5.06	7.13
19	3	4	12	25	35	20.5	-24.44	1	5.06	7.13
19	3	4	12	26	35	20.5	-24.56	0.69	5	7.13
19	3	4	12	27	35	20.5	-24.44	1.12	5.06	7.13
19	3	4	12	28	35	20.44	-24.56	0.81	5.06	7.13
19	3	4	12	29	36	20.44	-24.69	0.25	5.06	7.19
19	3	4	12	30	36	20.44	-24.56	0.94	5.06	7.19
19	3	4	12	31	36	20.37	-24.69	0.56	5.06	7.19
19	3	4	12	32	36	20.44	-24.5	1.06	5.06	7.19
19	3	4	12	33	37	20.44	-24.56	0.75	5.06	7.19
19	3	4	12	34	37	20.44	-24.44	1.19	5.06	7.19
19	3	4	12	35	37	20.44	-24.44	1.19	5.06	7.19
19	3	4	12	36	38	20.5	-24.44	1	5.06	7.19
19	3	4	12	37	38	20.5	-24.31	1.5	5.06	7.25
19	3	4	12	38	38	20.56	-24.56	1	5.06	7.25
19	3	4	12	39	38	20.56	-24.37	1.44	5.06	7.25
19	3	4	12	40	39	20.56	-24.56	1	5.06	7.25
19	3	4	12	41	39	20.62	-24.69	0.31	5.06	7.25
19	3	4	12	42	39	20.62	-24.5	1	5.06	7.25
19	3	4	12	43	39	20.62	-24.69	0.5	5.06	7.25
19	3	4	12	44	40	20.56	-24.5	1.12	5.06	7.25
19	3	4	12	45	40	20.56	-24.62	0.63	5.06	7.25
19	3	4	12	46	40	20.56	-24.44	1.12	5.06	7.25
19	3	4	12	47	40	20.56	-24.56	0.69	5.06	7.25
19	3	4	12	48	41	20.56	-24.75	-0.06	5.06	7.25
19	3	4	12	49	41	20.56	-24.56	0.75	5.06	7.25
19	3	4	12	50	41	20.56	-24.69	0.25	5.06	7.25
19	3	4	12	51	41	20.56	-24.5	0.81	5.06	7.25
19	3	4	12	52	42	20.62	-24.62	0.37	5.06	7.19
19	3	4	12	53	42	20.62	-24.37	0.94	5.06	7.19
19	3	4	12	54	42	20.62	-24.5	0.56	5.06	7.19
19	3	4	12	55	42	20.62	-24.62	-0.06	5.06	7.19
19	3	4	12	56	43	20.69	-24.44	0.69	5.06	7.19
19	3	4	12	57	43	20.69	-24.62	0.06	5.06	7.19
19	3	4	12	58	43	20.75	-24.37	0.75	5.06	7.19
19	3	4	12	59	43	20.81	-24.56	0.19	5.06	7.19
19	3	4	13	0	44	20.81	-24.31	0.75	5.06	7.19
19	3	4	13	1	44	20.87	-24.44	0.31	5.06	7.13
19	3	4	13	2	44	20.87	-24.62	-0.44	5.06	7.13
19	3	4	13	3	45	20.94	-24.37	0.37	5.06	7.13
19	3	4	13	4	45	20.94	-24.56	-0.25	5.06	7.13
19	3	4	13	5	45	21	-24.31	0.44	5.06	7.13
19	3	4	13	6	45	21	-24.44	-0.06	5.06	7.13
19	3	4	13	7	46	21	-24.25	0.25	5.06	7.13
19	3	4	13	8	46	21	-24.37	0.06	5.06	7.06
19	3	4	13	9	46	20.94	-24.56	-0.75	5	7.13
19	3	4	13	10	46	20.94	-24.31	0.12	5.06	7.06
19	3	4	13	11	47	20.94	-24.44	-0.56	5.06	7.06
19	3	4	13	12	47	20.94	-24.25	0.12	5.06	7.06
19	3	4	13	13	47	20.94	-24.37	-0.37	5.06	7.06
19	3	4	13	14	47	20.94	-24.12	0.06	5	7.06
19	3	4	13	15	48	20.87	-24.25	-0.19	5	7
19	3	4	13	16	48	20.87	-24.5	-1.06	5.06	7.06
19	3	4	13	17	48	20.87	-24.19	-0.12	5	7
19	3	4	13	18	48	20.87	-24.37	-0.81	5.06	7
19	3	4	13	19	49	20.87	-24.12	-0.06	5	7
19	3	4	13	20	49	20.87	-24.31	-0.63	5	7
19	3	4	13	21	49	20.87	-24.12	0	5	7
19	3	4	13	22	49	20.87	-24.19	-0.25	5	7
19	3	4	13	23	50	20.94	-24.44	-1.06	5	7
19	3	4	13	24	50	20.94	-24.12	-0.12	5	7
19	3	4	13	25	50	21	-24.31	-0.88	5	7
19	3	4	13	26	50	21	-24.06	-0.12	5	7
19	3	4	13	27	51	21.06	-24.25	-0.75	5	7
19	3	4	13	28	51	21.06	-24	-0.06	5	7
19	3	4	13	29	51	21.12	-24.19	-0.69	5	7
19	3	4	13	30	52	21.12	-24.37	-1.69	5	7
19	3	4	13	31	52	21.19	-24.06	-0.56	5	7
19	3	4	13	32	52	21.12	-24.25	-1.31	5	7
19	3	4	13	33	52	21.12	-24	-0.5	5	7
19	3	4	13	34	53	21.12	-24.19	-1.12	5	6.94
19	3	4	13	35	53	21.12	-23.94	-0.37	5	6.94
19	3	4	13	36	53	21.06	-24.06	-0.94	5	6.94
19	3	4	13	37	53	21.06	-24.12	-1.5	5	6.94
19	3	4	13	38	54	21.06	-24	-1.06	5	6.94
19	3	4	13	39	54	21.06	-24.19	-1.94	5	6.94
19	3	4	13	40	54	21	-23.87	-0.94	5	6.94
19	3	4	13	41	54	21	-24.06	-1.63	5	6.88
19	3	4	13	42	55	21	-23.81	-0.81	5	6.88
19	3	4	13	43	55	21	-24	-1.44	5	6.88
19	3	4	13	44	55	20.94	-23.75	-0.69	5	6.88
19	3	4	13	45	55	20.94	-23.87	-1.19	5	6.88
19	3	4	13	46	56	20.94	-24.12	-2.38	5	6.88
19	3	4	13	47	56	20.94	-23.81	-1.19	4.94	6.88
19	3	4	13	48	56	20.94	-24	-2.06	5	6.88
19	3	4	13	49	56	21	-23.75	-1.12	4.94	6.88
19	3	4	13	50	57	21	-23.94	-1.81	4.94	6.81
19	3	4	13	51	57	21	-23.69	-1	4.94	6.81
19	3	4	13	52	57	21.06	-23.81	-1.69	4.94	6.81
19	3	4	13	53	57	21.06	-23.69	-1.44	4.94	6.81
19	3	4	13	54	58	21.12	-23.75	-1.5	4.94	6.81
19	3	4	13	55	58	21.12	-24	-2.5	4.94	6.81
19	3	4	13	56	58	21.12	-23.69	-1.37	4.94	6.81
19	3	4	13	57	59	21.12	-23.87	-2.19	4.94	6.81
19	3	4	13	58	59	21.12	-23.62	-1.25	4.94	6.81
19	3	4	13	59	59	21.12	-23.81	-1.94	4.94	6.81
19	3	4	14	0	59	21.12	-23.56	-1.12	4.94	6.81
19	3	4	14	2	0	21.12	-23.69	-1.81	4.94	6.75
19	3	4	14	3	0	21.12	-23.81	-2.56	4.94	6.75
19	3	4	14	4	0	21.06	-23.62	-1.56	4.94	6.75
19	3	4	14	5	0	21.06	-23.81	-2.44	4.94	6.75
19	3	4	14	6	1	21.06	-23.56	-1.44	4.94	6.75
19	3	4	14	7	1	21.06	-23.75	-2.13	4.94	6.75
19	3	4	14	8	1	21.06	-23.5	-1.25	4.94	6.75
19	3	4	14	9	1	21.06	-23.62	-1.88	4.94	6.75
19	3	4	14	10	2	21.06	-23.44	-1.12	4.94	6.75
19	3	4	14	11	2	21.06	-23.56	-1.56	4.94	6.75
19	3	4	14	12	2	21.06	-23.75	-2.44	4.94	6.75
19	3	4	14	13	2	21.06	-23.5	-1.37	4.94	6.75
19	3	4	14	14	3	21.06	-23.69	-2.06	4.94	6.75
19	3	4	14	15	3	21.06	-23.44	-1.19	4.94	6.75
19	3	4	14	16	3	21.12	-23.56	-1.81	4.94	6.75
19	3	4	14	17	3	21.12	-23.37	-1.06	4.94	6.75
19	3	4	14	18	4	21.19	-23.5	-1.56	4.94	6.75
19	3	4	14	19	4	21.19	-23.69	-2.31	4.94	6.75
19	3	4	14	20	4	21.25	-23.44	-1.25	4.94	6.75
19	3	4	14	21	4	21.25	-23.62	-1.94	4.94	6.75
19	3	4	14	22	5	21.31	-23.37	-1.06	4.94	6.75
19	3	4	14	23	5	21.31	-23.62	-1.69	4.94	6.75
19	3	4	14	24	5	21.37	-23.37	-0.94	4.94	6.75
19	3	4	14	25	5	21.37	-23.5	-1.37	4.94	6.75
19	3	4	14	26	6	21.31	-23.5	-1.56	4.94	6.75
19	3	4	14	27	6	21.31	-23.44	-1.19	4.94	6.75
19	3	4	14	28	6	21.31	-23.62	-1.81	4.94	6.81
19	3	4	14	29	7	21.25	-23.37	-0.94	4.94	6.75
19	3	4	14	30	7	21.25	-23.56	-1.5	4.94	6.81
19	3	4	14	31	7	21.25	-23.37	-0.75	4.94	6.81
19	3	4	14	32	7	21.25	-23.5	-1.25	4.94	6.81
19	3	4	14	33	8	21.19	-23.31	-0.75	4.94	6.81
19	3	4	14	34	8	21.19	-23.44	-1	4.94	6.81
19	3	4	14	35	8	21.19	-23.62	-1.63	4.94	6.81
19	3	4	14	36	8	21.12	-23.37	-0.75	4.94	6.81
19	3	4	14	37	9	21.12	-23.56	-1.19	4.94	6.81
19	3	4	14	38	9	21.12	-23.31	-0.5	4.94	6.81
19	3	4	14	39	9	21.06	-23.5	-0.94	4.94	6.88
19	3	4	14	40	9	21.06	-23.31	-0.37	4.94	6.88
19	3	4	14	41	10	21.06	-23.44	-0.69	4.94	6.88
19	3	4	14	42	10	21.06	-23.62	-1.25	4.94	6.88
19	3	4	14	43	10	21.06	-23.37	-0.44	4.94	6.88
19	3	4	14	44	10	21.06	-23.56	-1	4.94	6.88
19	3	4	14	45	11	21.06	-23.37	-0.31	4.94	6.88
19	3	4	14	46	11	21.12	-23.5	-0.81	4.94	6.88
19	3	4	14	47	11	21.12	-23.31	-0.25	4.94	6.88
19	3	4	14	48	11	21.19	-23.44	-0.5	4.94	6.88
19	3	4	14	49	12	21.19	-23.62	-1	5	6.88
19	3	4	14	50	12	21.25	-23.37	-0.25	5	6.88
19	3	4	14	51	12	21.25	-23.56	-0.69	4.94	6.94
19	3	4	14	52	12	21.25	-23.37	-0.06	5	6.94
19	3	4	14	53	13	21.25	-23.5	-0.5	5	6.94
19	3	4	14	54	13	21.25	-23.37	-0.31	5	6.94
19	3	4	14	55	13	21.25	-23.44	-0.31	5	6.94
19	3	4	14	56	13	21.19	-23.56	-0.75	5	6.94
19	3	4	14	57	14	21.19	-23.37	-0.06	5	6.94
19	3	4	14	58	14	21.19	-23.5	-0.44	5	6.94
19	3	4	14	59	14	21.19	-23.31	0.12	5	6.94
19	3	4	15	0	15	21.12	-23.5	-0.19	5	7
19	3	4	15	1	15	21.12	-23.56	-0.69	5	7
19	3	4	15	2	15	21.12	-23.37	0.06	5	7
19	3	4	15	3	15	21.06	-23.56	-0.44	5	7
19	3	4	15	4	16	21.06	-23.37	0.19	5	7
19	3	4	15	5	16	21.06	-23.5	-0.25	5	7
19	3	4	15	6	16	21	-23.31	0.31	5	7.06
19	3	4	15	7	16	21	-23.37	0.37	5	7.06
19	3	4	15	8	17	21	-23.5	-0.12	5	7.06
19	3	4	15	9	17	21.06	-23.31	0.56	5	7.06
19	3	4	15	10	17	21.06	-23.5	0.19	5	7.06
19	3	4	15	11	17	21.12	-23.37	0.69	5.06	7.13
19	3	4	15	12	18	21.12	-23.5	0.31	5	7.13
19	3	4	15	13	18	21.12	-23.37	0.81	5.06	7.13
19	3	4	15	14	18	21.19	-23.37	0.81	5.06	7.13
19	3	4	15	15	18	21.19	-23.44	0.44	5.06	7.13
19	3	4	15	16	19	21.25	-23.37	1	5.06	7.19
19	3	4	15	17	19	21.25	-23.5	0.5	5.06	7.19
19	3	4	15	18	19	21.25	-23.37	1	5.06	7.19
19	3	4	15	19	19	21.31	-23.5	0.63	5.06	7.19
19	3	4	15	20	20	21.37	-23.62	0.06	5.06	7.25
19	3	4	15	21	20	21.37	-23.5	0.75	5.06	7.25
19	3	4	15	22	20	21.37	-23.62	0.25	5.06	7.25
19	3	4	15	23	20	21.44	-23.5	0.88	5.06	7.25
19	3	4	15	24	21	21.44	-23.62	0.5	5.06	7.25
19	3	4	15	25	21	21.5	-23.44	1	5.06	7.25
19	3	4	15	26	25	21.5	-23.56	0.69	5.06	7.25
19	3	4	15	27	25	21.56	-23.69	0.06	5.06	7.25
19	3	4	15	28	25	21.56	-23.5	0.75	5.06	7.25
19	3	4	15	29	26	21.56	-23.69	0.31	5.06	7.25
19	3	4	15	30	26	21.62	-23.44	0.88	5.06	7.25
19	3	4	15	31	26	21.62	-23.62	0.44	5.06	7.25
19	3	4	15	32	26	21.62	-23.44	0.63	5.06	7.25
19	3	4	15	33	27	21.69	-23.56	0.56	5.06	7.25
19	3	4	15	34	27	21.69	-23.69	0.06	5.06	7.25
19	3	4	15	35	27	21.69	-23.5	0.75	5.06	7.25
19	3	4	15	36	27	21.62	-23.56	0.25	5.06	7.25
19	3	4	15	37	28	21.62	-23.37	0.88	5.06	7.25
19	3	4	15	38	28	21.56	-23.5	0.31	5.06	7.25
19	3	4	15	39	28	21.56	-23.56	-0.12	5.06	7.25
19	3	4	15	40	28	21.56	-23.44	0.44	5.06	7.19
19	3	4	15	41	29	21.5	-23.62	-0.19	5.06	7.25
19	3	4	15	42	29	21.5	-23.37	0.56	5.06	7.19
19	3	4	15	43	29	21.5	-23.56	0.06	5.06	7.19
19	3	4	15	44	29	21.44	-23.31	0.63	5.06	7.19
19	3	4	15	45	30	21.37	-23.44	0.19	5.06	7.19
19	3	4	15	46	30	21.37	-23.44	-0.12	5.06	7.19
19	3	4	15	47	30	21.37	-23.37	0.56	5.06	7.19
19	3	4	15	48	31	21.31	-23.56	-0.19	5.06	7.19
19	3	4	15	49	31	21.31	-23.37	0.56	5.06	7.19
19	3	4	15	50	31	21.31	-23.5	-0.06	5.06	7.19
19	3	4	15	51	31	21.25	-23.31	0.56	5.06	7.19
19	3	4	15	52	32	21.25	-23.44	0.06	5.06	7.19
19	3	4	15	53	32	21.25	-23.5	-0.63	5.06	7.19
19	3	4	15	54	32	21.25	-23.31	0.12	5.06	7.19
19	3	4	15	55	32	21.19	-23.5	-0.56	5.06	7.19
19	3	4	15	56	33	21.19	-23.31	0.25	5.06	7.19
19	3	4	15	57	33	21.19	-23.44	-0.37	5.06	7.19
19	3	4	15	58	33	21.19	-23.19	0.25	5.06	7.13
19	3	4	15	59	33	21.19	-23.37	-0.25	5.06	7.13
19	3	4	16	0	34	21.19	-23.31	-0.56	5.06	7.13
19	3	4	16	1	34	21.19	-23.31	-0.31	5.06	7.13
19	3	4	16	2	34	21.19	-23.44	-1.06	5.06	7.13
19	3	4	16	3	34	21.19	-23.19	-0.19	5.06	7.13
19	3	4	16	4	35	21.19	-23.37	-0.75	5.06	7.13
19	3	4	16	5	35	21.19	-23.12	0	5.06	7.13
19	3	4	16	6	35	21.19	-23.31	-0.5	5.06	7.13
19	3	4	16	7	35	21.19	-23.06	0.06	5.06	7.06
19	3	4	16	8	36	21.12	-23.19	-0.37	5.06	7.13
19	3	4	16	9	36	21.12	-23.44	-1.31	5.06	7.13
19	3	4	16	10	36	21.12	-23.12	-0.31	5.06	7.06
19	3	4	16	11	36	21.06	-23.31	-1	5.06	7.06
19	3	4	16	12	37	21.06	-23.06	-0.25	5.06	7.06
19	3	4	16	13	37	21.06	-23.25	-0.81	5	7.06
19	3	4	16	14	37	21.06	-23	-0.12	5.06	7.06
19	3	4	16	15	37	21.06	-23.12	-0.69	5	7.06
19	3	4	16	16	38	21.06	-23.31	-1.75	5.06	7.06
19	3	4	16	17	38	21.06	-23.06	-0.63	5	7.06
19	3	4	16	18	38	21.06	-23.25	-1.37	5	7
19	3	4	16	19	39	21.06	-23	-0.56	5	7
19	3	4	16	20	39	21.06	-23.19	-1.25	5	7
19	3	4	16	21	39	21.06	-22.94	-0.5	5	7
19	3	4	16	22	39	21.06	-23.06	-1.06	5	7
19	3	4	16	23	40	21.06	-23	-1.12	5	7
19	3	4	16	24	40	21	-23	-0.88	5	6.94
19	3	4	16	25	40	21	-23.19	-1.75	5	6.94
19	3	4	16	26	40	21	-22.94	-0.75	5	6.94
19	3	4	16	27	41	20.94	-23.12	-1.63	5	6.94
19	3	4	16	28	41	20.94	-22.87	-0.81	5	6.94
19	3	4	16	29	41	20.94	-23.06	-1.5	5	6.94
19	3	4	16	30	41	20.87	-22.81	-0.75	5	6.94
19	3	4	16	31	42	20.87	-23	-1.37	5	6.94
19	3	4	16	32	42	20.81	-23.19	-2.38	5	6.94
19	3	4	16	33	42	20.81	-22.94	-1.19	5	6.88
19	3	4	16	34	42	20.81	-23.12	-2.13	5	6.88
19	3	4	16	35	43	20.81	-22.81	-1.12	5	6.88
19	3	4	16	36	43	20.87	-23	-1.88	5	6.88
19	3	4	16	37	43	20.87	-22.75	-1	5	6.88
19	3	4	16	38	43	20.87	-22.94	-1.75	5	6.88
19	3	4	16	39	44	20.87	-22.69	-1	5	6.88
19	3	4	16	40	44	20.81	-22.87	-1.63	4.94	6.81
19	3	4	16	41	44	20.81	-23.06	-2.75	4.94	6.88
19	3	4	16	42	44	20.81	-22.75	-1.5	4.94	6.81
19	3	4	16	43	45	20.81	-23	-2.38	4.94	6.81
19	3	4	16	44	45	20.81	-22.69	-1.37	4.94	6.81
19	3	4	16	45	45	20.75	-22.87	-2.06	4.94	6.81
19	3	4	16	46	46	20.75	-22.62	-1.19	4.94	6.81
19	3	4	16	47	46	20.69	-22.75	-1.88	4.94	6.81
19	3	4	16	48	46	20.69	-22.56	-1.44	4.94	6.81
19	3	4	16	49	46	20.69	-22.69	-1.69	4.94	6.75
19	3	4	16	50	47	20.69	-22.94	-2.63	4.94	6.81
19	3	4	16	51	47	20.69	-22.69	-1.5	4.94	6.75
19	3	4	16	52	47	20.69	-22.87	-2.25	4.94	6.75
19	3	4	16	53	47	20.62	-22.62	-1.37	4.94	6.75
19	3	4	16	54	48	20.69	-22.81	-2.06	4.94	6.75
19	3	4	16	55	48	20.69	-22.56	-1.25	4.94	6.75
19	3	4	16	56	48	20.69	-22.69	-1.88	4.94	6.75
19	3	4	16	57	48	20.69	-22.69	-2.06	4.94	6.75
19	3	4	16	58	49	20.75	-22.62	-1.69	4.94	6.75
19	3	4	16	59	49	20.75	-22.87	-2.44	4.94	6.75
19	3	4	17	0	49	20.81	-22.56	-1.44	4.94	6.75
19	3	4	17	1	49	20.81	-22.81	-2.19	4.94	6.75
19	3	4	17	2	50	20.81	-22.5	-1.31	4.94	6.75
19	3	4	17	3	50	20.81	-22.75	-1.94	4.94	6.75
19	3	4	17	4	50	20.81	-22.44	-1.19	4.94	6.75
19	3	4	17	5	50	20.81	-22.62	-1.75	4.94	6.75
19	3	4	17	6	51	20.81	-22.81	-2.63	4.94	6.75
19	3	4	17	7	51	20.75	-22.5	-1.44	4.94	6.75
19	3	4	17	8	51	20.75	-22.75	-2.19	4.94	6.75
19	3	4	17	9	51	20.75	-22.5	-1.31	4.94	6.75
19	3	4	17	10	52	20.69	-22.69	-1.94	4.94	6.75
19	3	4	17	11	52	20.69	-22.44	-1.12	4.94	6.75
19	3	4	17	12	52	20.69	-22.62	-1.69	4.94	6.75
19	3	4	17	13	52	20.62	-22.56	-1.75	4.94	6.75
19	3	4	17	14	53	20.62	-22.56	-1.5	4.94	6.75
19	3	4	17	15	53	20.56	-22.75	-2.25	4.94	6.81
19	3	4	17	16	53	20.56	-22.5	-1.25	4.94	6.81
19	3	4	17	17	54	20.56	-22.69	-1.88	4.94	6.75
19	3	4	17	18	54	20.56	-22.44	-1.06	4.94	6.75
19	3	4	17	19	54	20.56	-22.62	-1.56	4.94	6.81
19	3	4	17	20	54	20.5	-22.37	-0.88	4.94	6.75
19	3	4	17	21	55	20.5	-22.56	-1.12	4.94	6.81
19	3	4	17	22	55	20.56	-22.75	-1.94	4.94	6.81
19	3	4	17	23	55	20.56	-22.5	-0.94	4.94	6.81
19	3	4	17	24	55	20.62	-22.69	-1.63	4.94	6.81
19	3	4	17	25	56	20.62	-22.44	-0.88	4.94	6.81
19	3	4	17	26	56	20.69	-22.62	-1.44	4.94	6.81
19	3	4	17	27	56	20.75	-22.44	-0.75	4.94	6.81
19	3	4	17	28	56	20.75	-22.56	-1.19	4.94	6.81
19	3	4	17	29	57	20.81	-22.75	-1.94	4.94	6.81
19	3	4	17	30	57	20.81	-22.5	-0.94	4.94	6.81
19	3	4	17	31	57	20.87	-22.69	-1.56	4.94	6.81
19	3	4	17	32	57	20.94	-22.44	-0.75	4.94	6.81
19	3	4	17	33	58	20.94	-22.62	-1.31	4.94	6.88
19	3	4	17	34	58	21	-22.37	-0.63	4.94	6.81
19	3	4	17	35	58	21	-22.56	-1.12	4.94	6.81
19	3	4	17	36	58	21	-22.69	-1.75	5	6.81
19	3	4	17	37	59	21	-22.5	-0.75	4.94	6.81
19	3	4	17	38	59	20.94	-22.69	-1.31	4.94	6.81
19	3	4	17	39	59	20.94	-22.44	-0.56	5	6.81
19	3	4	17	40	59	20.94	-22.62	-1.06	5	6.88
19	3	4	17	42	0	20.87	-22.37	-0.37	5	6.88
19	3	4	17	43	0	20.87	-22.56	-0.81	5	6.88
19	3	4	17	44	0	20.87	-22.69	-1.5	5	6.88
19	3	4	17	45	0	20.81	-22.5	-0.56	5	6.88
19	3	4	17	46	1	20.81	-22.69	-1.12	5	6.88
19	3	4	17	47	1	20.75	-22.44	-0.37	5	6.88
19	3	4	17	48	1	20.75	-22.56	-0.69	5	6.88
19	3	4	17	49	2	20.75	-22.37	-0.12	5	6.94
19	3	4	17	50	2	20.69	-22.5	-0.5	5	6.94
19	3	4	17	51	2	20.69	-22.62	-1.06	5	6.94
19	3	4	17	52	2	20.62	-22.44	-0.19	5	6.94
19	3	4	17	53	3	20.62	-22.62	-0.69	5	6.94
19	3	4	17	54	3	20.62	-22.37	0	5	6.94
19	3	4	17	55	3	20.56	-22.56	-0.31	5	6.94
19	3	4	17	56	3	20.56	-22.37	0.19	5	7
19	3	4	17	57	4	20.56	-22.5	-0.19	5	7
19	3	4	17	58	4	20.5	-22.62	-0.63	5	7
19	3	4	17	59	4	20.5	-22.44	0.12	5	7
19	3	4	18	0	4	20.5	-22.62	-0.25	5	7
19	3	4	18	1	5	20.5	-22.44	0.31	5	7
19	3	4	18	2	5	20.5	-22.56	-0.06	5	7.06
19	3	4	18	3	5	20.5	-22.37	0.5	5.06	7.06
19	3	4	18	4	5	20.5	-22.5	0.31	5.06	7.06
19	3	4	18	5	6	20.56	-22.62	-0.12	5.06	7.06
19	3	4	18	6	6	20.56	-22.44	0.5	5.06	7.06
19	3	4	18	7	6	20.56	-22.69	0.12	5.06	7.06
19	3	4	18	8	6	20.56	-22.5	0.63	5.06	7.13
19	3	4	18	9	7	20.56	-22.62	0.25	5.06	7.13
19	3	4	18	10	7	20.56	-22.5	0.37	5.06	7.13
19	3	4	18	11	7	20.5	-22.62	0.37	5.06	7.13
19	3	4	18	12	7	20.5	-22.75	-0.19	5.06	7.13
19	3	4	18	13	8	20.5	-22.56	0.5	5.06	7.13
19	3	4	18	14	8	20.5	-22.69	0.12	5.06	7.13
19	3	4	18	15	8	20.5	-22.5	0.63	5.06	7.13
19	3	4	18	16	8	20.44	-22.62	0.31	5.06	7.13
19	3	4	18	17	9	20.44	-22.69	-0.31	5.06	7.13
19	3	4	18	18	9	20.44	-22.56	0.37	5.06	7.13
19	3	4	18	19	9	20.37	-22.75	-0.25	5.06	7.13
19	3	4	18	20	10	20.37	-22.5	0.44	5.06	7.13
19	3	4	18	21	10	20.44	-22.69	0	5.06	7.13
19	3	4	18	22	10	20.44	-22.5	0.56	5.06	7.13
19	3	4	18	23	10	20.44	-22.62	0.19	5.06	7.13
19	3	4	18	24	11	20.44	-22.75	-0.56	5.06	7.13
19	3	4	18	25	11	20.44	-22.5	0.31	5.06	7.13
19	3	4	18	26	11	20.5	-22.69	-0.25	5.06	7.13
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	3	4	18	34	25	20.56	-21.37	2.44	5.13	7.31
19	3	4	18	35	25	20.56	-21.31	2.63	5.13	7.38
19	3	4	18	36	26	20.56	-21.25	2.69	5.13	7.44
19	3	4	18	37	26	20.56	-21.25	2.81	5.19	7.44
19	3	4	18	38	26	20.56	-21.31	2.94	5.19	7.5
19	3	4	18	39	26	20.5	-21.37	3.06	5.19	7.5
19	3	4	18	40	27	20.5	-21.44	3.13	5.19	7.56
19	3	4	18	41	27	20.5	-21.37	3.19	5.19	7.56
19	3	4	18	42	27	20.5	-21.37	3.31	5.19	7.63
19	3	4	18	43	27	20.5	-21.37	3.38	5.19	7.63
19	3	4	18	44	28	20.5	-21.37	3.5	5.25	7.69
19	3	4	18	45	28	20.5	-21.37	3.56	5.19	7.69
19	3	4	18	46	28	20.5	-21.37	3.56	5.25	7.69
19	3	4	18	47	28	20.56	-21.37	3.69	5.25	7.75
19	3	4	18	48	29	20.56	-21.37	3.75	5.25	7.75
19	3	4	18	49	29	20.56	-21.44	3.88	5.25	7.81
19	3	4	18	50	29	20.62	-21.5	3.88	5.25	7.81
19	3	4	18	51	29	20.69	-21.5	4	5.25	7.88
19	3	4	18	52	30	20.69	-21.37	4.06	5.25	7.88
19	3	4	18	53	30	20.75	-21.44	4.13	5.25	7.88
19	3	4	18	54	30	20.75	-21.44	4.19	5.31	7.94
19	3	4	18	55	30	20.81	-21.5	4.25	5.31	7.94
19	3	4	18	56	31	20.81	-21.44	4.31	5.31	7.94
19	3	4	18	57	31	20.81	-21.37	4.38	5.31	8
19	3	4	18	58	31	20.81	-21.37	4.38	5.31	8
19	3	4	18	59	32	20.81	-21.37	4.44	5.31	8
19	3	4	19	0	32	20.75	-21.44	4.5	5.31	8
19	3	4	19	1	32	20.75	-21.44	4.5	5.31	8
19	3	4	19	2	32	20.75	-21.44	4.56	5.31	8.06
19	3	4	19	3	33	20.75	-21.37	4.63	5.31	8.06
19	3	4	19	4	33	20.75	-21.44	4.69	5.31	8.06
19	3	4	19	5	33	20.75	-21.5	4.69	5.31	8.06
19	3	4	19	6	33	20.75	-21.5	4.75	5.31	8.06
19	3	4	19	7	34	20.69	-21.5	4.81	5.31	8.06
19	3	4	19	8	34	20.69	-21.5	4.81	5.31	8.13
19	3	4	19	9	34	20.69	-21.5	4.88	5.38	8.13
19	3	4	19	10	34	20.69	-21.5	4.94	5.31	8.13
19	3	4	19	11	35	20.62	-21.5	4.94	5.31	8.13
19	3	4	19	12	35	20.62	-21.5	5	5.38	8.13
19	3	4	19	13	35	20.62	-21.5	5	5.31	8.13
19	3	4	19	14	35	20.62	-21.5	5.06	5.38	8.19
19	3	4	19	15	36	20.56	-21.5	5.13	5.38	8.19
19	3	4	19	16	36	20.56	-21.44	5.19	5.38	8.19
19	3	4	19	17	36	20.62	-21.31	5.19	5.38	8.19
19	3	4	19	18	36	20.62	-21.44	5.06	5.38	8.19
19	3	4	19	19	37	20.62	-21.44	5.19	5.38	8.19
19	3	4	19	20	37	20.69	-21.56	5.06	5.38	8.19
19	3	4	19	21	37	20.69	-21.62	4.94	5.38	8.19
19	3	4	19	22	37	20.75	-21.69	5.13	5.38	8.19
19	3	4	19	23	38	20.75	-21.75	5	5.38	8.19
19	3	4	19	24	38	20.81	-21.75	5.13	5.38	8.25
19	3	4	19	25	38	20.81	-21.81	5.06	5.38	8.25
19	3	4	19	26	39	20.81	-21.87	4.81	5.38	8.25
19	3	4	19	27	39	20.81	-21.94	5.06	5.38	8.25
19	3	4	19	28	39	20.81	-22.06	4.81	5.44	8.25
19	3	4	19	29	39	20.81	-22	5	5.44	8.19
19	3	4	19	30	40	20.75	-22.12	4.88	5.44	8.19
19	3	4	19	31	40	20.75	-22.06	4.56	5.44	8.19
19	3	4	19	32	40	20.75	-22.12	4.88	5.44	8.19
19	3	4	19	33	40	20.75	-22.12	4.69	5.44	8.19
19	3	4	19	34	41	20.75	-22.06	4.88	5.38	8.19
19	3	4	19	35	41	20.75	-22.12	4.75	5.44	8.19
19	3	4	19	36	41	20.75	-22.12	4.56	5.44	8.19
19	3	4	19	37	41	20.69	-22.19	4.69	5.44	8.19
19	3	4	19	38	42	20.69	-22.25	4.38	5.44	8.19
19	3	4	19	39	42	20.69	-22.12	4.63	5.44	8.19
19	3	4	19	40	42	20.69	-22.25	4.44	5.44	8.19
19	3	4	19	41	42	20.69	-22.19	4.63	5.44	8.19
19	3	4	19	42	43	20.69	-22.25	4.44	5.44	8.13
19	3	4	19	43	43	20.62	-22.31	4.13	5.44	8.13
19	3	4	19	44	43	20.62	-22.19	4.44	5.44	8.13
19	3	4	19	45	43	20.62	-22.25	4.19	5.44	8.13
19	3	4	19	46	44	20.62	-22.19	4.44	5.44	8.13
19	3	4	19	47	44	20.62	-22.31	4.13	5.44	8.13
19	3	4	19	48	44	20.62	-22.37	3.69	5.44	8.13
19	3	4	19	49	44	20.69	-22.31	4.13	5.44	8.13
19	3	4	19	50	45	20.69	-22.37	3.75	5.44	8.06
19	3	4	19	51	45	20.75	-22.31	4.06	5.44	8.06
19	3	4	19	52	45	20.75	-22.31	3.81	5.44	8.06
19	3	4	19	53	45	20.81	-22.37	3.38	5.44	8.06
19	3	4	19	54	46	20.81	-22.31	3.81	5.44	8.06
19	3	4	19	55	46	20.87	-22.44	3.5	5.44	8.06
19	3	4	19	56	46	20.87	-22.31	3.88	5.44	8
19	3	4	19	57	47	20.94	-22.44	3.5	5.44	8.06
19	3	4	19	58	47	20.94	-22.31	3.81	5.44	8
19	3	4	19	59	47	21	-22.37	3.44	5.44	8
19	3	4	20	0	47	21	-22.5	3	5.44	8
19	3	4	20	1	48	21.06	-22.37	3.5	5.44	8
19	3	4	20	2	48	21.06	-22.44	3.06	5.44	8
19	3	4	20	3	48	21.06	-22.31	3.5	5.44	8
19	3	4	20	4	48	21.06	-22.44	3.19	5.44	7.94
19	3	4	20	5	49	21.06	-22.5	2.63	5.44	7.94
19	3	4	20	6	49	21.06	-22.31	3.25	5.44	7.94
19	3	4	20	7	49	21	-22.44	2.75	5.44	7.94
19	3	4	20	8	49	21	-22.31	3.19	5.44	7.94
19	3	4	20	9	50	21	-22.37	2.81	5.44	7.88
19	3	4	20	10	50	21	-22.25	3.19	5.44	7.88
19	3	4	20	11	50	21	-22.37	2.81	5.44	7.88
19	3	4	20	12	50	20.94	-22.5	2.31	5.44	7.88
19	3	4	20	13	51	20.94	-22.31	2.88	5.44	7.88
19	3	4	20	14	51	20.94	-22.44	2.5	5.44	7.88
19	3	4	20	15	51	20.94	-22.25	2.94	5.44	7.81
19	3	4	20	16	51	20.87	-22.31	2.56	5.44	7.81
19	3	4	20	17	52	20.87	-22.31	2	5.44	7.81
19	3	4	20	18	52	20.87	-22.25	2.56	5.44	7.81
19	3	4	20	19	52	20.87	-22.37	2	5.44	7.81
19	3	4	20	20	53	20.81	-22.19	2.56	5.44	7.81
19	3	4	20	21	53	20.81	-22.31	2.13	5.44	7.81
19	3	4	20	22	53	20.81	-22.12	2.63	5.44	7.75
19	3	4	20	23	53	20.75	-22.25	2.25	5.44	7.75
19	3	4	20	24	54	20.75	-22.37	1.56	5.44	7.75
19	3	4	20	25	54	20.75	-22.19	2.25	5.44	7.75
19	3	4	20	26	54	20.75	-22.31	1.75	5.44	7.75
19	3	4	20	27	54	20.75	-22.12	2.25	5.44	7.75
19	3	4	20	28	55	20.75	-22.25	1.75	5.44	7.75
19	3	4	20	29	55	20.69	-22.06	2.25	5.44	7.69
19	3	4	20	30	55	20.69	-22.12	2	5.44	7.69
19	3	4	20	31	55	20.69	-22.31	1.31	5.44	7.69
19	3	4	20	32	56	20.69	-22.06	2	5.44	7.69
19	3	4	20	33	56	20.69	-22.19	1.44	5.44	7.69
19	3	4	20	34	56	20.75	-22	2	5.44	7.69
19	3	4	20	35	56	20.81	-22.12	1.56	5.38	7.69
19	3	4	20	36	57	20.81	-21.94	1.75	5.44	7.63
19	3	4	20	37	57	20.87	-22.06	1.56	5.44	7.63
19	3	4	20	38	57	20.87	-22.19	0.88	5.44	7.63
19	3	4	20	39	57	20.87	-22	1.63	5.38	7.63
19	3	4	20	40	58	20.94	-22.12	1.06	5.38	7.63
19	3	4	20	41	58	20.94	-21.94	1.63	5.38	7.63
19	3	4	20	42	58	20.94	-22	1.19	5.44	7.63
19	3	4	20	43	58	20.94	-21.94	1.12	5.38	7.56
19	3	4	20	44	59	20.94	-22	1.12	5.38	7.56
19	3	4	20	45	59	20.87	-22.12	0.5	5.38	7.56
19	3	4	20	46	59	20.94	-21.94	1.25	5.38	7.56
19	3	4	20	48	0	20.87	-22.06	0.63	5.38	7.56
19	3	4	20	49	0	20.87	-21.81	1.31	5.38	7.56
19	3	4	20	50	0	20.87	-21.94	0.75	5.38	7.5
19	3	4	20	51	0	20.81	-21.81	0.81	5.38	7.5
19	3	4	20	52	1	20.81	-21.87	0.88	5.38	7.5
19	3	4	20	53	1	20.81	-22.06	0.19	5.38	7.5
19	3	4	20	54	1	20.81	-21.81	0.94	5.38	7.5
19	3	4	20	55	1	20.81	-21.94	0.44	5.38	7.5
19	3	4	20	56	2	20.75	-21.69	1.06	5.38	7.5
19	3	4	20	57	2	20.75	-21.81	0.56	5.38	7.5
19	3	4	20	58	2	20.75	-21.62	0.81	5.38	7.44
19	3	4	20	59	2	20.75	-21.75	0.69	5.38	7.44
19	3	4	21	0	3	20.75	-21.94	-0.12	5.38	7.44
19	3	4	21	1	3	20.75	-21.69	0.75	5.38	7.44
19	3	4	21	2	3	20.81	-21.81	0.12	5.38	7.44
19	3	4	21	3	3	20.81	-21.56	0.81	5.38	7.44
19	3	4	21	4	4	20.81	-21.75	0.25	5.38	7.44
19	3	4	21	5	4	20.87	-21.5	0.88	5.38	7.38
19	3	4	21	6	4	20.87	-21.62	0.44	5.38	7.38
19	3	4	21	7	4	20.94	-21.81	-0.44	5.38	7.38
19	3	4	21	8	5	20.87	-21.56	0.5	5.38	7.38
19	3	4	21	9	5	20.87	-21.75	-0.19	5.38	7.38
19	3	4	21	10	5	20.87	-21.5	0.56	5.38	7.38
19	3	4	21	11	5	20.87	-21.62	0	5.38	7.38
19	3	4	21	12	6	20.87	-21.37	0.63	5.38	7.31
19	3	4	21	13	6	20.87	-21.56	0.06	5.38	7.31
19	3	4	21	14	6	20.81	-21.75	-0.81	5.38	7.31
19	3	4	21	15	6	20.81	-21.44	0.25	5.38	7.31
19	3	4	21	16	7	20.81	-21.62	-0.5	5.38	7.31
19	3	4	21	17	7	20.81	-21.37	0.31	5.38	7.31
19	3	4	21	18	7	20.75	-21.56	-0.37	5.38	7.31
19	3	4	21	19	8	20.75	-21.31	0.37	5.38	7.31
19	3	4	21	20	8	20.75	-21.44	-0.25	5.38	7.31
19	3	4	21	21	8	20.75	-21.44	-0.37	5.38	7.31
19	3	4	21	22	8	20.69	-21.37	-0.06	5.38	7.31
19	3	4	21	23	9	20.69	-21.56	-0.94	5.38	7.31
19	3	4	21	24	9	20.69	-21.31	0	5.38	7.31
19	3	4	21	25	9	20.62	-21.44	-0.63	5.38	7.31
19	3	4	21	26	9	20.62	-21.19	0.12	5.38	7.31
19	3	4	21	27	10	20.62	-21.37	-0.56	5.38	7.25
19	3	4	21	28	10	20.62	-21.12	0.12	5.31	7.25
19	3	4	21	29	10	20.62	-21.31	-0.5	5.31	7.25
19	3	4	21	30	10	20.62	-21.5	-1.5	5.38	7.25
19	3	4	21	31	11	20.62	-21.19	-0.37	5.31	7.25
19	3	4	21	32	11	20.69	-21.37	-1.12	5.31	7.25
19	3	4	21	33	11	20.69	-21.12	-0.25	5.31	7.25
19	3	4	21	34	11	20.69	-21.25	-0.94	5.31	7.25
19	3	4	21	35	12	20.69	-21.06	-0.19	5.31	7.19
19	3	4	21	36	12	20.69	-21.19	-0.75	5.31	7.19
19	3	4	21	37	12	20.69	-21	-0.31	5.31	7.19
19	3	4	21	38	12	20.69	-21.06	-0.63	5.31	7.19
19	3	4	21	39	13	20.69	-21.31	-1.75	5.31	7.19
19	3	4	21	40	13	20.62	-21	-0.63	5.31	7.19
19	3	4	21	41	13	20.62	-21.25	-1.44	5.31	7.19
19	3	4	21	42	13	20.62	-20.94	-0.5	5.31	7.13
19	3	4	21	43	14	20.56	-21.12	-1.25	5.31	7.19
19	3	4	21	44	14	20.56	-20.87	-0.44	5.31	7.19
19	3	4	21	45	14	20.56	-21	-1	5.31	7.13
19	3	4	21	46	14	20.5	-20.94	-0.94	5.31	7.13
19	3	4	21	47	15	20.5	-20.94	-0.88	5.31	7.13
19	3	4	21	48	15	20.5	-21.12	-1.94	5.31	7.19
19	3	4	21	49	15	20.5	-20.87	-0.81	5.31	7.13
19	3	4	21	50	16	20.5	-21.06	-1.69	5.31	7.13
19	3	4	21	51	16	20.5	-20.81	-0.75	5.31	7.13
19	3	4	21	52	16	20.5	-20.94	-1.5	5.31	7.06
19	3	4	21	53	16	20.56	-20.75	-0.69	5.31	7.06
19	3	4	21	54	17	20.62	-20.87	-1.31	5.25	7.06
19	3	4	21	55	17	20.62	-20.75	-1.19	5.25	7.06
19	3	4	21	56	17	20.62	-20.75	-1.19	5.31	7.06
19	3	4	21	57	17	20.62	-20.94	-2.31	5.25	7.06
19	3	4	21	58	18	20.69	-20.69	-1.12	5.25	7.06
19	3	4	21	59	18	20.69	-20.87	-2.06	5.25	7.06
19	3	4	22	0	18	20.69	-20.62	-1.06	5.25	7
19	3	4	22	1	18	20.69	-20.81	-1.88	5.25	7
19	3	4	22	2	19	20.69	-20.56	-1	5.25	7
19	3	4	22	3	19	20.69	-20.75	-1.69	5.25	7
19	3	4	22	4	19	20.69	-20.5	-1	5.25	7
19	3	4	22	5	19	20.69	-20.62	-1.5	5.25	7
19	3	4	22	6	20	20.62	-20.87	-2.63	5.25	7
19	3	4	22	7	20	20.62	-20.56	-1.37	5.25	6.94
19	3	4	22	8	20	20.62	-20.81	-2.31	5.25	7
19	3	4	22	9	20	20.56	-20.5	-1.25	5.25	6.94
19	3	4	22	10	21	20.56	-20.69	-2.13	5.25	6.94
19	3	4	22	11	21	20.56	-20.44	-1.19	5.25	6.94
19	3	4	22	12	21	20.56	-20.62	-1.94	5.25	6.94
19	3	4	22	13	21	20.56	-20.44	-1.06	5.25	6.94
19	3	4	22	14	22	20.62	-20.56	-1.75	5.25	6.94
19	3	4	22	15	22	20.62	-20.69	-2.38	5.25	6.94
19	3	4	22	16	24	20.62	-20.5	-1.69	5.19	6.94
19	3	4	22	17	25	20.69	-20.75	-2.75	5.25	6.94
19	3	4	22	18	25	20.69	-20.44	-1.56	5.19	6.88
19	3	4	22	19	25	20.69	-20.62	-2.5	5.25	6.88
19	3	4	22	20	25	20.75	-20.37	-1.5	5.19	6.88
19	3	4	22	21	26	20.69	-20.5	-2.38	5.19	6.88
19	3	4	22	22	26	20.69	-20.31	-1.44	5.19	6.88
19	3	4	22	23	26	20.69	-20.44	-2.13	5.19	6.88
19	3	4	22	24	26	20.69	-20.19	-1.44	5.19	6.88
19	3	4	22	25	27	20.62	-20.37	-1.94	5.19	6.88
19	3	4	22	26	27	20.62	-20.56	-3.25	5.19	6.88
19	3	4	22	27	27	20.62	-20.31	-1.94	5.19	6.88
19	3	4	22	28	27	20.62	-20.5	-3	5.19	6.88
19	3	4	22	29	28	20.56	-20.25	-1.81	5.19	6.88
19	3	4	22	30	28	20.56	-20.44	-2.75	5.19	6.88
19	3	4	22	31	28	20.5	-20.19	-1.75	5.19	6.81
19	3	4	22	32	28	20.5	-20.31	-2.56	5.19	6.88
19	3	4	22	33	29	20.5	-20.06	-1.63	5.19	6.81
19	3	4	22	34	29	20.44	-20.25	-2.38	5.19	6.81
19	3	4	22	35	29	20.5	-20.06	-1.63	5.19	6.81
19	3	4	22	36	29	20.44	-20.19	-2.25	5.13	6.81
19	3	4	22	37	30	20.44	-20.37	-3.38	5.19	6.81
19	3	4	22	38	30	20.5	-20.12	-2.13	5.13	6.81
19	3	4	22	39	30	20.5	-20.31	-3.31	5.13	6.81
19	3	4	22	40	30	20.5	-20.06	-2.06	5.19	6.75
19	3	4	22	41	31	20.56	-20.25	-3.06	5.13	6.75
19	3	4	22	42	31	20.56	-20	-2	5.13	6.75
19	3	4	22	43	31	20.56	-20.19	-2.88	5.13	6.75
19	3	4	22	44	32	20.56	-19.94	-1.88	5.13	6.75
19	3	4	22	45	32	20.56	-20.06	-2.75	5.13	6.75
19	3	4	22	46	32	20.5	-19.87	-1.81	5.13	6.75
19	3	4	22	47	32	20.5	-20	-2.5	5.13	6.75
19	3	4	22	48	33	20.5	-20.06	-2.94	5.13	6.75
19	3	4	22	49	33	20.44	-19.94	-2.44	5.13	6.69
19	3	4	22	50	33	20.44	-20.19	-3.75	5.13	6.69
19	3	4	22	51	33	20.44	-19.87	-2.38	5.13	6.69
19	3	4	22	52	34	20.37	-20.06	-3.44	5.13	6.69
19	3	4	22	53	34	20.37	-19.81	-2.25	5.13	6.69
19	3	4	22	54	34	20.31	-20	-3.19	5.13	6.69
19	3	4	22	55	34	20.31	-19.75	-2.13	5.13	6.69
19	3	4	22	56	35	20.31	-19.87	-3	5.13	6.69
19	3	4	22	57	35	20.25	-19.69	-2	5.13	6.63
19	3	4	22	58	35	20.25	-19.81	-2.88	5.06	6.63
19	3	4	22	59	35	20.25	-19.62	-2	5.13	6.63
19	3	4	23	0	36	20.25	-19.75	-2.75	5.13	6.63
19	3	4	23	1	36	20.25	-19.94	-3.44	5.06	6.63
19	3	4	23	2	36	20.31	-19.69	-2.63	5.06	6.63
19	3	4	23	3	36	20.31	-19.94	-3.94	5.06	6.63
19	3	4	23	4	37	20.31	-19.62	-2.5	5.06	6.56
19	3	4	23	5	37	20.37	-19.81	-3.63	5.06	6.63
19	3	4	23	6	37	20.37	-19.56	-2.44	5.06	6.56
19	3	4	23	7	37	20.37	-19.75	-3.5	5.06	6.56
19	3	4	23	8	38	20.37	-19.5	-2.38	5.06	6.56
19	3	4	23	9	38	20.37	-19.69	-3.38	5.06	6.56
19	3	4	23	10	38	20.37	-19.44	-2.31	5.06	6.56
19	3	4	23	11	39	20.37	-19.62	-3.25	5.06	6.56
19	3	4	23	12	39	20.31	-19.37	-2.25	5.06	6.56
19	3	4	23	13	39	20.31	-19.56	-3	5.06	6.56
19	3	4	23	14	39	20.31	-19.56	-3.06	5.06	6.56
19	3	4	23	15	40	20.25	-19.5	-2.94	5.06	6.5
19	3	4	23	16	40	20.25	-19.75	-4.31	5	6.5
19	3	4	23	17	40	20.25	-19.44	-2.88	5.06	6.5
19	3	4	23	18	40	20.25	-19.69	-4.19	5.06	6.5
19	3	4	23	19	41	20.25	-19.37	-2.81	5	6.5
19	3	4	23	20	41	20.25	-19.56	-4	5	6.5
19	3	4	23	21	41	20.25	-19.31	-2.75	5	6.44
19	3	4	23	22	41	20.25	-19.5	-3.81	5	6.44
19	3	4	23	23	42	20.19	-19.25	-2.69	5	6.44
19	3	4	23	24	42	20.19	-19.44	-3.63	5	6.44
19	3	4	23	25	42	20.19	-19.19	-2.63	5	6.44
19	3	4	23	26	42	20.19	-19.31	-3.5	5	6.44
19	3	4	23	27	43	20.25	-19.12	-2.5	5	6.44
19	3	4	23	28	43	20.25	-19.25	-3.31	5	6.44
19	3	4	23	29	43	20.25	-19.25	-3.19	5	6.44
19	3	4	23	30	43	20.31	-19.19	-3.19	5	6.38
19	3	4	23	31	44	20.31	-19.44	-4.31	5	6.44
19	3	4	23	32	44	20.31	-19.19	-3.25	5	6.38
19	3	4	23	33	44	20.31	-19.44	-4.63	5	6.38
19	3	4	23	34	44	20.37	-19.12	-3.13	5	6.38
19	3	4	23	35	45	20.37	-19.31	-4.38	5	6.38
19	3	4	23	36	45	20.37	-19.06	-3.06	5	6.38
19	3	4	23	37	45	20.44	-19.25	-4.25	5	6.38
19	3	4	23	38	45	20.37	-19	-3	5	6.31
19	3	4	23	39	46	20.44	-19.19	-4.06	4.94	6.31
19	3	4	23	40	46	20.37	-18.94	-2.94	4.94	6.31
19	3	4	23	41	46	20.37	-19.12	-3.94	4.94	6.31
19	3	4	23	42	47	20.37	-18.87	-2.88	4.94	6.31
19	3	4	23	43	47	20.31	-19	-3.81	4.94	6.31
19	3	4	23	44	47	20.31	-18.81	-2.81	4.94	6.31
19	3	4	23	45	47	20.31	-18.94	-3.75	4.94	6.31
19	3	4	23	46	48	20.25	-18.87	-3.31	4.94	6.31
19	3	4	23	47	48	20.25	-18.87	-3.56	4.94	6.31
19	3	4	23	48	48	20.19	-19.06	-4.38	4.94	6.25
19	3	4	23	49	48	20.19	-18.81	-3.44	4.94	6.25
19	3	4	23	50	49	20.19	-19.06	-4.88	4.94	6.25
19	3	4	23	51	49	20.12	-18.75	-3.38	4.94	6.25
19	3	4	23	52	49	20.12	-19	-4.56	4.94	6.25
19	3	4	23	53	49	20.12	-18.69	-3.25	4.94	6.25
19	3	4	23	54	50	20.12	-18.94	-4.38	4.88	6.25
19	3	4	23	55	50	20.12	-18.69	-3.19	4.88	6.25
19	3	4	23	56	50	20.19	-18.87	-4.13	4.94	6.25
19	3	4	23	57	50	20.19	-18.62	-3	4.94	6.25
19	3	4	23	58	51	20.19	-18.81	-3.94	4.88	6.25
19	3	4	23	59	51	20.19	-18.56	-2.94	4.88	6.25
19	3	5	0	0	51	20.19	-18.75	-3.75	4.88	6.25
19	3	5	0	1	51	20.12	-18.56	-2.81	4.88	6.25
19	3	5	0	2	52	20.12	-18.69	-3.56	4.88	6.25
19	3	5	0	3	52	20.12	-18.62	-3.31	4.94	6.25
19	3	5	0	4	52	20.06	-18.69	-3.44	4.88	6.19
19	3	5	0	5	52	20.06	-18.87	-4.31	4.94	6.19
19	3	5	0	6	53	20.06	-18.62	-3.31	4.88	6.19
19	3	5	0	7	53	20	-18.87	-4.56	4.88	6.19
19	3	5	0	8	53	20	-18.56	-3.13	4.88	6.19
19	3	5	0	9	54	20	-18.81	-4.25	4.88	6.19
19	3	5	0	10	54	19.94	-18.5	-3	4.88	6.19
19	3	5	0	11	54	19.94	-18.75	-4.06	4.88	6.19
19	3	5	0	12	54	19.94	-18.5	-2.94	4.88	6.19
19	3	5	0	13	55	19.94	-18.62	-3.81	4.88	6.19
19	3	5	0	14	55	20	-18.44	-2.81	4.88	6.19
19	3	5	0	15	55	20	-18.62	-3.63	4.88	6.19
19	3	5	0	16	55	20	-18.44	-2.69	4.88	6.19
19	3	5	0	17	56	20	-18.56	-3.44	4.88	6.19
19	3	5	0	18	56	20.06	-18.44	-3	4.88	6.19
19	3	5	0	19	56	20.06	-18.5	-3.25	4.88	6.19
19	3	5	0	20	56	20.12	-18.75	-4.19	4.88	6.19
19	3	5	0	21	57	20.06	-18.5	-3.19	4.88	6.13
19	3	5	0	22	57	20.06	-18.75	-4.38	4.88	6.19
19	3	5	0	23	57	20.06	-18.44	-3.06	4.88	6.13
19	3	5	0	24	57	20.06	-18.69	-4.13	4.88	6.13
19	3	5	0	25	58	20.06	-18.44	-2.94	4.88	6.13
19	3	5	0	26	58	20	-18.56	-3.31	4.88	6.13
19	3	5	0	27	58	20	-18.37	-2.44	4.88	6.13
19	3	5	0	28	58	20	-18.56	-3.25	4.88	6.19
19	3	5	0	29	59	20	-18.37	-2.38	4.88	6.19
19	3	5	0	30	59	20	-18.5	-3.19	4.88	6.19
19	3	5	0	31	59	20	-18.31	-2.38	4.88	6.19
19	3	5	0	32	59	20	-18.5	-3	4.88	6.19
19	3	5	0	34	0	20	-18.62	-3.69	4.88	6.19
19	3	5	0	35	0	20	-18.44	-2.88	4.88	6.19
19	3	5	0	36	0	20	-18.69	-3.94	4.88	6.19
19	3	5	0	37	0	20	-18.44	-2.69	4.88	6.19
19	3	5	0	38	1	20.06	-18.62	-3.63	4.88	6.19
19	3	5	0	39	1	20.06	-18.37	-2.56	4.88	6.19
19	3	5	0	40	1	20.12	-18.56	-3.38	4.88	6.19
19	3	5	0	41	2	20.12	-18.37	-2.38	4.88	6.19
19	3	5	0	42	2	20.12	-18.5	-3.19	4.88	6.19
19	3	5	0	43	2	20.19	-18.31	-2.25	4.88	6.19
19	3	5	0	44	2	20.19	-18.5	-3	4.88	6.19
19	3	5	0	45	3	20.25	-18.31	-2.44	4.88	6.19
19	3	5	0	46	3	20.25	-18.44	-2.81	4.88	6.19
19	3	5	0	47	3	20.25	-18.62	-3.81	4.88	6.19
19	3	5	0	48	3	20.25	-18.37	-2.63	4.88	6.19
19	3	5	0	49	4	20.25	-18.62	-3.63	4.88	6.19
19	3	5	0	50	4	20.25	-18.37	-2.44	4.88	6.19
19	3	5	0	51	4	20.25	-18.56	-3.31	4.88	6.25
19	3	5	0	52	4	20.19	-18.31	-2.31	4.88	6.25
19	3	5	0	53	5	20.19	-18.5	-3.13	4.88	6.25
19	3	5	0	54	5	20.19	-18.31	-2.19	4.88	6.25
19	3	5	0	55	5	20.12	-18.44	-2.94	4.88	6.25
19	3	5	0	56	5	20.12	-18.31	-2.06	4.88	6.25
19	3	5	0	57	6	20.06	-18.44	-2.63	4.88	6.25
19	3	5	0	58	6	20.06	-18.62	-3.25	4.88	6.25
19	3	5	0	59	6	20.06	-18.44	-2.5	4.88	6.25
19	3	5	1	0	6	20	-18.69	-3.5	4.88	6.25
19	3	5	1	1	7	20	-18.44	-2.31	4.88	6.25
19	3	5	1	2	7	19.94	-18.62	-3.19	4.88	6.25
19	3	5	1	3	7	19.94	-18.44	-2.19	4.88	6.25
19	3	5	1	4	7	19.87	-18.62	-3	4.88	6.25
19	3	5	1	5	8	19.87	-18.37	-2.06	4.88	6.25
19	3	5	1	6	8	19.87	-18.56	-2.75	4.88	6.25
19	3	5	1	7	8	19.87	-18.37	-1.94	4.88	6.25
19	3	5	1	8	8	19.94	-18.5	-2.5	4.88	6.25
19	3	5	1	9	9	19.94	-18.56	-3.13	4.88	6.25
19	3	5	1	10	9	19.94	-18.44	-2.38	4.88	6.25
19	3	5	1	11	9	20	-18.69	-3.25	4.88	6.25
19	3	5	1	12	10	20	-18.37	-2.13	4.88	6.25
19	3	5	1	13	10	20.06	-18.62	-2.94	4.88	6.25
19	3	5	1	14	10	20	-18.37	-2	4.88	6.25
19	3	5	1	15	10	20	-18.62	-2.69	4.88	6.31
19	3	5	1	16	11	20	-18.37	-1.81	4.88	6.25
19	3	5	1	17	11	20	-18.56	-2.44	4.88	6.31
19	3	5	1	18	11	19.94	-18.37	-1.69	4.88	6.31
19	3	5	1	19	11	19.94	-18.5	-2.19	4.88	6.31
19	3	5	1	20	12	19.94	-18.75	-3.13	4.88	6.31
19	3	5	1	21	12	19.87	-18.5	-1.94	4.88	6.31
19	3	5	1	22	12	19.87	-18.69	-2.81	4.88	6.31
19	3	5	1	23	12	19.87	-18.44	-1.81	4.88	6.31
19	3	5	1	24	13	19.81	-18.62	-2.56	4.88	6.31
19	3	5	1	25	13	19.81	-18.44	-1.69	4.88	6.38
19	3	5	1	26	13	19.81	-18.56	-2.31	4.88	6.31
19	3	5	1	27	13	19.81	-18.37	-1.5	4.88	6.38
19	3	5	1	28	14	19.81	-18.5	-2.13	4.88	6.38
19	3	5	1	29	14	19.81	-18.62	-2.56	4.88	6.38
19	3	5	1	30	14	19.81	-18.5	-1.94	4.88	6.38
19	3	5	1	31	14	19.87	-18.69	-2.75	4.88	6.38
19	3	5	1	32	15	19.87	-18.44	-1.69	4.88	6.38
19	3	5	1	33	15	19.94	-18.62	-2.44	4.88	6.38
19	3	5	1	34	15	19.94	-18.44	-1.56	4.88	6.38
19	3	5	1	35	15	20	-18.56	-2.25	4.88	6.38
19	3	5	1	36	16	20	-18.37	-1.44	4.94	6.38
19	3	5	1	37	16	20.06	-18.5	-1.56	4.88	6.38
19	3	5	1	38	16	20.06	-18.44	-1.63	4.94	6.38
19	3	5	1	39	16	20.06	-18.44	-1.5	4.88	6.38
19	3	5	1	40	17	20.06	-18.69	-2.31	4.94	6.44
19	3	5	1	41	17	20.06	-18.44	-1.31	4.94	6.44
19	3	5	1	42	17	20.06	-18.69	-2.13	4.94	6.44
19	3	5	1	43	18	20	-18.5	-1.25	4.94	6.44
19	3	5	1	44	18	20	-18.62	-1.94	4.94	6.44
19	3	5	1	45	18	20	-18.44	-1.19	4.94	6.44
19	3	5	1	46	18	19.94	-18.56	-1.75	4.94	6.44
19	3	5	1	47	19	19.94	-18.56	-1.69	4.94	6.44
19	3	5	1	48	19	19.94	-18.56	-1.5	4.94	6.44
19	3	5	1	49	19	19.87	-18.75	-2.25	4.94	6.44
19	3	5	1	50	19	19.87	-18.5	-1.25	4.94	6.44
19	3	5	1	51	20	19.81	-18.69	-1.94	4.94	6.44
19	3	5	1	52	20	19.81	-18.44	-1.12	4.94	6.44
19	3	5	1	53	20	19.81	-18.62	-1.69	4.94	6.44
19	3	5	1	54	20	19.81	-18.44	-0.94	4.94	6.44
19	3	5	1	55	21	19.81	-18.56	-1.44	4.94	6.44
19	3	5	1	56	21	19.81	-18.69	-2	4.94	6.44
19	3	5	1	57	21	19.81	-18.5	-1.37	4.94	6.44
19	3	5	1	58	21	19.81	-18.69	-2	4.94	6.5
19	3	5	1	59	22	19.81	-18.5	-1.12	4.94	6.5
19	3	5	2	0	22	19.81	-18.62	-1.69	4.94	6.5
19	3	5	2	1	22	19.81	-18.44	-0.94	4.94	6.5
19	3	5	2	2	22	19.75	-18.56	-1.37	4.94	6.5
19	3	5	2	3	23	19.75	-18.37	-0.75	4.94	6.5
19	3	5	2	4	23	19.75	-18.44	-1.06	4.94	6.5
19	3	5	2	5	23	19.69	-18.62	-1.88	4.94	6.5
19	3	5	2	6	27	19.69	-18.44	-0.81	4.94	6.5
19	3	5	2	7	27	19.69	-18.62	-1.44	4.94	6.5
19	3	5	2	8	27	19.62	-18.44	-0.69	4.94	6.5
19	3	5	2	9	28	19.62	-18.56	-1.19	4.94	6.56
19	3	5	2	10	28	19.62	-18.37	-0.5	4.94	6.56
19	3	5	2	11	28	19.62	-18.5	-1	4.94	6.56
19	3	5	2	12	29	19.62	-18.69	-1.69	4.94	6.56
19	3	5	2	13	29	19.62	-18.44	-0.75	5	6.56
19	3	5	2	14	29	19.62	-18.62	-1.37	5	6.56
19	3	5	2	15	29	19.62	-18.44	-0.63	5	6.56
19	3	5	2	16	30	19.62	-18.56	-1.12	5	6.56
19	3	5	2	17	30	19.69	-18.44	-0.44	5	6.56
19	3	5	2	18	30	19.69	-18.5	-0.88	5	6.56
19	3	5	2	19	30	19.69	-18.44	-0.69	5	6.63
19	3	5	2	20	31	19.69	-18.5	-0.63	5	6.56
19	3	5	2	21	31	19.69	-18.62	-1.19	5	6.63
19	3	5	2	22	31	19.69	-18.44	-0.44	5	6.56
19	3	5	2	23	31	19.69	-18.62	-0.94	5	6.63
19	3	5	2	24	32	19.62	-18.44	-0.31	5	6.63
19	3	5	2	25	32	19.62	-18.56	-0.75	5	6.63
19	3	5	2	26	32	19.62	-18.44	-0.12	5	6.63
19	3	5	2	27	32	19.56	-18.5	-0.31	5	6.63
19	3	5	2	28	33	19.56	-18.62	-0.94	5	6.63
19	3	5	2	29	33	19.5	-18.44	-0.12	5	6.69
19	3	5	2	30	33	19.5	-18.62	-0.69	5	6.69
19	3	5	2	31	33	19.5	-18.44	0	5	6.69
19	3	5	2	32	34	19.44	-18.56	-0.44	5	6.69
19	3	5	2	33	34	19.44	-18.44	0.06	5	6.69
19	3	5	2	34	34	19.5	-18.56	-0.31	5	6.69
19	3	5	2	35	34	19.5	-18.69	-0.94	5	6.69
19	3	5	2	36	35	19.56	-18.5	-0.12	5	6.69
19	3	5	2	37	35	19.56	-18.62	-0.5	5	6.69
19	3	5	2	38	35	19.62	-18.5	0.12	5	6.69
19	3	5	2	39	35	19.62	-18.62	-0.31	5	6.69
19	3	5	2	40	36	19.69	-18.44	0.25	5	6.75
19	3	5	2	41	36	19.75	-18.56	-0.12	5	6.69
19	3	5	2	42	36	19.75	-18.69	-0.69	5	6.75
19	3	5	2	43	36	19.75	-18.5	0.12	5	6.75
19	3	5	2	44	37	19.81	-18.62	-0.37	5.06	6.75
19	3	5	2	45	37	19.81	-18.5	0.25	5.06	6.75
19	3	5	2	46	37	19.87	-18.56	-0.19	5.06	6.75
19	3	5	2	47	38	19.87	-18.44	0.37	5.06	6.75
19	3	5	2	48	38	19.87	-18.5	0	5.06	6.75
19	3	5	2	49	38	19.87	-18.62	-0.69	5	6.75
19	3	5	2	50	38	19.94	-18.5	0.12	5.06	6.75
19	3	5	2	51	39	19.94	-18.62	-0.37	5.06	6.75
19	3	5	2	52	39	19.94	-18.44	0.31	5.06	6.75
19	3	5	2	53	39	20	-18.5	-0.12	5.06	6.75
19	3	5	2	54	39	20	-18.37	0.44	5.06	6.75
19	3	5	2	55	40	20	-18.5	0	5.06	6.75
19	3	5	2	56	40	20	-18.69	-0.75	5.06	6.75
19	3	5	2	57	40	20	-18.5	0.12	5.06	6.75
19	3	5	2	58	40	20	-18.62	-0.44	5.06	6.81
19	3	5	2	59	41	20.06	-18.44	0.25	5.06	6.81
19	3	5	3	0	41	20.06	-18.56	-0.19	5.06	6.75
19	3	5	3	1	41	20.06	-18.44	0.37	5.06	6.75
19	3	5	3	2	41	20.06	-18.5	0	5.06	6.75
19	3	5	3	3	41	20	-18.62	-0.75	5.06	6.75
19	3	5	3	4	41	20	-18.44	0.12	5.06	6.81
19	3	5	3	5	41	19.94	-18.62	-0.5	5.06	6.81
19	3	5	3	6	41	19.94	-18.44	0.25	5.06	6.81
19	3	5	3	7	42	19.87	-18.56	-0.31	5.06	6.81
19	3	5	3	8	42	19.87	-18.37	0.31	5.06	6.81
19	3	5	3	9	42	19.81	-18.5	-0.19	5.06	6.81
19	3	5	3	10	42	19.81	-18.44	-0.12	5.06	6.81
19	3	5	3	11	43	19.75	-18.44	0	5.06	6.75
19	3	5	3	12	43	19.75	-18.56	-0.75	5.06	6.81
19	3	5	3	13	43	19.69	-18.37	0.06	5.06	6.81
19	3	5	3	14	43	19.69	-18.5	-0.44	5.06	6.81
19	3	5	3	15	44	19.62	-18.31	0.19	5.06	6.75
19	3	5	3	16	44	19.62	-18.44	-0.25	5.06	6.75
19	3	5	3	17	44	19.56	-18.25	0.31	5.06	6.81
19	3	5	3	18	44	19.5	-18.37	0	5.06	6.81
19	3	5	3	19	45	19.5	-18.5	-0.88	5.06	6.81
19	3	5	3	20	45	19.5	-18.31	0.06	5.06	6.75
19	3	5	3	21	45	19.44	-18.44	-0.63	5.06	6.75
19	3	5	3	22	46	19.44	-18.25	0.06	5.06	6.75
19	3	5	3	23	46	19.37	-18.37	-0.5	5.06	6.75
19	3	5	3	24	46	19.37	-18.25	0.19	5.06	6.75
19	3	5	3	25	46	19.37	-18.37	-0.31	5.06	6.75
19	3	5	3	26	47	19.37	-18.37	-0.75	5.06	6.75
19	3	5	3	27	47	19.37	-18.31	-0.37	5.06	6.75
19	3	5	3	28	47	19.37	-18.5	-1.19	5.06	6.75
19	3	5	3	29	47	19.37	-18.19	-0.25	5.06	6.75
19	3	5	3	30	48	19.44	-18.37	-0.88	5.06	6.75
19	3	5	3	31	48	19.44	-18.12	-0.12	5.06	6.75
19	3	5	3	32	48	19.5	-18.25	-0.75	5.06	6.75
19	3	5	3	33	48	19.5	-18.12	-0.06	5.06	6.75
19	3	5	3	34	49	19.56	-18.19	-0.56	5.06	6.69
19	3	5	3	35	49	19.56	-18.37	-1.5	5.06	6.69
19	3	5	3	36	49	19.62	-18.19	-0.44	5.06	6.69
19	3	5	3	37	49	19.62	-18.31	-1.19	5.06	6.69
19	3	5	3	38	50	19.69	-18.06	-0.37	5.06	6.69
19	3	5	3	39	50	19.69	-18.25	-1	5.06	6.69
19	3	5	3	40	50	19.69	-18.06	-0.31	5	6.63
19	3	5	3	41	50	19.69	-18.12	-0.88	5.06	6.63
19	3	5	3	42	51	19.69	-17.94	-0.19	5	6.63
19	3	5	3	43	51	19.69	-18	-0.56	5	6.63
19	3	5	3	44	51	19.62	-18.25	-1.56	5.06	6.63
19	3	5	3	45	51	19.62	-18	-0.5	5.06	6.63
19	3	5	3	46	52	19.56	-18.19	-1.25	5	6.63
19	3	5	3	47	52	19.56	-17.94	-0.44	5	6.63
19	3	5	3	48	52	19.56	-18.06	-1.19	5	6.63
19	3	5	3	49	52	19.5	-17.87	-0.44	5	6.63
19	3	5	3	50	53	19.5	-17.94	-1.12	5	6.63
19	3	5	3	51	53	19.44	-17.81	-0.44	5.06	6.56
19	3	5	3	52	53	19.44	-17.87	-1	5	6.56
19	3	5	3	53	53	19.44	-18.06	-2	5	6.56
19	3	5	3	54	54	19.44	-17.81	-0.81	5	6.56
19	3	5	3	55	54	19.44	-18	-1.69	5	6.56
19	3	5	3	56	54	19.44	-17.75	-0.75	5	6.56
19	3	5	3	57	55	19.44	-17.87	-1.5	5	6.56
19	3	5	3	58	55	19.5	-17.69	-0.69	5	6.56
19	3	5	3	59	55	19.5	-17.81	-1.31	5	6.5
19	3	5	4	0	55	19.5	-17.62	-0.63	5	6.5
19	3	5	4	1	56	19.5	-17.69	-1.12	5	6.5
19	3	5	4	2	56	19.5	-17.87	-2.06	5	6.5
19	3	5	4	3	56	19.5	-17.62	-1	5	6.5
19	3	5	4	4	56	19.5	-17.81	-1.88	5	6.5
19	3	5	4	5	57	19.5	-17.56	-0.94	5	6.5
19	3	5	4	6	57	19.5	-17.75	-1.69	5	6.5
19	3	5	4	7	57	19.44	-17.5	-0.88	5	6.5
19	3	5	4	8	57	19.44	-17.62	-1.56	5	6.5
19	3	5	4	9	58	19.37	-17.5	-0.81	5	6.5
19	3	5	4	10	58	19.37	-17.56	-1.44	5	6.44
19	3	5	4	11	58	19.37	-17.56	-1.56	5	6.5
19	3	5	4	12	58	19.31	-17.5	-1.31	5	6.5
19	3	5	4	13	59	19.31	-17.69	-2.31	5	6.44
19	3	5	4	14	59	19.31	-17.44	-1.19	5	6.44
19	3	5	4	15	59	19.31	-17.56	-2	5	6.44
19	3	5	4	16	59	19.31	-17.37	-1.06	5	6.44
19	3	5	4	18	0	19.31	-17.5	-1.88	5	6.44
19	3	5	4	19	0	19.31	-17.31	-1	5	6.44
19	3	5	4	20	0	19.31	-17.37	-1.75	4.94	6.44
19	3	5	4	21	0	19.37	-17.25	-1	5	6.38
19	3	5	4	22	1	19.37	-17.31	-1.44	4.94	6.38
19	3	5	4	23	1	19.44	-17.44	-2.31	4.94	6.44
19	3	5	4	24	1	19.44	-17.25	-1.56	4.94	6.38
19	3	5	4	25	1	19.44	-17.37	-2.5	4.94	6.38
19	3	5	4	26	2	19.5	-17.19	-1.37	4.94	6.38
19	3	5	4	27	2	19.5	-17.31	-2.25	4.94	6.38
19	3	5	4	28	2	19.5	-17.12	-1.31	4.94	6.38
19	3	5	4	29	3	19.5	-17.25	-2.06	4.94	6.38
19	3	5	4	30	3	19.5	-17.06	-1.19	4.94	6.31
19	3	5	4	31	3	19.5	-17.12	-1.88	4.94	6.31
19	3	5	4	32	3	19.5	-17	-1.25	4.94	6.31
19	3	5	4	33	4	19.5	-17.06	-1.69	4.94	6.31
19	3	5	4	34	4	19.44	-17.31	-2.75	4.94	6.31
19	3	5	4	35	4	19.44	-17	-1.5	4.94	6.31
19	3	5	4	36	4	19.37	-17.19	-2.44	4.94	6.31
19	3	5	4	37	5	19.37	-17	-1.44	4.94	6.31
19	3	5	4	38	5	19.37	-17.12	-2.13	4.94	6.31
19	3	5	4	39	5	19.37	-16.94	-1.25	4.94	6.31
19	3	5	4	40	5	19.37	-17.06	-1.88	4.94	6.31
19	3	5	4	41	6	19.37	-16.87	-1.12	4.94	6.31
19	3	5	4	42	6	19.37	-17	-1.69	4.94	6.31
19	3	5	4	43	6	19.37	-17	-1.88	4.94	6.31
19	3	5	4	44	6	19.44	-16.94	-1.44	4.94	6.31
19	3	5	4	45	7	19.44	-17.06	-2.31	4.94	6.31
19	3	5	4	46	7	19.44	-16.87	-1.25	4.94	6.31
19	3	5	4	47	7	19.44	-17.06	-2	4.94	6.31
19	3	5	4	48	7	19.37	-16.87	-1.12	4.94	6.31
19	3	5	4	49	8	19.37	-17	-1.75	4.94	6.31
19	3	5	4	50	8	19.37	-16.81	-1	4.94	6.31
19	3	5	4	51	8	19.37	-16.94	-1.56	4.94	6.31
19	3	5	4	52	8	19.31	-16.81	-1.12	4.94	6.31
19	3	5	4	53	9	19.31	-16.87	-1.37	4.94	6.31
19	3	5	4	54	9	19.25	-17	-2.25	4.94	6.31
19	3	5	4	55	9	19.25	-16.75	-1.19	4.94	6.31
19	3	5	4	56	9	19.25	-16.94	-1.94	4.94	6.31
19	3	5	4	57	10	19.25	-16.75	-1.06	4.94	6.31
19	3	5	4	58	10	19.25	-16.87	-1.69	4.94	6.31
19	3	5	4	59	10	19.25	-16.75	-0.94	4.94	6.38
19	3	5	5	0	11	19.25	-16.87	-1.44	4.94	6.31
19	3	5	5	1	11	19.25	-16.69	-0.81	4.94	6.31
19	3	5	5	2	11	19.25	-16.81	-1.25	4.94	6.31
19	3	5	5	3	11	19.25	-16.94	-2.06	4.94	6.38
19	3	5	5	4	12	19.19	-16.75	-1	4.94	6.38
19	3	5	5	5	12	19.19	-16.87	-1.75	4.94	6.38
19	3	5	5	6	12	19.19	-16.69	-0.88	4.94	6.38
19	3	5	5	7	12	19.12	-16.81	-1.44	4.94	6.38
19	3	5	5	8	13	19.12	-16.69	-0.75	4.94	6.38
19	3	5	5	9	13	19.12	-16.75	-1.19	4.94	6.38
19	3	5	5	10	13	19.06	-16.62	-0.75	4.94	6.38
19	3	5	5	11	13	19.12	-16.69	-1	4.94	6.38
19	3	5	5	12	14	19.12	-16.87	-1.81	4.94	6.38
19	3	5	5	13	14	19.12	-16.62	-0.81	4.94	6.38
19	3	5	5	14	14	19.19	-16.75	-1.5	4.94	6.38
19	3	5	5	15	14	19.19	-16.62	-0.69	4.94	6.38
19	3	5	5	16	15	19.19	-16.75	-1.31	5	6.38
19	3	5	5	17	15	19.19	-16.56	-0.56	4.94	6.38
19	3	5	5	18	15	19.19	-16.69	-1	4.94	6.38
19	3	5	5	19	15	19.25	-16.62	-1.12	4.94	6.38
19	3	5	5	20	16	19.25	-16.62	-0.75	4.94	6.38
19	3	5	5	21	16	19.25	-16.75	-1.37	5	6.38
19	3	5	5	22	16	19.19	-16.56	-0.56	5	6.38
19	3	5	5	23	16	19.19	-16.69	-1.12	5	6.38
19	3	5	5	24	17	19.19	-16.5	-0.44	5	6.38
19	3	5	5	25	17	19.19	-16.62	-0.88	5	6.38
19	3	5	5	26	17	19.19	-16.5	-0.25	4.94	6.44
19	3	5	5	27	18	19.19	-16.62	-0.94	5	6.38
19	3	5	5	28	18	19.19	-16.69	-1.5	5	6.44
19	3	5	5	29	18	19.19	-16.56	-0.56	5	6.44
19	3	5	5	30	18	19.25	-16.69	-1.12	5	6.44
19	3	5	5	31	19	19.25	-16.5	-0.37	5	6.44
19	3	5	5	32	19	19.25	-16.62	-0.88	5	6.44
19	3	5	5	33	19	19.25	-16.44	-0.19	5	6.44
19	3	5	5	34	19	19.31	-16.62	-0.69	5	6.44
19	3	5	5	35	20	19.31	-16.5	-0.37	5	6.44
19	3	5	5	36	20	19.25	-16.56	-0.44	5	6.5
19	3	5	5	37	20	19.25	-16.62	-1	5	6.5
19	3	5	5	38	20	19.25	-16.5	-0.19	5	6.5
19	3	5	5	39	21	19.19	-16.62	-0.69	5	6.5
19	3	5	5	40	21	19.19	-16.5	-0.06	5	6.5
19	3	5	5	41	21	19.19	-16.62	-0.44	5	6.5
19	3	5	5	42	21	19.12	-16.44	0.12	5	6.5
19	3	5	5	43	22	19.12	-16.56	-0.19	5	6.56
19	3	5	5	44	22	19.12	-16.69	-0.88	5	6.56
19	3	5	5	45	22	19.12	-16.5	0	5	6.56
19	3	5	5	46	22	19.12	-16.62	-0.56	5	6.56
19	3	5	5	47	23	19.12	-16.5	0.06	5	6.56
19	3	5	5	48	23	19.12	-16.62	-0.31	5	6.56
19	3	5	5	49	23	19.12	-16.5	0.25	5	6.56
19	3	5	5	50	23	19.12	-16.56	-0.12	5	6.56
19	3	5	5	51	24	19.19	-16.62	-0.75	5	6.56
19	3	5	5	52	24	19.19	-16.5	0.12	5	6.56
19	3	5	5	53	24	19.19	-16.62	-0.44	5	6.56
19	3	5	5	54	24	19.19	-16.5	0.25	5	6.56
19	3	5	5	55	25	19.19	-16.62	-0.19	5.06	6.63
19	3	5	5	56	28	19.19	-16.5	0.44	5.06	6.63
19	3	5	5	57	29	19.19	-16.56	0.06	5.06	6.63
19	3	5	5	58	29	19.19	-16.62	-0.69	5.06	6.63
19	3	5	5	59	29	19.19	-16.5	0.19	5.06	6.63
19	3	5	6	0	29	19.12	-16.62	-0.31	5.06	6.63
19	3	5	6	1	30	19.12	-16.44	0.31	5.06	6.63
19	3	5	6	2	30	19.12	-16.56	-0.12	5.06	6.63
19	3	5	6	3	30	19.12	-16.44	0.44	5.06	6.63
19	3	5	6	4	31	19.12	-16.56	0.12	5.06	6.63
19	3	5	6	5	31	19.12	-16.62	-0.63	5.06	6.63
19	3	5	6	6	31	19.12	-16.5	0.25	5.06	6.63
19	3	5	6	7	31	19.12	-16.62	-0.44	5.06	6.63
19	3	5	6	8	32	19.12	-16.44	0.31	5.06	6.63
19	3	5	6	9	32	19.19	-16.56	-0.19	5.06	6.63
19	3	5	6	10	32	19.19	-16.44	0.37	5.06	6.63
19	3	5	6	11	32	19.25	-16.5	-0.06	5.06	6.63
19	3	5	6	12	33	19.25	-16.5	-0.06	5.06	6.63
19	3	5	6	13	33	19.31	-16.5	0.12	5.06	6.63
19	3	5	6	14	33	19.31	-16.62	-0.56	5.06	6.63
19	3	5	6	15	33	19.31	-16.44	0.19	5.06	6.63
19	3	5	6	16	34	19.37	-16.56	-0.31	5.06	6.63
19	3	5	6	17	34	19.37	-16.37	0.31	5.06	6.63
19	3	5	6	18	34	19.37	-16.44	-0.19	5.06	6.63
19	3	5	6	19	34	19.37	-16.31	0.37	5.06	6.63
19	3	5	6	20	35	19.44	-16.37	0.06	5.06	6.63
19	3	5	6	21	35	19.44	-16.56	-0.75	5.06	6.63
19	3	5	6	22	35	19.37	-16.37	0.12	5.06	6.63
19	3	5	6	23	35	19.37	-16.5	-0.44	5.06	6.63
19	3	5	6	24	36	19.37	-16.31	0.25	5.06	6.63
19	3	5	6	25	36	19.31	-16.44	-0.31	5.06	6.63
19	3	5	6	26	36	19.31	-16.25	0.31	5.06	6.63
19	3	5	6	27	36	19.25	-16.37	-0.19	5.06	6.63
19	3	5	6	28	37	19.25	-16.37	-0.44	5.06	6.63
19	3	5	6	29	37	19.19	-16.31	-0.06	5.06	6.63
19	3	5	6	30	37	19.19	-16.44	-0.88	5.06	6.63
19	3	5	6	31	37	19.12	-16.25	0	5.06	6.63
19	3	5	6	32	38	19.12	-16.31	-0.63	5.06	6.63
19	3	5	6	33	38	19.12	-16.19	0.12	5.06	6.56
19	3	5	6	34	38	19.12	-16.25	-0.44	5.06	6.63
19	3	5	6	35	38	19.12	-16.12	0.19	5.06	6.63
19	3	5	6	36	39	19.12	-16.19	-0.25	5.06	6.56
19	3	5	6	37	39	19.19	-16.37	-1.19	5.06	6.56
19	3	5	6	38	39	19.19	-16.12	-0.19	5.06	6.56
19	3	5	6	39	40	19.19	-16.25	-0.94	5.06	6.56
19	3	5	6	40	40	19.19	-16.06	-0.12	5.06	6.56
19	3	5	6	41	40	19.19	-16.19	-0.75	5.06	6.56
19	3	5	6	42	40	19.25	-16	0	5.06	6.56
19	3	5	6	43	41	19.25	-16.06	-0.56	5.06	6.56
19	3	5	6	44	41	19.25	-15.94	0.06	5.06	6.56
19	3	5	6	45	41	19.31	-16	-0.25	5.06	6.56
19	3	5	6	46	41	19.31	-16.19	-1.31	5.06	6.56
19	3	5	6	47	42	19.31	-15.94	-0.25	5.06	6.56
19	3	5	6	48	42	19.31	-16.12	-1.06	5.06	6.56
19	3	5	6	49	42	19.31	-15.94	-0.19	5.06	6.56
19	3	5	6	50	42	19.31	-16	-0.88	5.06	6.56
19	3	5	6	51	43	19.31	-15.88	-0.12	5	6.56
19	3	5	6	52	43	19.25	-15.94	-0.75	5.06	6.56
19	3	5	6	53	43	19.25	-15.75	-0.06	5.06	6.56
19	3	5	6	54	43	19.25	-15.81	-0.5	5.06	6.5
19	3	5	6	55	44	19.19	-16	-1.44	5.06	6.5
19	3	5	6	56	44	19.19	-15.81	-0.37	5.06	6.5
19	3	5	6	57	44	19.12	-15.94	-1.25	5.06	6.5
19	3	5	6	58	44	19.12	-15.75	-0.37	5.06	6.5
19	3	5	6	59	45	19.12	-15.88	-1.12	5.06	6.5
19	3	5	7	0	45	19.12	-15.69	-0.31	5.06	6.5
19	3	5	7	1	45	19.12	-15.81	-0.94	5	6.5
19	3	5	7	2	45	19.12	-15.63	-0.25	5.06	6.5
19	3	5	7	3	46	19.12	-15.69	-0.81	5	6.44
19	3	5	7	4	46	19.19	-15.81	-1.5	5.06	6.44
19	3	5	7	5	46	19.19	-15.63	-0.81	5	6.44
19	3	5	7	6	47	19.25	-15.75	-1.31	5	6.44
19	3	5	7	7	47	19.25	-15.56	-0.44	5.06	6.44
19	3	5	7	8	47	19.31	-15.69	-1.25	5	6.44
19	3	5	7	9	47	19.31	-15.5	-0.44	5	6.38
19	3	5	7	10	48	19.31	-15.63	-1.12	5	6.44
19	3	5	7	11	48	19.37	-15.44	-0.44	5	6.38
19	3	5	7	12	48	19.37	-15.56	-1.06	5	6.38
19	3	5	7	13	48	19.44	-15.44	-0.81	5	6.38
19	3	5	7	14	49	19.44	-15.5	-0.94	5	6.38
19	3	5	7	15	49	19.5	-15.63	-1.88	5	6.38
19	3	5	7	16	49	19.5	-15.38	-0.81	5	6.31
19	3	5	7	17	49	19.5	-15.56	-1.69	5	6.31
19	3	5	7	18	50	19.5	-15.31	-0.75	5	6.31
19	3	5	7	19	50	19.44	-15.44	-1.44	5	6.31
19	3	5	7	20	50	19.44	-15.25	-0.63	5	6.31
19	3	5	7	21	50	19.44	-15.38	-1.31	5	6.31
19	3	5	7	22	51	19.44	-15.19	-0.56	5	6.31
19	3	5	7	23	51	19.37	-15.25	-1.19	5	6.31
19	3	5	7	24	51	19.37	-15.25	-1.56	5	6.31
19	3	5	7	25	51	19.31	-15.19	-1	5	6.25
19	3	5	7	26	52	19.31	-15.31	-2	5	6.31
19	3	5	7	27	52	19.25	-15.13	-0.94	5	6.25
19	3	5	7	28	52	19.25	-15.25	-1.81	5	6.25
19	3	5	7	29	52	19.25	-15.06	-0.88	5	6.25
19	3	5	7	30	53	19.25	-15.19	-1.63	5	6.25
19	3	5	7	31	53	19.25	-15	-0.81	5	6.25
19	3	5	7	32	53	19.31	-15.13	-1.44	5	6.25
19	3	5	7	33	53	19.31	-15	-0.69	5	6.25
19	3	5	7	34	54	19.25	-15	-1.31	5	6.25
19	3	5	7	35	54	19.25	-15	-1.69	5	6.25
19	3	5	7	36	54	19.25	-14.94	-1.19	5	6.25
19	3	5	7	37	55	19.25	-15.13	-2.25	5	6.25
19	3	5	7	38	55	19.19	-14.88	-1.12	5	6.19
19	3	5	7	39	55	19.19	-15	-2	5	6.19
19	3	5	7	40	55	19.12	-14.81	-1.06	5	6.19
19	3	5	7	41	56	19.12	-14.94	-1.81	5	6.19
19	3	5	7	42	56	19.12	-14.75	-0.94	5	6.19
19	3	5	7	43	56	19.12	-14.88	-1.69	5	6.19
19	3	5	7	44	56	19.12	-14.69	-0.88	5	6.19
19	3	5	7	45	57	19.12	-14.81	-1.5	5	6.19
19	3	5	7	46	57	19.19	-14.75	-1.31	5	6.19
19	3	5	7	47	57	19.19	-14.69	-1.44	5	6.19
19	3	5	7	48	57	19.19	-14.81	-2.5	5	6.13
19	3	5	7	49	58	19.25	-14.63	-1.37	4.94	6.13
19	3	5	7	50	58	19.25	-14.75	-2.31	5	6.13
19	3	5	7	51	58	19.25	-14.56	-1.25	5	6.13
19	3	5	7	52	58	19.31	-14.69	-2.13	4.94	6.13
19	3	5	7	53	59	19.31	-14.5	-1.19	4.94	6.13
19	3	5	7	54	59	19.37	-14.63	-1.94	4.94	6.13
19	3	5	7	55	59	19.37	-14.44	-1.06	4.94	6.06
19	3	5	7	56	59	19.37	-14.5	-1.81	4.94	6.06
19	3	5	7	58	0	19.37	-14.38	-1	4.94	6.06
19	3	5	7	59	0	19.37	-14.44	-1.63	4.94	6.06
19	3	5	8	0	0	19.37	-14.5	-1.81	4.94	6.06
19	3	5	8	1	0	19.37	-14.38	-1.44	4.94	6.06
19	3	5	8	2	1	19.44	-14.5	-2.5	4.94	6.06
19	3	5	8	3	1	19.37	-14.31	-1.31	4.94	6.06
19	3	5	8	4	1	19.31	-14.44	-2.25	4.94	6.06
19	3	5	8	5	1	19.31	-14.25	-1.25	4.94	6.06
19	3	5	8	6	2	19.25	-14.38	-2.06	4.94	6.06
19	3	5	8	7	2	19.25	-14.19	-1.12	4.94	6.06
19	3	5	8	8	2	19.25	-14.25	-1.81	4.94	6.06
19	3	5	8	9	2	19.19	-14.13	-1	4.94	6.06
19	3	5	8	10	3	19.19	-14.19	-1.69	4.94	6.06
19	3	5	8	11	3	19.19	-14.06	-1	4.94	6.06
19	3	5	8	12	3	19.19	-14.13	-1.5	4.94	6.06
19	3	5	8	13	3	19.19	-14.25	-2.31	4.94	6.06
19	3	5	8	14	4	19.19	-14.06	-1.31	4.94	6.06
19	3	5	8	15	4	19.19	-14.25	-2.25	4.94	6.06
19	3	5	8	16	4	19.19	-14.06	-1.19	4.94	6.06
19	3	5	8	17	5	19.19	-14.13	-1.94	4.94	6.06
19	3	5	8	18	5	19.25	-14	-1.06	4.94	6.06
19	3	5	8	19	5	19.25	-14.06	-1.75	4.94	6.06
19	3	5	8	20	5	19.25	-13.94	-0.94	4.94	6.06
19	3	5	8	21	6	19.19	-14.06	-1.56	4.94	6.06
19	3	5	8	22	6	19.19	-13.88	-0.81	4.94	6.06
19	3	5	8	23	6	19.19	-14	-1.37	4.94	6.06
19	3	5	8	24	6	19.12	-14.06	-2.06	4.94	6.06
19	3	5	8	25	7	19.12	-13.94	-1.19	4.94	6.06
19	3	5	8	26	7	19.12	-14.06	-2.06	4.94	6.06
19	3	5	8	27	7	19.06	-13.88	-1.06	4.94	6.06
19	3	5	8	28	7	19.06	-14	-1.75	4.94	6.13
19	3	5	8	29	8	19	-13.81	-0.94	4.94	6.06
19	3	5	8	30	8	19	-13.88	-1.56	4.94	6.06
19	3	5	8	31	8	18.94	-13.75	-0.81	4.94	6.06
19	3	5	8	32	8	18.94	-13.81	-1.37	4.94	6.06
19	3	5	8	33	9	18.94	-13.69	-0.75	5	6.06
19	3	5	8	34	9	18.94	-13.81	-1.19	4.94	6.06
19	3	5	8	35	9	19	-13.88	-2.06	4.94	6.13
19	3	5	8	36	9	19	-13.69	-1	4.94	6.06
19	3	5	8	37	10	19	-13.81	-1.75	5	6.13
19	3	5	8	38	10	19.06	-13.69	-0.81	4.94	6.06
19	3	5	8	39	10	19.06	-13.75	-1.5	4.94	6.06
19	3	5	8	40	10	19.12	-13.63	-0.69	4.94	6.06
19	3	5	8	41	11	19.12	-13.69	-1.25	4.94	6.13
19	3	5	8	42	11	19.19	-13.56	-0.56	4.94	6.13
19	3	5	8	43	11	19.19	-13.63	-1.06	4.94	6.13
19	3	5	8	44	11	19.19	-13.69	-1.37	5	6.13
19	3	5	8	45	12	19.12	-13.56	-0.88	4.94	6.13
19	3	5	8	46	12	19.12	-13.69	-1.63	5	6.13
19	3	5	8	47	12	19.12	-13.56	-0.69	4.94	6.13
19	3	5	8	48	13	19.12	-13.69	-1.31	4.94	6.13
19	3	5	8	49	13	19.06	-13.56	-0.56	5	6.13
19	3	5	8	50	13	19.06	-13.63	-1.12	5	6.13
19	3	5	8	51	13	19.06	-13.5	-0.44	4.94	6.13
19	3	5	8	52	14	19.06	-13.56	-0.94	5	6.19
19	3	5	8	53	14	19	-13.5	-0.63	5	6.19
19	3	5	8	54	14	19.06	-13.56	-0.75	5	6.19
19	3	5	8	55	14	19.06	-13.63	-1.5	5	6.19
19	3	5	8	56	15	19.06	-13.44	-0.56	5	6.19
19	3	5	8	57	15	19.06	-13.56	-1.19	5	6.19
19	3	5	8	58	15	19.06	-13.44	-0.37	5	6.25
19	3	5	8	59	15	19.12	-13.56	-0.94	5	6.25
19	3	5	9	0	16	19.12	-13.44	-0.25	5	6.25
19	3	5	9	1	16	19.12	-13.5	-0.69	5	6.25
19	3	5	9	2	16	19.12	-13.44	-0.31	5	6.25
19	3	5	9	3	16	19.06	-13.5	-0.5	5	6.25
19	3	5	9	4	17	19.06	-13.63	-1.25	5	6.25
19	3	5	9	5	17	19.06	-13.44	-0.31	5	6.25
19	3	5	9	6	17	19	-13.56	-1	5	6.25
19	3	5	9	7	17	19	-13.44	-0.19	5	6.25
19	3	5	9	8	18	19	-13.5	-0.69	5	6.25
19	3	5	9	9	18	19	-13.38	-0.06	5	6.25
19	3	5	9	10	18	19	-13.44	-0.56	5	6.25
19	3	5	9	11	18	19	-13.38	-0.31	5	6.31
19	3	5	9	12	19	19	-13.38	-0.31	5	6.25
19	3	5	9	13	19	19	-13.5	-1	5	6.31
19	3	5	9	14	19	19	-13.38	-0.12	5	6.25
19	3	5	9	15	19	19	-13.44	-0.75	5	6.31
19	3	5	9	16	20	19	-13.31	0	5	6.31
19	3	5	9	17	20	19.06	-13.44	-0.5	5	6.31
19	3	5	9	18	20	19.06	-13.31	0.12	5	6.31
19	3	5	9	19	21	19.12	-13.38	-0.25	5	6.31
19	3	5	9	20	21	19.12	-13.38	-0.63	5	6.31
19	3	5	9	21	21	19.19	-13.31	-0.19	5	6.31
19	3	5	9	22	21	19.19	-13.44	-0.75	5	6.31
19	3	5	9	23	22	19.25	-13.25	0	5	6.31
19	3	5	9	24	22	19.25	-13.38	-0.5	5	6.31
19	3	5	9	25	22	19.25	-13.25	0.12	5	6.31
19	3	5	9	26	22	19.25	-13.31	-0.31	5.06	6.31
19	3	5	9	27	23	19.31	-13.19	0.31	5	6.31
19	3	5	9	28	23	19.25	-13.25	0	5	6.31
19	3	5	9	29	23	19.25	-13.38	-0.81	5.06	6.31
19	3	5	9	30	23	19.25	-13.19	0.12	5.06	6.31
19	3	5	9	31	24	19.25	-13.31	-0.5	5	6.31
19	3	5	9	32	24	19.25	-13.19	0.19	5.06	6.31
19	3	5	9	33	24	19.19	-13.25	-0.31	5	6.31
19	3	5	9	34	24	19.19	-13.13	0.31	5.06	6.38
19	3	5	9	35	25	19.19	-13.19	-0.12	5	6.38
19	3	5	9	36	25	19.19	-13.13	0.06	5.06	6.38
19	3	5	9	37	25	19.19	-13.13	-0.19	5	6.31
19	3	5	9	38	25	19.19	-13.25	-0.88	5.06	6.38
19	3	5	9	39	26	19.19	-13.06	0	5.06	6.31
19	3	5	9	40	26	19.19	-13.13	-0.56	5.06	6.38
19	3	5	9	41	26	19.19	-13	0.19	5.06	6.38
19	3	5	9	42	26	19.25	-13.13	-0.31	5.06	6.38
19	3	5	9	43	27	19.25	-12.94	0.31	5.06	6.38
19	3	5	9	44	27	19.25	-13	-0.19	5.06	6.38
19	3	5	9	45	27	19.25	-12.94	0.06	5.06	6.38
19	3	5	9	46	31	19.25	-12.94	0.06	5.06	6.38
19	3	5	9	47	31	19.25	-13.06	-0.69	5.06	6.38
19	3	5	9	48	31	19.25	-12.94	0.12	5.06	6.38
19	3	5	9	49	32	19.25	-13	-0.44	5.06	6.38
19	3	5	9	50	32	19.19	-12.88	0.25	5.06	6.38
19	3	5	9	51	32	19.19	-12.94	-0.25	5.06	6.38
19	3	5	9	52	32	19.19	-12.81	0.37	5.06	6.38
19	3	5	9	53	33	19.12	-12.88	-0.06	5.06	6.38
19	3	5	9	54	33	19.12	-12.88	-0.44	5.06	6.38
19	3	5	9	55	33	19.12	-12.81	0	5.06	6.38
19	3	5	9	56	34	19.12	-12.88	-0.75	5.06	6.38
19	3	5	9	57	34	19.12	-12.75	0.12	5.06	6.38
19	3	5	9	58	34	19.12	-12.81	-0.5	5.06	6.38
19	3	5	9	59	34	19.12	-12.69	0.25	5.06	6.38
19	3	5	10	0	35	19.12	-12.75	-0.31	5.06	6.38
19	3	5	10	1	35	19.12	-12.63	0.31	5.06	6.38
19	3	5	10	2	35	19.12	-12.69	-0.19	5.06	6.38
19	3	5	10	3	35	19.12	-12.63	-0.12	5.06	6.38
19	3	5	10	4	36	19.19	-12.56	-0.06	5.06	6.31
19	3	5	10	5	36	19.19	-12.69	-0.88	5.06	6.31
19	3	5	10	6	36	19.19	-12.5	0.06	5.06	6.31
19	3	5	10	7	36	19.19	-12.63	-0.69	5.06	6.31
19	3	5	10	8	37	19.19	-12.44	0.12	5.06	6.31
19	3	5	10	9	37	19.25	-12.5	-0.5	5.06	6.31
19	3	5	10	10	37	19.25	-12.38	0.19	5.06	6.31
19	3	5	10	11	37	19.31	-12.44	-0.37	5.06	6.31
19	3	5	10	12	38	19.31	-12.31	0.25	5.06	6.31
19	3	5	10	13	38	19.31	-12.38	-0.06	5.06	6.31
19	3	5	10	14	38	19.31	-12.44	-0.94	5.06	6.31
19	3	5	10	15	38	19.31	-12.31	0	5.06	6.31
19	3	5	10	16	39	19.37	-12.38	-0.81	5.06	6.31
19	3	5	10	17	39	19.37	-12.25	0.06	5.06	6.31
19	3	5	10	18	39	19.37	-12.31	-0.63	5.06	6.31
19	3	5	10	19	39	19.31	-12.19	0.12	5.06	6.31
19	3	5	10	20	40	19.31	-12.25	-0.5	5.06	6.31
19	3	5	10	21	40	19.31	-12.13	0.19	5.06	6.25
19	3	5	10	22	40	19.31	-12.19	-0.37	5.06	6.25
19	3	5	10	23	40	19.25	-12.06	-0.12	5.06	6.25
19	3	5	10	24	41	19.25	-12.06	-0.25	5.06	6.25
19	3	5	10	25	41	19.19	-12.19	-1.12	5.06	6.25
19	3	5	10	26	41	19.19	-12	-0.12	5.06	6.25
19	3	5	10	27	42	19.19	-12.13	-0.94	5.06	6.25
19	3	5	10	28	42	19.12	-11.94	-0.06	5.06	6.25
19	3	5	10	29	42	19.12	-12	-0.69	5.06	6.25
19	3	5	10	30	42	19.12	-11.88	0.06	5.06	6.25
19	3	5	10	31	43	19.12	-11.94	-0.56	5.06	6.25
19	3	5	10	32	43	19.12	-11.81	0.06	5.06	6.19
19	3	5	10	33	43	19.12	-11.88	-0.44	5.06	6.25
19	3	5	10	34	43	19.12	-11.81	-0.25	5.06	6.25
19	3	5	10	35	44	19.12	-11.81	-0.31	5	6.19
19	3	5	10	36	44	19.19	-11.88	-1	5.06	6.25
19	3	5	10	37	44	19.19	-11.75	-0.06	5.06	6.19
19	3	5	10	38	44	19.19	-11.81	-0.94	5.06	6.25
19	3	5	10	39	45	19.19	-11.63	-0.06	5.06	6.19
19	3	5	10	40	45	19.19	-11.75	-0.75	5.06	6.25
19	3	5	10	41	45	19.19	-11.63	0	5.06	6.19
19	3	5	10	42	45	19.12	-11.69	-0.63	5.06	6.19
19	3	5	10	43	46	19.12	-11.5	0.06	5.06	6.19
19	3	5	10	44	46	19.12	-11.56	-0.44	5.06	6.19
19	3	5	10	45	46	19.06	-11.44	0.12	5.06	6.19
19	3	5	10	46	46	19.06	-11.5	-0.37	5.06	6.19
19	3	5	10	47	47	19.06	-11.56	-1.06	5.06	6.19
19	3	5	10	48	47	19	-11.44	-0.37	5.06	6.19
19	3	5	10	49	47	19	-11.5	-1.25	5.06	6.19
19	3	5	10	50	47	19	-11.31	-0.31	5.06	6.13
19	3	5	10	51	48	19	-11.44	-1.06	5.06	6.13
19	3	5	10	52	48	19.06	-11.25	-0.19	5.06	6.13
19	3	5	10	53	48	19.06	-11.31	-0.88	5.06	6.13
19	3	5	10	54	48	19.06	-11.19	-0.06	5.06	6.13
19	3	5	10	55	49	19.06	-11.19	-0.69	5.06	6.13
19	3	5	10	56	49	19.06	-11.06	0	5.06	6.13
19	3	5	10	57	49	19.06	-11.13	-0.63	5	6.13
19	3	5	10	58	50	19.06	-11.06	-0.12	5.06	6.13
19	3	5	10	59	50	19.06	-11.06	-0.5	5	6.06
19	3	5	11	0	50	19.06	-11.13	-1.12	5.06	6.13
19	3	5	11	1	50	19.06	-11	-0.44	5	6.06
19	3	5	11	2	51	19	-11.06	-1.37	5	6.06
19	3	5	11	3	51	19.06	-10.88	-0.31	5	6.06
19	3	5	11	4	51	19.06	-10.94	-1.12	5.06	6.06
19	3	5	11	5	51	19.06	-10.81	-0.25	5.06	6.06
19	3	5	11	6	52	19.06	-10.88	-1.25	5	6.06
19	3	5	11	7	52	19.12	-10.75	-0.37	5	6.06
19	3	5	11	8	52	19.12	-10.75	-1	5	6
19	3	5	11	9	52	19.19	-10.63	-0.25	5	6
19	3	5	11	10	53	19.19	-10.63	-0.81	5	6
19	3	5	11	11	53	19.25	-10.56	-0.12	5	6
19	3	5	11	12	53	19.25	-10.56	-0.69	5	6
19	3	5	11	13	54	19.25	-10.5	-0.44	5	6
19	3	5	11	14	54	19.25	-10.5	-0.56	5	6
19	3	5	11	15	54	19.25	-10.56	-1.25	5	6
19	3	5	11	16	54	19.19	-10.44	-0.5	5	6
19	3	5	11	17	55	19.19	-10.5	-1.44	5	6
19	3	5	11	18	55	19.19	-10.38	-0.44	5	5.94
19	3	5	11	19	55	19.19	-10.44	-1.19	5	6
19	3	5	11	20	55	19.12	-10.25	-0.31	5	6
19	3	5	11	21	56	19.12	-10.31	-1	5	6
19	3	5	11	22	56	19.12	-10.19	-0.25	5	5.94
19	3	5	11	23	56	19.12	-10.25	-0.88	5	6
19	3	5	11	24	56	19.19	-10.13	-0.12	5	6
19	3	5	11	25	57	19.19	-10.13	-0.75	5	5.94
19	3	5	11	26	57	19.19	-10.06	-0.06	5	5.94
19	3	5	11	27	57	19.25	-10.06	-0.69	5	5.94
19	3	5	11	28	57	19.19	-10	-0.19	5	5.94
19	3	5	11	29	58	19.19	-10	-0.56	5	5.94
19	3	5	11	30	58	19.19	-10	-0.94	5	5.94
19	3	5	11	31	58	19.19	-9.94	-0.5	5	5.94
19	3	5	11	32	58	19.19	-10	-1.44	5	5.94
19	3	5	11	33	59	19.12	-9.88	-0.37	5	5.94
19	3	5	11	34	59	19.12	-9.88	-1.25	5	5.94
19	3	5	11	35	59	19.12	-9.75	-0.31	5	5.94
19	3	5	11	36	59	19.12	-9.81	-0.94	5	5.94
19	3	5	11	38	0	19.06	-9.69	-0.12	5	5.94
19	3	5	11	39	0	19.06	-9.75	-0.88	5	5.94
19	3	5	11	40	0	19.06	-9.63	-0.12	5	5.94
19	3	5	11	41	0	19.06	-9.63	-0.81	5	5.94
19	3	5	11	42	1	19.12	-9.56	-0.12	5	5.94
19	3	5	11	43	1	19.12	-9.56	-0.75	5.06	5.94
19	3	5	11	44	1	19.12	-9.5	-0.06	5	5.94
19	3	5	11	45	2	19.12	-9.5	-0.63	5	5.94
19	3	5	11	46	2	19.12	-9.44	-0.12	5	5.94
19	3	5	11	47	2	19.12	-9.38	-0.56	5	5.88
19	3	5	11	48	2	19.12	-9.38	-0.88	5	5.88
19	3	5	11	49	3	19.12	-9.31	-0.5	5	5.88
19	3	5	11	50	3	19.19	-9.31	-1.31	5	5.88
19	3	5	11	51	3	19.12	-9.25	-0.44	5	5.88
19	3	5	11	52	3	19.12	-9.25	-1.31	5	5.88
19	3	5	11	53	4	19.12	-9.13	-0.37	5	5.88
19	3	5	11	54	4	19.12	-9.19	-1.12	5	5.88
19	3	5	11	55	4	19.12	-9.06	-0.25	5	5.88
19	3	5	11	56	4	19.12	-9.06	-1	5.06	5.88
19	3	5	11	57	5	19.12	-8.94	-0.19	5	5.88
19	3	5	11	58	5	19.12	-9	-0.94	5	5.88
19	3	5	11	59	5	19.19	-8.88	-0.19	5	5.88
19	3	5	12	0	5	19.19	-8.88	-0.88	5	5.88
19	3	5	12	1	6	19.25	-8.81	-0.12	5	5.88
19	3	5	12	2	6	19.25	-8.81	-0.75	5	5.81
19	3	5	12	3	6	19.25	-8.75	-0.06	5	5.88
19	3	5	12	4	6	19.25	-8.69	-0.69	5	5.81
19	3	5	12	5	7	19.25	-8.69	0	5	5.81
19	3	5	12	6	7	19.25	-8.63	-0.56	5	5.81
19	3	5	12	7	7	19.25	-8.63	-0.5	5	5.88
19	3	5	12	8	7	19.19	-8.56	-0.5	5	5.81
19	3	5	12	9	8	19.19	-8.56	-0.94	5	5.88
19	3	5	12	10	8	19.19	-8.5	-0.44	5	5.81
19	3	5	12	11	8	19.19	-8.5	-1.25	5	5.81
19	3	5	12	12	8	19.12	-8.38	-0.37	5.06	5.81
19	3	5	12	13	9	19.12	-8.38	-1.31	5	5.81
19	3	5	12	14	9	19.12	-8.31	-0.31	5	5.81
19	3	5	12	15	9	19.12	-8.31	-1.19	5	5.81
19	3	5	12	16	9	19.12	-8.25	-0.25	5	5.81
19	3	5	12	17	10	19.19	-8.25	-1.06	5	5.81
19	3	5	12	18	10	19.19	-8.19	-0.25	5	5.81
19	3	5	12	19	10	19.19	-8.13	-1	5	5.81
19	3	5	12	20	11	19.25	-8.06	-0.19	5	5.81
19	3	5	12	21	11	19.25	-8.06	-0.88	5	5.81
19	3	5	12	22	11	19.25	-8	-0.12	5	5.81
19	3	5	12	23	11	19.25	-8	-0.75	5	5.81
19	3	5	12	24	12	19.31	-7.94	-0.06	5	5.81
19	3	5	12	25	12	19.31	-7.88	-0.69	5	5.81
19	3	5	12	26	12	19.31	-7.81	0	5	5.81
19	3	5	12	27	12	19.37	-7.81	-0.63	5	5.81
19	3	5	12	28	13	19.37	-7.75	0	5	5.81
19	3	5	12	29	13	19.37	-7.75	-0.56	5	5.81
19	3	5	12	30	13	19.44	-7.69	0	5	5.81
19	3	5	12	31	13	19.44	-7.69	-0.5	5	5.81
19	3	5	12	32	14	19.44	-7.69	-0.31	5	5.81
19	3	5	12	33	14	19.44	-7.63	-0.44	5	5.81
19	3	5	12	34	14	19.5	-7.63	-0.63	5	5.81
19	3	5	12	35	14	19.5	-7.5	-0.37	5	5.81
19	3	5	12	36	15	19.5	-7.5	-0.88	5	5.81
19	3	5	12	37	15	19.5	-7.44	-0.31	5	5.81
19	3	5	12	38	15	19.5	-7.38	-1.12	5	5.81
19	3	5	12	39	15	19.56	-7.31	-0.37	5	5.81
19	3	5	12	40	16	19.56	-7.31	-1.25	5	5.81
19	3	5	12	41	16	19.56	-7.19	-0.31	5	5.81
19	3	5	12	42	16	19.56	-7.25	-1.12	5	5.81
19	3	5	12	43	16	19.56	-7.13	-0.25	4.94	5.81
19	3	5	12	44	17	19.62	-7.13	-1	4.94	5.81
19	3	5	12	45	17	19.62	-7.06	-0.19	4.94	5.81
19	3	5	12	46	17	19.62	-7.06	-0.88	4.94	5.81
19	3	5	12	47	18	19.62	-7	-0.12	4.94	5.81
19	3	5	12	48	18	19.62	-6.94	-0.81	4.94	5.81
19	3	5	12	49	18	19.69	-6.88	-0.06	4.94	5.81
19	3	5	12	50	18	19.69	-6.88	-0.69	4.94	5.81
19	3	5	12	51	19	19.69	-6.81	0	4.94	5.81
19	3	5	12	52	19	19.75	-6.81	-0.63	4.94	5.81
19	3	5	12	53	19	19.75	-6.75	0.06	4.94	5.81
19	3	5	12	54	19	19.81	-6.75	-0.56	4.94	5.81
19	3	5	12	55	20	19.81	-6.75	0.06	4.94	5.81
19	3	5	12	56	20	19.87	-6.69	-0.5	4.94	5.88
19	3	5	12	57	20	19.87	-6.63	0.12	4.94	5.81
19	3	5	12	58	20	19.87	-6.56	-0.44	4.94	5.88
19	3	5	12	59	21	19.94	-6.56	0.19	4.94	5.88
19	3	5	13	0	21	20	-6.5	-0.31	4.88	5.88
19	3	5	13	1	21	20	-6.5	0.25	4.94	5.88
19	3	5	13	2	21	20	-6.44	-0.19	4.88	5.88
19	3	5	13	3	22	20.06	-6.44	0.12	4.94	5.88
19	3	5	13	4	22	20.06	-6.38	-0.12	4.88	5.88
19	3	5	13	5	22	20.06	-6.44	-0.19	4.94	5.88
19	3	5	13	6	22	20.06	-6.31	-0.06	4.88	5.88
19	3	5	13	7	23	20.12	-6.38	-0.44	4.88	5.88
19	3	5	13	8	23	20.12	-6.25	0	4.88	5.88
19	3	5	13	9	23	20.12	-6.25	-0.63	4.88	5.88
19	3	5	13	10	23	20.12	-6.19	0.12	4.88	5.88
19	3	5	13	11	24	20.12	-6.19	-0.63	4.88	5.94
19	3	5	13	12	24	20.12	-6.13	0.19	4.88	5.88
19	3	5	13	13	24	20.19	-6.19	-0.5	4.88	5.94
19	3	5	13	14	24	20.12	-6.13	0.25	4.88	5.94
19	3	5	13	15	25	20.19	-6.06	-0.37	4.88	5.94
19	3	5	13	16	25	20.19	-6	0.31	4.88	5.94
19	3	5	13	17	25	20.19	-6	-0.25	4.88	5.94
19	3	5	13	18	26	20.19	-6	0.44	4.94	5.94
19	3	5	13	19	26	20.19	-5.94	-0.12	4.94	5.94
19	3	5	13	20	26	20.19	-5.94	0.5	4.94	5.94
19	3	5	13	21	26	20.19	-5.88	-0.06	4.94	5.94
19	3	5	13	22	27	20.19	-5.88	0.56	4.94	5.94
19	3	5	13	23	27	20.19	-5.81	0.06	4.94	5.94
19	3	5	13	24	27	20.19	-5.81	0.63	4.94	6
19	3	5	13	25	27	20.25	-5.75	0.12	4.94	6
19	3	5	13	26	28	20.19	-5.75	0.69	4.94	6
19	3	5	13	27	28	20.25	-5.75	0.5	4.94	6
19	3	5	13	28	28	20.25	-5.81	0.63	4.94	6
19	3	5	13	29	28	20.25	-5.75	0.5	4.94	6
19	3	5	13	30	29	20.25	-5.75	0.19	4.94	6
19	3	5	13	31	29	20.25	-5.69	0.56	4.94	6.06
19	3	5	13	32	29	20.25	-5.75	-0.06	5	6.06
19	3	5	13	33	29	20.25	-5.69	0.63	5	6.06
19	3	5	13	34	30	20.25	-5.69	-0.06	5	6.06
19	3	5	13	35	30	20.25	-5.63	0.69	5	6.06
19	3	5	13	36	32	20.25	-5.63	0.12	5	6.06
19	3	5	13	37	32	20.31	-5.56	0.75	5	6.06
19	3	5	13	38	33	20.31	-5.56	0.25	5	6.06
19	3	5	13	39	33	20.25	-5.56	0.88	5	6.06
19	3	5	13	40	33	20.25	-5.56	0.37	5	6.06
19	3	5	13	41	33	20.25	-5.5	0.94	5	6.06
19	3	5	13	42	34	20.19	-5.5	0.5	5	6.06
19	3	5	13	43	34	20.12	-5.5	1	5	6.13
19	3	5	13	44	34	20.12	-5.44	0.56	5	6.06
19	3	5	13	45	34	20.06	-5.44	1.06	5	6.13
19	3	5	13	46	35	20.06	-5.44	0.63	5.06	6.13
19	3	5	13	47	35	20	-5.38	0.88	5.06	6.13
19	3	5	13	48	35	20	-5.38	0.75	5.06	6.13
19	3	5	13	49	36	19.94	-5.38	0.44	5.06	6.13
19	3	5	13	50	36	19.94	-5.31	0.88	5.06	6.13
19	3	5	13	51	36	19.87	-5.31	0.25	5.06	6.13
19	3	5	13	52	36	19.81	-5.25	1	5.06	6.13
19	3	5	13	53	37	19.81	-5.31	0.37	5.06	6.13
19	3	5	13	54	37	19.75	-5.19	1	5.06	6.13
19	3	5	13	55	37	19.75	-5.25	0.5	5.06	6.13
19	3	5	13	56	37	19.69	-5.13	1.12	5.06	6.13
19	3	5	13	57	38	19.69	-5.19	0.63	5.06	6.13
19	3	5	13	58	38	19.69	-5.13	1.19	5.06	6.13
19	3	5	13	59	38	19.62	-5.13	0.75	5.06	6.13
19	3	5	14	0	38	19.62	-5.06	1.25	5.06	6.13
19	3	5	14	1	39	19.62	-5.06	0.88	5.06	6.13
19	3	5	14	2	39	19.62	-5	1.31	5.13	6.19
19	3	5	14	3	39	19.62	-5	0.94	5.13	6.19
19	3	5	14	4	39	19.62	-5	1.31	5.13	6.19
19	3	5	14	5	40	19.62	-4.94	1	5.13	6.19
19	3	5	14	6	40	19.56	-5	0.94	5.13	6.19
19	3	5	14	7	40	19.56	-4.94	1.12	5.13	6.19
19	3	5	14	8	40	19.56	-4.94	0.56	5.13	6.19
19	3	5	14	9	41	19.56	-4.88	1.19	5.13	6.19
19	3	5	14	10	41	19.56	-4.88	0.5	5.13	6.19
19	3	5	14	11	41	19.62	-4.81	1.19	5.13	6.19
19	3	5	14	12	41	19.62	-4.81	0.63	5.13	6.19
19	3	5	14	13	42	19.62	-4.75	1.25	5.13	6.19
19	3	5	14	14	42	19.62	-4.75	0.75	5.13	6.19
19	3	5	14	15	42	19.62	-4.69	1.31	5.13	6.19
19	3	5	14	16	43	19.69	-4.63	0.81	5.13	6.19
19	3	5	14	17	43	19.69	-4.56	1.37	5.13	6.19
19	3	5	14	18	43	19.69	-4.56	0.88	5.13	6.19
19	3	5	14	19	43	19.69	-4.5	1.37	5.13	6.19
19	3	5	14	20	44	19.69	-4.44	0.88	5.13	6.19
19	3	5	14	21	44	19.69	-4.44	1.37	5.13	6.19
19	3	5	14	22	44	19.62	-4.38	1	5.13	6.19
19	3	5	14	23	44	19.62	-4.38	1.44	5.13	6.19
19	3	5	14	24	45	19.56	-4.31	1.06	5.13	6.19
19	3	5	14	25	45	19.56	-4.31	1.06	5.13	6.19
19	3	5	14	26	45	19.56	-4.25	1.12	5.13	6.19
19	3	5	14	27	45	19.5	-4.25	0.75	5.13	6.19
19	3	5	14	28	46	19.5	-4.19	1.19	5.13	6.19
19	3	5	14	29	46	19.5	-4.13	0.5	5.13	6.19
19	3	5	14	30	46	19.5	-4.13	1.19	5.13	6.19
19	3	5	14	31	46	19.5	-4.06	0.56	5.13	6.19
19	3	5	14	32	47	19.5	-4.06	1.25	5.13	6.19
19	3	5	14	33	47	19.5	-4	0.63	5.13	6.19
19	3	5	14	34	47	19.56	-4	1.31	5.13	6.19
19	3	5	14	35	47	19.56	-3.94	0.75	5.13	6.19
19	3	5	14	36	48	19.56	-3.94	1.31	5.13	6.19
19	3	5	14	37	48	19.56	-3.88	0.81	5.13	6.19
19	3	5	14	38	48	19.62	-3.88	1.37	5.13	6.19
19	3	5	14	39	48	19.62	-3.81	0.88	5.13	6.19
19	3	5	14	40	49	19.62	-3.81	1.37	5.13	6.19
19	3	5	14	41	49	19.62	-3.75	0.94	5.13	6.19
19	3	5	14	42	49	19.56	-3.75	1.44	5.13	6.19
19	3	5	14	43	49	19.56	-3.69	1	5.19	6.19
19	3	5	14	44	50	19.56	-3.69	1.44	5.19	6.19
19	3	5	14	45	50	19.56	-3.63	1.06	5.13	6.19
19	3	5	14	46	50	19.5	-3.63	1.5	5.19	6.19
19	3	5	14	47	51	19.5	-3.63	1.12	5.19	6.13
19	3	5	14	48	51	19.44	-3.63	1.44	5.19	6.19
19	3	5	14	49	51	19.44	-3.56	1.12	5.19	6.19
19	3	5	14	50	51	19.44	-3.56	1.12	5.19	6.19
19	3	5	14	51	52	19.37	-3.56	1.19	5.19	6.19
19	3	5	14	52	52	19.37	-3.56	0.94	5.19	6.19
19	3	5	14	53	52	19.31	-3.5	1.25	5.19	6.19
19	3	5	14	54	52	19.31	-3.5	0.75	5.19	6.19
19	3	5	14	55	53	19.31	-3.44	1.25	5.19	6.19
19	3	5	14	56	53	19.31	-3.44	0.63	5.19	6.19
19	3	5	14	57	53	19.37	-3.44	1.31	5.19	6.19
19	3	5	14	58	53	19.37	-3.44	0.63	5.19	6.19
19	3	5	14	59	54	19.44	-3.38	1.31	5.19	6.19
19	3	5	15	0	54	19.44	-3.31	0.69	5.19	6.19
19	3	5	15	1	54	19.5	-3.31	1.37	5.19	6.19
19	3	5	15	2	54	19.5	-3.25	0.75	5.19	6.19
19	3	5	15	3	55	19.5	-3.19	1.37	5.19	6.19
19	3	5	15	4	55	19.5	-3.19	0.81	5.19	6.19
19	3	5	15	5	55	19.56	-3.13	1.44	5.19	6.19
19	3	5	15	6	55	19.5	-3.06	0.88	5.19	6.19
19	3	5	15	7	56	19.5	-3.06	1.44	5.19	6.19
19	3	5	15	8	56	19.5	-3	0.94	5.19	6.19
19	3	5	15	9	56	19.5	-3	1.44	5.19	6.19
19	3	5	15	10	56	19.44	-2.94	1	5.19	6.13
19	3	5	15	11	57	19.44	-2.94	1.5	5.19	6.13
19	3	5	15	12	57	19.44	-2.88	1.06	5.19	6.19
19	3	5	15	13	57	19.44	-2.88	1.5	5.19	6.19
19	3	5	15	14	57	19.44	-2.81	1.06	5.19	6.19
19	3	5	15	15	58	19.5	-2.81	1.56	5.19	6.13
19	3	5	15	16	58	19.5	-2.75	1.12	5.19	6.13
19	3	5	15	17	58	19.5	-2.75	1.56	5.19	6.13
19	3	5	15	18	59	19.5	-2.69	1.19	5.19	6.13
19	3	5	15	19	59	19.44	-2.69	1.63	5.19	6.13
19	3	5	15	20	59	19.44	-2.63	1.19	5.19	6.13
19	3	5	15	21	59	19.44	-2.63	1.63	5.19	6.13
19	3	5	15	23	0	19.37	-2.56	1.25	5.19	6.13
19	3	5	15	24	0	19.37	-2.56	1.69	5.25	6.13
19	3	5	15	25	0	19.37	-2.5	1.31	5.19	6.13
19	3	5	15	26	0	19.37	-2.5	1.56	5.25	6.13
19	3	5	15	27	1	19.31	-2.44	1.37	5.25	6.13
19	3	5	15	28	1	19.31	-2.5	1.44	5.19	6.13
19	3	5	15	29	1	19.31	-2.38	1.37	5.19	6.13
19	3	5	15	30	1	19.31	-2.44	1.31	5.25	6.13
19	3	5	15	31	2	19.37	-2.31	1.44	5.25	6.13
19	3	5	15	32	2	19.37	-2.38	1.25	5.25	6.13
19	3	5	15	33	2	19.37	-2.31	1.44	5.25	6.13
19	3	5	15	34	2	19.37	-2.31	1.19	5.25	6.13
19	3	5	15	35	3	19.44	-2.25	1.5	5.25	6.13
19	3	5	15	36	3	19.44	-2.25	1.12	5.25	6.13
19	3	5	15	37	3	19.44	-2.19	1.5	5.25	6.13
19	3	5	15	38	3	19.44	-2.19	1.12	5.25	6.13
19	3	5	15	39	4	19.5	-2.13	1.5	5.25	6.13
19	3	5	15	40	4	19.5	-2.06	1.12	5.25	6.13
19	3	5	15	41	4	19.5	-2.06	1.56	5.25	6.13
19	3	5	15	42	4	19.56	-2	1.12	5.25	6.13
19	3	5	15	43	5	19.56	-1.94	1.56	5.25	6.13
19	3	5	15	44	5	19.56	-1.94	1.12	5.25	6.13
19	3	5	15	45	5	19.62	-1.88	1.56	5.25	6.06
19	3	5	15	46	5	19.62	-1.88	1.12	5.25	6.13
19	3	5	15	47	6	19.62	-1.81	1.56	5.25	6.06
19	3	5	15	48	6	19.62	-1.75	1.12	5.25	6.06
19	3	5	15	49	6	19.62	-1.75	1.56	5.25	6.06
19	3	5	15	50	6	19.69	-1.69	1.19	5.25	6.06
19	3	5	15	51	7	19.69	-1.69	1.63	5.25	6.06
19	3	5	15	52	7	19.69	-1.56	1.25	5.25	6.06
19	3	5	15	53	7	19.69	-1.56	1.63	5.25	6.06
19	3	5	15	54	8	19.75	-1.5	1.25	5.25	6.06
19	3	5	15	55	8	19.75	-1.5	1.69	5.25	6.06
19	3	5	15	56	8	19.75	-1.44	1.31	5.25	6.06
19	3	5	15	57	8	19.75	-1.44	1.69	5.25	6
19	3	5	15	58	9	19.81	-1.37	1.31	5.25	6.06
19	3	5	15	59	9	19.81	-1.37	1.75	5.25	6
19	3	5	16	0	9	19.81	-1.31	1.31	5.25	6
19	3	5	16	1	9	19.81	-1.31	1.75	5.25	6
19	3	5	16	2	10	19.87	-1.25	1.37	5.25	6
19	3	5	16	3	10	19.87	-1.25	1.75	5.25	6
19	3	5	16	4	10	19.94	-1.19	1.37	5.25	6
19	3	5	16	5	10	19.94	-1.19	1.81	5.25	6
19	3	5	16	6	11	20	-1.12	1.44	5.25	6
19	3	5	16	7	11	20	-1.12	1.81	5.25	6
19	3	5	16	8	11	20.06	-1.06	1.44	5.25	6
19	3	5	16	9	11	20.06	-1.06	1.81	5.25	6
19	3	5	16	10	12	20.06	-1	1.5	5.25	6
19	3	5	16	11	12	20.12	-1.06	1.88	5.25	6
19	3	5	16	12	12	20.12	-1	1.5	5.25	6
19	3	5	16	13	12	20.12	-1	1.88	5.25	6
19	3	5	16	14	13	20.12	-0.94	1.5	5.25	6
19	3	5	16	15	13	20.19	-0.94	1.94	5.25	6
19	3	5	16	16	13	20.19	-0.88	1.56	5.25	6
19	3	5	16	17	13	20.19	-0.88	1.94	5.25	6
19	3	5	16	18	14	20.19	-0.81	1.56	5.25	6
19	3	5	16	19	14	20.19	-0.88	1.94	5.25	6
19	3	5	16	20	14	20.19	-0.81	1.63	5.25	6
19	3	5	16	21	14	20.19	-0.81	2	5.25	6
19	3	5	16	22	15	20.19	-0.81	1.63	5.25	6.06
19	3	5	16	23	15	20.25	-0.75	2	5.25	6
19	3	5	16	24	15	20.25	-0.75	1.69	5.25	6
19	3	5	16	25	16	20.25	-0.75	2	5.25	6
19	3	5	16	26	16	20.25	-0.69	1.75	5.31	6
19	3	5	16	27	16	20.25	-0.69	2.06	5.25	6
19	3	5	16	28	16	20.25	-0.63	1.75	5.25	6
19	3	5	16	29	17	20.25	-0.63	2.06	5.31	6
19	3	5	16	30	17	20.25	-0.63	1.81	5.31	6.06
19	3	5	16	31	17	20.25	-0.56	2.06	5.25	6
19	3	5	16	32	17	20.25	-0.56	1.81	5.31	6
19	3	5	16	33	18	20.25	-0.5	2.06	5.25	6.06
19	3	5	16	34	18	20.19	-0.56	1.88	5.31	6.06
19	3	5	16	35	18	20.19	-0.5	2.06	5.31	6.06
19	3	5	16	36	18	20.12	-0.5	1.94	5.31	6.06
19	3	5	16	37	19	20.12	-0.44	2.13	5.31	6.06
19	3	5	16	38	19	20.06	-0.5	2	5.31	6.06
19	3	5	16	39	19	20	-0.44	2.13	5.31	6.06
19	3	5	16	40	19	20	-0.44	2.06	5.31	6.06
19	3	5	16	41	20	19.94	-0.37	2.13	5.31	6.06
19	3	5	16	42	20	19.94	-0.44	2.13	5.31	6.06
19	3	5	16	43	20	19.87	-0.31	2.13	5.31	6
19	3	5	16	44	20	19.87	-0.37	2.19	5.31	6.06
19	3	5	16	45	21	19.81	-0.31	2.13	5.31	6.06
19	3	5	16	46	21	19.81	-0.37	2.25	5.31	6.06
19	3	5	16	47	21	19.75	-0.25	2.13	5.31	6.06
19	3	5	16	48	21	19.69	-0.31	2.38	5.31	6.06
19	3	5	16	49	22	19.69	-0.19	2.13	5.31	6.06
19	3	5	16	50	22	19.62	-0.25	2.44	5.31	6.06
19	3	5	16	51	22	19.62	-0.12	2.13	5.31	6.06
19	3	5	16	52	22	19.56	-0.19	2.5	5.31	6.06
19	3	5	16	53	23	19.56	-0.12	2.19	5.31	6.06
19	3	5	16	54	23	19.5	-0.19	2.5	5.31	6.06
19	3	5	16	55	23	19.5	-0.06	2.19	5.31	6.06
19	3	5	16	56	24	19.44	-0.12	2.5	5.31	6.06
19	3	5	16	57	24	19.44	0	2.19	5.31	6.06
19	3	5	16	58	24	19.37	-0.06	2.5	5.31	6.06
19	3	5	16	59	24	19.37	0	2.19	5.31	6.06
19	3	5	17	0	25	19.37	0	2.5	5.31	6.06
19	3	5	17	1	25	19.37	0.06	2.19	5.31	6.06
19	3	5	17	2	25	19.37	0	2.5	5.31	6.06
19	3	5	17	3	25	19.37	0.12	2.19	5.31	6.06
19	3	5	17	4	26	19.44	0.06	2.5	5.31	6.06
19	3	5	17	5	26	19.44	0.12	2.13	5.31	6.06
19	3	5	17	6	26	19.44	0.12	2.5	5.31	6.06
19	3	5	17	7	26	19.5	0.19	2.13	5.31	6.06
19	3	5	17	8	27	19.5	0.12	2.5	5.38	6.06
19	3	5	17	9	27	19.56	0.25	2.13	5.31	6.06
19	3	5	17	10	27	19.56	0.19	2.56	5.38	6.06
19	3	5	17	11	27	19.62	0.25	2.13	5.38	6.06
19	3	5	17	12	28	19.62	0.25	2.5	5.38	6.06
19	3	5	17	13	28	19.62	0.31	2.13	5.38	6.06
19	3	5	17	14	28	19.69	0.25	2.56	5.38	6.06
19	3	5	17	15	28	19.62	0.31	2.13	5.38	6.06
19	3	5	17	16	29	19.62	0.31	2.56	5.38	6.06
19	3	5	17	17	29	19.62	0.37	2.13	5.38	6.06
19	3	5	17	18	29	19.62	0.31	2.56	5.38	6.06
19	3	5	17	19	29	19.56	0.37	2.13	5.38	6.06
19	3	5	17	20	30	19.56	0.37	2.56	5.38	6.06
19	3	5	17	21	30	19.56	0.37	2.25	5.38	6.06
19	3	5	17	22	30	19.5	0.37	2.5	5.38	6.06
19	3	5	17	23	30	19.5	0.44	2.31	5.38	6.06
19	3	5	17	24	31	19.5	0.44	2.5	5.38	6.06
19	3	5	17	25	31	19.44	0.44	2.38	5.38	6.06
19	3	5	17	26	35	19.44	0.44	2.56	5.38	6.06
19	3	5	17	27	35	19.37	0.44	2.44	5.38	6.06
19	3	5	17	28	35	19.44	0.5	2.56	5.38	6.06
19	3	5	17	29	36	19.44	0.5	2.5	5.38	6.06
19	3	5	17	30	36	19.44	0.56	2.56	5.38	6.06
19	3	5	17	31	36	19.44	0.5	2.63	5.38	6.06
19	3	5	17	32	36	19.44	0.56	2.56	5.38	6.06
19	3	5	17	33	37	19.44	0.5	2.81	5.38	6.06
19	3	5	17	34	37	19.44	0.63	2.56	5.38	6.06
19	3	5	17	35	37	19.44	0.56	2.81	5.38	6.06
19	3	5	17	36	37	19.44	0.63	2.5	5.38	6.06
19	3	5	17	37	38	19.44	0.56	2.81	5.38	6.06
19	3	5	17	38	38	19.44	0.69	2.5	5.38	6.06
19	3	5	17	39	38	19.44	0.63	2.81	5.38	6.06
19	3	5	17	40	38	19.44	0.75	2.5	5.38	6.06
19	3	5	17	41	39	19.44	0.69	2.81	5.38	6.06
19	3	5	17	42	39	19.37	0.75	2.5	5.44	6.06
19	3	5	17	43	39	19.37	0.69	2.81	5.38	6.06
19	3	5	17	44	39	19.37	0.81	2.5	5.38	6.06
19	3	5	17	45	40	19.37	0.75	2.81	5.44	6.06
19	3	5	17	46	40	19.31	0.81	2.5	5.38	6.13
19	3	5	17	47	40	19.31	0.81	2.81	5.38	6.06
19	3	5	17	48	40	19.31	0.88	2.44	5.44	6.13
19	3	5	17	49	41	19.31	0.81	2.81	5.44	6.06
19	3	5	17	50	41	19.31	0.88	2.44	5.44	6.13
19	3	5	17	51	41	19.31	0.88	2.81	5.44	6.06
19	3	5	17	52	41	19.31	0.88	2.56	5.44	6.13
19	3	5	17	53	42	19.31	0.94	2.81	5.44	6.06
19	3	5	17	54	42	19.37	0.88	2.63	5.44	6.13
19	3	5	17	55	42	19.37	0.94	2.81	5.44	6.06
19	3	5	17	56	43	19.37	0.88	2.81	5.44	6.13
19	3	5	17	57	43	19.37	0.94	2.81	5.44	6.13
19	3	5	17	58	43	19.37	0.94	2.94	5.44	6.13
19	3	5	17	59	43	19.37	1	2.81	5.44	6.13
19	3	5	18	0	44	19.37	0.94	3.06	5.44	6.13
19	3	5	18	1	44	19.37	1.06	2.81	5.44	6.13
19	3	5	18	2	44	19.37	1	3.06	5.44	6.13
19	3	5	18	3	44	19.37	1.06	2.75	5.44	6.13
19	3	5	18	4	45	19.37	1.06	3.06	5.44	6.13
19	3	5	18	5	45	19.37	1.12	2.75	5.44	6.13
19	3	5	18	6	45	19.31	1.06	3.06	5.44	6.13
19	3	5	18	7	45	19.31	1.19	2.75	5.44	6.13
19	3	5	18	8	46	19.37	1.12	3.06	5.44	6.13
19	3	5	18	9	46	19.31	1.19	2.69	5.44	6.13
19	3	5	18	10	46	19.37	1.19	3.06	5.44	6.13
19	3	5	18	11	46	19.37	1.25	2.69	5.44	6.13
19	3	5	18	12	47	19.37	1.19	3.06	5.44	6.13
19	3	5	18	13	47	19.37	1.25	2.75	5.5	6.13
19	3	5	18	14	47	19.44	1.25	3	5.5	6.13
19	3	5	18	15	47	19.44	1.25	2.81	5.5	6.13
19	3	5	18	16	48	19.5	1.31	3	5.44	6.13
19	3	5	18	17	48	19.5	1.31	2.94	5.5	6.13
19	3	5	18	18	48	19.5	1.37	3	5.44	6.13
19	3	5	18	19	48	19.5	1.31	3.06	5.5	6.13
19	3	5	18	20	49	19.5	1.44	3	5.5	6.13
19	3	5	18	21	49	19.5	1.31	3.25	5.5	6.13
19	3	5	18	22	49	19.5	1.44	3	5.5	6.13
19	3	5	18	23	49	19.5	1.37	3.25	5.5	6.19
19	3	5	18	24	50	19.5	1.5	3	5.5	6.13
19	3	5	18	25	50	19.5	1.44	3.25	5.5	6.13
19	3	5	18	26	50	19.5	1.5	2.94	5.5	6.19
19	3	5	18	27	51	19.44	1.5	3.25	5.5	6.19
Yr	Mo	Dy	Hr	Mn	Sc	Amb	Ice	Pip	Bot	Top
19	3	5	18	37	50	19.62	1.19	4.25	5.56	6.63
19	3	5	18	38	50	19.62	1.19	4.31	5.56	6.69
19	3	5	18	39	50	19.69	1.19	4.44	5.56	6.75
19	3	5	18	40	51	19.69	1.19	4.5	5.56	6.75
19	3	5	18	41	51	19.69	1.19	4.5	5.63	6.81
19	3	5	18	42	51	19.69	1.19	4.63	5.63	6.88
19	3	5	18	43	51	19.69	1.19	4.69	5.63	6.94
19	3	5	18	44	52	19.62	1.19	4.69	5.63	7
19	3	5	18	45	52	19.62	1.19	4.75	5.63	7.06
19	3	5	18	46	52	19.62	1.19	4.81	5.63	7.06
19	3	5	18	47	52	19.56	1.19	4.88	5.69	7.13
19	3	5	18	48	53	19.56	1.19	4.94	5.69	7.19
19	3	5	18	49	53	19.56	1.19	4.94	5.63	7.19
19	3	5	18	50	53	19.56	1.19	5.06	5.69	7.25
19	3	5	18	51	53	19.5	1.19	5.06	5.69	7.25
19	3	5	18	52	54	19.5	1.19	5.13	5.69	7.31
19	3	5	18	53	54	19.5	1.19	5.19	5.69	7.38
19	3	5	18	54	54	19.5	1.19	5.25	5.69	7.38
19	3	5	18	55	54	19.5	1.19	5.31	5.69	7.44
19	3	5	18	56	55	19.5	1.19	5.31	5.69	7.44
19	3	5	18	57	55	19.56	1.19	5.38	5.75	7.5
19	3	5	18	58	55	19.56	1.19	5.38	5.75	7.5
19	3	5	18	59	55	19.56	1.19	5.44	5.75	7.56
19	3	5	19	0	56	19.56	1.19	5.5	5.75	7.56
19	3	5	19	1	56	19.56	1.19	5.5	5.75	7.63
19	3	5	19	2	56	19.56	1.19	5.56	5.75	7.63
19	3	5	19	3	56	19.56	1.19	5.63	5.75	7.63
19	3	5	19	4	57	19.5	1.19	5.63	5.75	7.69
19	3	5	19	5	57	19.5	1.19	5.69	5.75	7.69
19	3	5	19	6	57	19.5	1.19	5.75	5.75	7.75
19	3	5	19	7	57	19.5	1.19	5.75	5.75	7.75
19	3	5	19	8	58	19.5	1.19	5.81	5.81	7.75
19	3	5	19	9	58	19.44	1.19	5.81	5.81	7.75
19	3	5	19	10	58	19.44	1.25	5.88	5.81	7.81
19	3	5	19	11	59	19.44	1.25	5.81	5.81	7.81
19	3	5	19	12	59	19.5	1.25	5.88	5.81	7.81
19	3	5	19	13	59	19.5	1.25	5.88	5.81	7.81
19	3	5	19	14	59	19.5	1.25	5.88	5.81	7.88
19	3	5	19	16	0	19.56	1.25	5.94	5.81	7.88
19	3	5	19	17	0	19.56	1.25	5.94	5.81	7.88
19	3	5	19	18	0	19.56	1.25	6	5.81	7.88
19	3	5	19	19	0	19.62	1.25	6	5.81	7.88
19	3	5	19	20	1	19.62	1.25	5.94	5.81	7.94
19	3	5	19	21	1	19.62	1.25	6	5.81	7.94
19	3	5	19	22	1	19.62	1.25	5.94	5.81	7.94
19	3	5	19	23	1	19.62	1.25	6.06	5.81	7.94
19	3	5	19	24	2	19.62	1.25	6	5.88	7.94
19	3	5	19	25	2	19.62	1.19	5.94	5.81	7.94
19	3	5	19	26	2	19.62	1.19	6.06	5.88	7.94
19	3	5	19	27	3	19.56	1.19	6	5.88	8
19	3	5	19	28	3	19.56	1.19	6.06	5.88	7.94
19	3	5	19	29	3	19.56	1.19	6	5.88	8
19	3	5	19	30	3	19.5	1.19	5.94	5.88	8
19	3	5	19	31	4	19.5	1.19	6.06	5.88	8
19	3	5	19	32	4	19.5	1.12	5.94	5.88	8
19	3	5	19	33	4	19.44	1.19	6.06	5.88	8
19	3	5	19	34	4	19.44	1.12	6	5.88	8
19	3	5	19	35	5	19.37	1.19	5.94	5.88	8
19	3	5	19	36	5	19.37	1.12	6.06	5.88	8
19	3	5	19	37	5	19.37	1.12	5.94	5.88	8
19	3	5	19	38	5	19.37	1.19	6.06	5.88	8
19	3	5	19	39	6	19.37	1.19	6.06	5.88	8
19	3	5	19	40	6	19.44	1.19	6.13	5.94	8
19	3	5	19	41	6	19.44	1.19	6	5.94	8
19	3	5	19	42	6	19.44	1.19	5.94	5.94	8.06
19	3	5	19	43	7	19.5	1.19	6.06	5.94	8
19	3	5	19	44	7	19.5	1.19	5.94	5.94	8
19	3	5	19	45	7	19.5	1.19	6.06	5.94	8.06
19	3	5	19	46	7	19.5	1.19	6	5.94	8
19	3	5	19	47	8	19.56	1.19	5.88	5.94	8.06
19	3	5	19	48	8	19.5	1.25	6	5.94	8
19	3	5	19	49	8	19.5	1.25	5.88	5.94	8
19	3	5	19	50	8	19.5	1.25	6	5.94	8
19	3	5	19	51	9	19.5	1.25	5.94	5.94	8
19	3	5	19	52	9	19.44	1.25	5.88	5.94	8
19	3	5	19	53	9	19.44	1.31	5.94	5.94	8
19	3	5	19	54	10	19.44	1.25	5.81	5.94	8
19	3	5	19	55	10	19.44	1.31	6	5.94	8
19	3	5	19	56	10	19.44	1.31	5.88	6	8
19	3	5	19	57	10	19.37	1.31	6	6	8
19	3	5	19	58	11	19.37	1.37	5.88	6	8
19	3	5	19	59	11	19.44	1.31	5.75	6	8
19	3	5	20	0	11	19.44	1.37	5.94	6	8
19	3	5	20	1	11	19.44	1.37	5.81	6	8
19	3	5	20	2	12	19.44	1.37	5.94	6	8
19	3	5	20	3	12	19.5	1.37	5.81	6	8
19	3	5	20	4	12	19.5	1.44	5.75	6	8
19	3	5	20	5	12	19.5	1.37	5.81	6	8
19	3	5	20	6	13	19.56	1.44	5.75	6	8
19	3	5	20	7	13	19.5	1.44	5.88	6	8
19	3	5	20	8	13	19.5	1.44	5.75	6	8
19	3	5	20	9	13	19.5	1.44	5.88	6	8
19	3	5	20	10	14	19.5	1.5	5.75	6	8
19	3	5	20	11	14	19.5	1.5	5.63	6	8
19	3	5	20	12	14	19.5	1.5	5.81	6	8
19	3	5	20	13	14	19.44	1.5	5.69	6	8
19	3	5	20	14	15	19.44	1.5	5.81	6	7.94
19	3	5	20	15	15	19.44	1.56	5.69	6	7.94
19	3	5	20	16	15	19.44	1.56	5.81	6	7.94
19	3	5	20	17	15	19.44	1.56	5.75	6.06	7.94
19	3	5	20	18	16	19.44	1.56	5.63	6.06	7.94
19	3	5	20	19	16	19.44	1.56	5.81	6.06	8
19	3	5	20	20	16	19.5	1.63	5.69	6.06	7.94
19	3	5	20	21	16	19.5	1.63	5.81	6.06	7.94
19	3	5	20	22	17	19.5	1.63	5.69	6.06	7.94
19	3	5	20	23	17	19.56	1.63	5.56	6.06	7.94
19	3	5	20	24	17	19.56	1.63	5.75	6.06	7.94
19	3	5	20	25	18	19.56	1.63	5.56	6.06	7.94
19	3	5	20	26	18	19.56	1.69	5.75	6.06	7.94
19	3	5	20	27	18	19.56	1.69	5.63	6.06	7.94
19	3	5	20	28	18	19.56	1.69	5.81	6.06	7.94
19	3	5	20	29	19	19.56	1.69	5.75	6.06	7.94
19	3	5	20	30	19	19.5	1.69	5.56	6.06	7.94
19	3	5	20	31	19	19.5	1.69	5.81	6.06	7.94
19	3	5	20	32	19	19.5	1.75	5.56	6.06	7.94
19	3	5	20	33	20	19.5	1.75	5.75	6.13	7.94
19	3	5	20	34	20	19.44	1.75	5.63	6.13	7.94
19	3	5	20	35	20	19.5	1.81	5.81	6.06	7.94
19	3	5	20	36	20	19.5	1.81	5.69	6.13	7.94
19	3	5	20	37	21	19.5	1.81	5.44	6.13	7.94
19	3	5	20	38	21	19.5	1.81	5.69	6.13	7.94
19	3	5	20	39	21	19.5	1.81	5.5	6.13	7.94
19	3	5	20	40	21	19.5	1.88	5.75	6.13	7.94
19	3	5	20	41	22	19.5	1.88	5.56	6.13	7.94
19	3	5	20	42	22	19.5	1.88	5.75	6.13	7.94
19	3	5	20	43	22	19.5	1.94	5.63	6.13	7.94
19	3	5	20	44	22	19.5	1.94	5.38	6.13	7.94
19	3	5	20	45	23	19.44	1.94	5.63	6.13	7.94
19	3	5	20	46	23	19.44	1.94	5.5	6.13	7.94
19	3	5	20	47	23	19.44	2	5.69	6.13	7.94
19	3	5	20	48	27	19.44	2	5.56	6.13	7.94
19	3	5	20	49	27	19.44	2	5.75	6.13	7.94
19	3	5	20	50	27	19.44	2.06	5.63	6.13	7.94
19	3	5	20	51	28	19.44	2.06	5.38	6.13	7.88
19	3	5	20	52	28	19.44	2.06	5.63	6.13	7.88
19	3	5	20	53	28	19.44	2.06	5.5	6.13	7.88
19	3	5	20	54	28	19.44	2.13	5.69	6.19	7.94
19	3	5	20	55	29	19.5	2.13	5.5	6.19	7.88
19	3	5	20	56	29	19.5	2.13	5.69	6.19	7.88
19	3	5	20	57	29	19.5	2.19	5.56	6.19	7.88
19	3	5	20	58	30	19.5	2.19	5.38	6.19	7.88
19	3	5	20	59	30	19.56	2.19	5.63	6.19	7.88
19	3	5	21	0	30	19.5	2.25	5.44	6.19	7.88
19	3	5	21	1	30	19.5	2.25	5.69	6.19	7.88
19	3	5	21	2	31	19.5	2.25	5.5	6.19	7.88
19	3	5	21	3	31	19.5	2.31	5.69	6.19	7.88
19	3	5	21	4	31	19.44	2.31	5.56	6.19	7.88
19	3	5	21	5	31	19.44	2.31	5.38	6.19	7.88
19	3	5	21	6	32	19.44	2.38	5.63	6.19	7.88
19	3	5	21	7	32	19.37	2.38	5.44	6.19	7.88
19	3	5	21	8	32	19.37	2.38	5.63	6.25	7.88
19	3	5	21	9	32	19.37	2.44	5.5	6.25	7.88
19	3	5	21	10	33	19.37	2.44	5.69	6.19	7.88
19	3	5	21	11	33	19.37	2.5	5.56	6.25	7.88
19	3	5	21	12	33	19.37	2.5	5.56	6.25	7.88
19	3	5	21	13	33	19.44	2.5	5.63	6.25	7.88
19	3	5	21	14	34	19.44	2.5	5.38	6.25	7.88
19	3	5	21	15	34	19.44	2.56	5.63	6.25	7.88
19	3	5	21	16	34	19.5	2.56	5.5	6.25	7.88
19	3	5	21	17	34	19.5	2.63	5.69	6.25	7.88
19	3	5	21	18	35	19.56	2.63	5.56	6.25	7.88
19	3	5	21	19	35	19.56	2.69	5.69	6.25	7.88
19	3	5	21	20	35	19.56	2.69	5.63	6.25	7.88
19	3	5	21	21	35	19.62	2.69	5.38	6.25	7.88
19	3	5	21	22	36	19.56	2.75	5.69	6.25	7.88
19	3	5	21	23	36	19.62	2.75	5.5	6.25	7.88
19	3	5	21	24	36	19.56	2.75	5.69	6.25	7.88
19	3	5	21	25	36	19.56	2.81	5.56	6.25	7.88
19	3	5	21	26	37	19.56	2.81	5.75	6.25	7.88
19	3	5	21	27	37	19.56	2.88	5.63	6.31	7.88
19	3	5	21	28	37	19.56	2.88	5.63	6.31	7.88
19	3	5	21	29	38	19.5	2.88	5.63	6.31	7.88
19	3	5	21	30	38	19.5	2.88	5.44	6.31	7.88
19	3	5	21	31	38	19.5	2.94	5.69	6.31	7.88
19	3	5	21	32	38	19.5	2.94	5.5	6.31	7.88
19	3	5	21	33	39	19.5	2.94	5.75	6.31	7.88
19	3	5	21	34	39	19.5	3	5.56	6.31	7.88
19	3	5	21	35	39	19.5	3	5.75	6.31	7.88
19	3	5	21	36	39	19.5	3.06	5.63	6.31	7.88
19	3	5	21	37	40	19.56	3.06	5.56	6.31	7.88
19	3	5	21	38	40	19.56	3.06	5.69	6.31	7.88
19	3	5	21	39	40	19.56	3.13	5.5	6.31	7.88
19	3	5	21	40	40	19.56	3.13	5.75	6.31	7.88
19	3	5	21	41	41	19.56	3.13	5.56	6.31	7.88
19	3	5	21	42	41	19.62	3.19	5.75	6.31	7.88
19	3	5	21	43	41	19.56	3.19	5.63	6.38	7.88
19	3	5	21	44	41	19.56	3.25	5.81	6.38	7.88
19	3	5	21	45	42	19.56	3.25	5.69	6.38	7.88
19	3	5	21	46	42	19.56	3.25	5.63	6.38	7.88
19	3	5	21	47	42	19.56	3.31	5.75	6.38	7.88
19	3	5	21	48	42	19.5	3.31	5.63	6.38	7.88
19	3	5	21	49	43	19.5	3.31	5.88	6.38	7.88
19	3	5	21	50	43	19.44	3.31	5.69	6.38	7.94
19	3	5	21	51	43	19.44	3.31	5.88	6.38	7.88
19	3	5	21	52	43	19.44	3.38	5.69	6.38	7.94
19	3	5	21	53	44	19.44	3.38	5.88	6.38	7.94
19	3	5	21	54	44	19.44	3.44	5.75	6.38	7.88
19	3	5	21	55	44	19.44	3.44	5.75	6.38	7.94
19	3	5	21	56	44	19.44	3.5	5.81	6.38	7.94
19	3	5	21	57	45	19.44	3.44	5.63	6.38	7.94
19	3	5	21	58	45	19.5	3.5	5.88	6.38	7.94
19	3	5	21	59	45	19.5	3.5	5.69	6.44	7.94
19	3	5	22	0	46	19.5	3.56	5.88	6.38	7.94
19	3	5	22	1	46	19.5	3.56	5.75	6.44	7.94
19	3	5	22	2	46	19.5	3.63	5.94	6.44	7.94
19	3	5	22	3	46	19.5	3.63	5.81	6.44	7.94
19	3	5	22	4	47	19.5	3.63	5.94	6.44	7.94
19	3	5	22	5	47	19.5	3.69	5.81	6.44	7.94
19	3	5	22	6	47	19.5	3.69	5.69	6.44	7.94
19	3	5	22	7	47	19.44	3.75	5.88	6.44	7.94
19	3	5	22	8	48	19.44	3.75	5.75	6.44	7.94
19	3	5	22	9	48	19.37	3.75	6	6.44	7.94
19	3	5	22	10	48	19.37	3.75	5.81	6.44	7.94
19	3	5	22	11	48	19.37	3.81	6	6.44	7.94
19	3	5	22	12	49	19.37	3.81	5.88	6.5	7.94
19	3	5	22	13	49	19.31	3.81	6	6.44	7.94
19	3	5	22	14	49	19.31	3.88	5.88	6.44	7.94
19	3	5	22	15	49	19.31	3.88	5.94	6.5	7.94
19	3	5	22	16	50	19.37	3.94	5.94	6.5	7.94
19	3	5	22	17	50	19.37	3.94	5.75	6.5	7.94
19	3	5	22	18	50	19.37	3.94	6	6.5	7.94
19	3	5	22	19	50	19.44	4	5.81	6.5	7.94
19	3	5	22	20	51	19.44	4	6	6.5	7.94
19	3	5	22	21	51	19.5	4.06	5.88	6.5	7.94
19	3	5	22	22	51	19.5	4.06	6.06	6.5	7.94
19	3	5	22	23	51	19.5	4.06	5.94	6.5	7.94
19	3	5	22	24	52	19.5	4.13	6.06	6.5	7.94
19	3	5	22	25	52	19.56	4.13	6	6.5	7.94
19	3	5	22	26	52	19.5	4.13	6	6.5	7.94
19	3	5	22	27	52	19.5	4.19	6.06	6.5	7.94
19	3	5	22	28	53	19.5	4.19	5.88	6.5	7.94
19	3	5	22	29	53	19.5	4.25	6.06	6.56	7.94
19	3	5	22	30	53	19.5	4.25	5.88	6.56	7.94
19	3	5	22	31	54	19.44	4.25	6.13	6.56	8
19	3	5	22	32	54	19.44	4.31	6	6.56	8
19	3	5	22	33	54	19.44	4.31	6.13	6.56	8
19	3	5	22	34	54	19.44	4.31	6	6.56	7.94
19	3	5	22	35	55	19.44	4.38	6.19	6.56	8
19	3	5	22	36	55	19.44	4.38	6.06	6.56	8
19	3	5	22	37	55	19.5	4.38	6.13	6.56	8
19	3	5	22	38	55	19.5	4.44	6.13	6.56	8
19	3	5	22	39	56	19.5	4.5	6	6.56	8
19	3	5	22	40	56	19.5	4.5	6.19	6.56	8
19	3	5	22	41	56	19.56	4.5	6	6.63	8
19	3	5	22	42	56	19.56	4.56	6.19	6.63	8
19	3	5	22	43	57	19.56	4.56	6.06	6.63	8
19	3	5	22	44	57	19.62	4.63	6.25	6.63	8
19	3	5	22	45	57	19.62	4.63	6.13	6.63	8
19	3	5	22	46	57	19.62	4.63	6.25	6.63	8
19	3	5	22	47	58	19.69	4.69	6.19	6.63	8
19	3	5	22	48	58	19.69	4.69	6.31	6.63	8
19	3	5	22	49	58	19.69	4.69	6.25	6.63	8
19	3	5	22	50	58	19.69	4.75	6.19	6.63	8.06
19	3	5	22	51	59	19.69	4.75	6.25	6.63	8
19	3	5	22	52	59	19.62	4.81	6.13	6.63	8
19	3	5	22	53	59	19.62	4.81	6.31	6.63	8
19	3	5	22	54	59	19.62	4.88	6.19	6.63	8.06
19	3	5	22	56	0	19.62	4.88	6.38	6.69	8
19	3	5	22	57	0	19.56	4.88	6.25	6.69	8.06
19	3	5	22	58	0	19.56	4.94	6.38	6.69	8.06
19	3	5	22	59	1	19.5	4.94	6.31	6.69	8.06
19	3	5	23	0	1	19.5	5	6.44	6.69	8.06
19	3	5	23	1	1	19.5	5	6.31	6.69	8.06
19	3	5	23	2	1	19.44	5	6.5	6.69	8.06
19	3	5	23	3	2	19.44	5.06	6.38	6.69	8.06
19	3	5	23	4	2	19.37	5.06	6.44	6.69	8.06
19	3	5	23	5	2	19.37	5.13	6.44	6.69	8.06
19	3	5	23	6	2	19.37	5.13	6.31	6.69	8.06
19	3	5	23	7	3	19.31	5.13	6.44	6.69	8.06
19	3	5	23	8	3	19.31	5.19	6.31	6.69	8.06
19	3	5	23	9	3	19.31	5.19	6.5	6.75	8.06
19	3	5	23	10	3	19.37	5.25	6.38	6.69	8.06
19	3	5	23	11	4	19.37	5.25	6.5	6.75	8.06
19	3	5	23	12	4	19.37	5.31	6.44	6.75	8.13
19	3	5	23	13	4	19.37	5.31	6.56	6.75	8.13
19	3	5	23	14	4	19.44	5.31	6.5	6.75	8.13
19	3	5	23	15	5	19.44	5.31	6.63	6.75	8.13
19	3	5	23	16	5	19.5	5.38	6.5	6.75	8.13
19	3	5	23	17	5	19.5	5.38	6.63	6.75	8.13
19	3	5	23	18	5	19.5	5.44	6.56	6.75	8.13
19	3	5	23	19	6	19.5	5.44	6.56	6.75	8.13
19	3	5	23	20	6	19.5	5.5	6.63	6.75	8.13
19	3	5	23	21	6	19.5	5.5	6.56	6.75	8.13
19	3	5	23	22	6	19.5	5.5	6.63	6.81	8.13
19	3	5	23	23	7	19.5	5.5	6.5	6.81	8.13
19	3	5	23	24	7	19.44	5.56	6.69	6.81	8.13
19	3	5	23	25	7	19.44	5.56	6.56	6.81	8.13
19	3	5	23	26	7	19.44	5.63	6.69	6.81	8.13
19	3	5	23	27	8	19.37	5.63	6.63	6.81	8.13
19	3	5	23	28	8	19.37	5.63	6.75	6.81	8.13
19	3	5	23	29	8	19.37	5.69	6.69	6.81	8.19
19	3	5	23	30	9	19.37	5.69	6.75	6.81	8.19
19	3	5	23	31	9	19.44	5.69	6.69	6.81	8.19
19	3	5	23	32	9	19.44	5.75	6.81	6.81	8.19
19	3	5	23	33	9	19.44	5.75	6.75	6.81	8.19
19	3	5	23	34	10	19.44	5.75	6.81	6.88	8.19
19	3	5	23	35	10	19.44	5.81	6.81	6.81	8.19
19	3	5	23	36	10	19.44	5.81	6.81	6.88	8.19
19	3	5	23	37	10	19.44	5.81	6.81	6.88	8.19
19	3	5	23	38	11	19.44	5.81	6.75	6.88	8.19
19	3	5	23	39	11	19.44	5.88	6.88	6.88	8.19
19	3	5	23	40	11	19.37	5.88	6.75	6.88	8.19
19	3	5	23	41	11	19.37	5.94	6.88	6.88	8.25
19	3	5	23	42	12	19.37	5.94	6.81	6.88	8.25
19	3	5	23	43	12	19.31	5.94	6.94	6.88	8.25
19	3	5	23	44	12	19.31	6	6.81	6.88	8.25
19	3	5	23	45	12	19.31	6	6.94	6.88	8.25
19	3	5	23	46	13	19.31	6	6.88	6.88	8.25
19	3	5	23	47	13	19.37	6	7	6.94	8.25
19	3	5	23	48	13	19.37	6.06	6.88	6.94	8.25
19	3	5	23	49	13	19.37	6.06	7	6.94	8.25
19	3	5	23	50	14	19.31	6.06	6.94	6.94	8.25
19	3	5	23	51	14	19.31	6.06	7.06	6.94	8.25
19	3	5	23	52	14	19.31	6.13	6.94	6.94	8.25
19	3	5	23	53	14	19.31	6.13	7.06	6.94	8.25
19	3	5	23	54	15	19.25	6.13	7	6.94	8.25
19	3	5	23	55	15	19.25	6.19	7	6.94	8.25
19	3	5	23	56	15	19.25	6.19	7	6.94	8.25
19	3	5	23	57	15	19.25	6.19	7	6.94	8.31
19	3	5	23	58	16	19.25	6.19	7.06	6.94	8.25
19	3	5	23	59	16	19.25	6.25	7	6.94	8.31
19	3	6	0	0	16	19.25	6.25	7.06	7	8.31
19	3	6	0	1	17	19.25	6.25	7	7	8.31
19	3	6	0	2	17	19.31	6.25	7.06	7	8.31
19	3	6	0	3	17	19.31	6.31	7	7	8.31
19	3	6	0	4	17	19.31	6.31	7.13	7	8.31
19	3	6	0	5	18	19.31	6.31	7.06	7	8.31
19	3	6	0	6	18	19.31	6.38	7.13	7	8.31
19	3	6	0	7	18	19.31	6.38	7.06	7	8.31
19	3	6	0	8	18	19.31	6.38	7.19	7	8.31
19	3	6	0	9	19	19.31	6.38	7.13	7	8.31
19	3	6	0	10	19	19.25	6.44	7.19	7	8.31
19	3	6	0	11	19	19.25	6.44	7.13	7	8.31
19	3	6	0	12	19	19.25	6.44	7.19	7	8.31
19	3	6	0	13	20	19.25	6.44	7.13	7	8.31
19	3	6	0	14	20	19.25	6.5	7.25	7.06	8.38
19	3	6	0	15	20	19.25	6.5	7.19	7.06	8.38
19	3	6	0	16	20	19.25	6.5	7.25	7.06	8.31
19	3	6	0	17	21	19.25	6.5	7.19	7.06	8.38
19	3	6	0	18	21	19.25	6.5	7.25	7.06	8.31
19	3	6	0	19	21	19.31	6.56	7.25	7.06	8.38
19	3	6	0	20	21	19.31	6.56	7.25	7.06	8.38
19	3	6	0	21	22	19.31	6.56	7.25	7.06	8.38
19	3	6	0	22	22	19.31	6.56	7.25	7.06	8.38
19	3	6	0	23	22	19.31	6.63	7.25	7.13	8.38
19	3	6	0	24	22	19.31	6.63	7.25	7.06	8.38
19	3	6	0	25	23	19.31	6.63	7.31	7.13	8.38
19	3	6	0	26	23	19.31	6.63	7.25	7.13	8.38
19	3	6	0	27	23	19.31	6.69	7.31	7.13	8.38
19	3	6	0	28	23	19.25	6.69	7.25	7.13	8.38
19	3	6	0	29	24	19.25	6.69	7.38	7.13	8.38
19	3	6	0	30	24	19.25	6.69	7.25	7.13	8.38
19	3	6	0	31	24	19.19	6.75	7.38	7.13	8.44
19	3	6	0	32	25	19.19	6.75	7.31	7.13	8.38
19	3	6	0	33	25	19.19	6.75	7.38	7.13	8.44
19	3	6	0	34	25	19.19	6.75	7.38	7.13	8.44
19	3	6	0	35	25	19.25	6.81	7.44	7.13	8.44
19	3	6	0	36	26	19.25	6.81	7.38	7.13	8.44
19	3	6	0	37	26	19.25	6.81	7.44	7.13	8.44
19	3	6	0	38	30	19.25	6.81	7.38	7.19	8.44
19	3	6	0	39	30	19.25	6.88	7.5	7.19	8.44
19	3	6	0	40	30	19.31	6.88	7.44	7.19	8.44
19	3	6	0	41	30	19.31	6.88	7.5	7.19	8.44
19	3	6	0	42	31	19.31	6.88	7.44	7.19	8.44
19	3	6	0	43	31	19.37	6.88	7.5	7.19	8.44
19	3	6	0	44	31	19.37	6.94	7.44	7.19	8.44
19	3	6	0	45	31	19.37	6.94	7.5	7.19	8.44
19	3	6	0	46	32	19.44	6.94	7.5	7.19	8.44
19	3	6	0	47	32	19.44	6.94	7.56	7.19	8.44
19	3	6	0	48	32	19.44	7	7.5	7.25	8.5
19	3	6	0	49	32	19.5	7	7.56	7.19	8.5
19	3	6	0	50	33	19.5	7	7.56	7.25	8.5
19	3	6	0	51	33	19.5	7	7.56	7.25	8.5
19	3	6	0	52	33	19.56	7	7.56	7.25	8.5
19	3	6	0	53	33	19.56	7	7.56	7.25	8.5
19	3	6	0	54	34	19.56	7.06	7.56	7.25	8.5
19	3	6	0	55	34	19.56	7.06	7.56	7.25	8.5
19	3	6	0	56	34	19.62	7.06	7.63	7.25	8.5
19	3	6	0	57	34	19.62	7.06	7.56	7.25	8.5
19	3	6	0	58	35	19.62	7.06	7.63	7.25	8.5
19	3	6	0	59	35	19.62	7.13	7.69	7.25	8.5
19	3	6	1	0	35	19.56	7.13	7.69	7.25	8.56
19	3	6	1	1	36	19.56	7.13	7.75	7.25	8.56
19	3	6	1	2	36	19.56	7.13	7.75	7.25	8.56
19	3	6	1	3	36	19.56	7.13	7.75	7.25	8.63
19	3	6	1	4	36	19.56	7.13	7.75	7.25	8.63
19	3	6	1	5	37	19.5	7.13	7.75	7.25	8.69
19	3	6	1	6	37	19.5	7.19	7.81	7.25	8.69
19	3	6	1	7	37	19.5	7.19	7.81	7.25	8.75
19	3	6	1	8	37	19.5	7.19	7.81	7.19	8.75
19	3	6	1	9	38	19.44	7.19	7.81	7.25	8.75
19	3	6	1	10	38	19.44	7.19	7.88	7.25	8.81
19	3	6	1	11	38	19.44	7.19	7.88	7.25	8.81
19	3	6	1	12	38	19.44	7.19	7.88	7.25	8.88
19	3	6	1	13	39	19.44	7.19	7.88	7.25	8.88
19	3	6	1	14	39	19.37	7.19	7.88	7.25	8.88
19	3	6	1	15	39	19.37	7.19	7.88	7.25	8.94
19	3	6	1	16	39	19.37	7.19	7.94	7.25	8.94
19	3	6	1	17	40	19.37	7.19	7.94	7.25	9
19	3	6	1	18	40	19.37	7.25	7.94	7.25	9
19	3	6	1	19	40	19.37	7.25	7.94	7.25	9
19	3	6	1	20	40	19.37	7.25	8	7.25	9.06
19	3	6	1	21	41	19.44	7.25	8	7.25	9.06
19	3	6	1	22	41	19.44	7.25	8	7.25	9.06
19	3	6	1	23	41	19.44	7.25	8	7.25	9.06
19	3	6	1	24	41	19.44	7.25	8.06	7.25	9.13
19	3	6	1	25	42	19.44	7.31	8.06	7.25	9.13
19	3	6	1	26	42	19.5	7.31	8.06	7.25	9.13
19	3	6	1	27	42	19.5	7.31	8.06	7.25	9.13
19	3	6	1	28	42	19.5	7.31	8.06	7.25	9.19
19	3	6	1	29	43	19.5	7.31	8.06	7.25	9.19
19	3	6	1	30	43	19.5	7.38	8.06	7.25	9.19
19	3	6	1	31	43	19.5	7.38	8.13	7.25	9.25
19	3	6	1	32	44	19.44	7.38	8.13	7.25	9.25
19	3	6	1	33	44	19.44	7.38	8.13	7.25	9.25
19	3	6	1	34	44	19.37	7.44	8.13	7.31	9.25
19	3	6	1	35	44	19.37	7.44	8.13	7.25	9.25
19	3	6	1	36	45	19.37	7.44	8.19	7.31	9.25
19	3	6	1	37	45	19.37	7.44	8.19	7.31	9.31
19	3	6	1	38	45	19.37	7.5	8.19	7.31	9.31
19	3	6	1	39	45	19.31	7.5	8.19	7.31	9.31
19	3	6	1	40	46	19.31	7.5	8.19	7.31	9.31
19	3	6	1	41	46	19.31	7.56	8.19	7.31	9.31
19	3	6	1	42	46	19.31	7.56	8.25	7.31	9.31
19	3	6	1	43	46	19.37	7.56	8.25	7.31	9.31
19	3	6	1	44	47	19.37	7.63	8.25	7.31	9.38
19	3	6	1	45	47	19.37	7.63	8.25	7.31	9.38
19	3	6	1	46	47	19.37	7.63	8.25	7.31	9.38
19	3	6	1	47	47	19.44	7.69	8.25	7.31	9.38
19	3	6	1	48	48	19.44	7.69	8.25	7.31	9.38
19	3	6	1	49	48	19.44	7.69	8.31	7.31	9.38
19	3	6	1	50	48	19.44	7.75	8.31	7.38	9.44
19	3	6	1	51	48	19.44	7.75	8.31	7.31	9.44
19	3	6	1	52	49	19.37	7.75	8.31	7.38	9.44
19	3	6	1	53	49	19.37	7.81	8.31	7.38	9.44
19	3	6	1	54	49	19.37	7.81	8.31	7.38	9.44
19	3	6	1	55	49	19.37	7.88	8.31	7.38	9.44
19	3	6	1	56	50	19.31	7.88	8.38	7.38	9.44
19	3	6	1	57	50	19.31	7.88	8.38	7.38	9.44
19	3	6	1	58	50	19.31	7.88	8.38	7.38	9.5
19	3	6	1	59	50	19.31	7.94	8.38	7.38	9.44
19	3	6	2	0	51	19.25	7.94	8.38	7.38	9.5
19	3	6	2	1	51	19.25	7.94	8.38	7.38	9.5
19	3	6	2	2	51	19.25	8	8.38	7.38	9.5
19	3	6	2	3	52	19.25	8	8.38	7.38	9.5
19	3	6	2	4	52	19.31	8.06	8.44	7.38	9.5
19	3	6	2	5	52	19.31	8.06	8.44	7.38	9.5
19	3	6	2	6	52	19.31	8.13	8.44	7.38	9.5
19	3	6	2	7	53	19.37	8.13	8.44	7.38	9.56
19	3	6	2	8	53	19.37	8.13	8.44	7.44	9.5
19	3	6	2	9	53	19.37	8.19	8.44	7.44	9.56
19	3	6	2	10	53	19.44	8.19	8.44	7.44	9.5
19	3	6	2	11	54	19.44	8.19	8.44	7.44	9.56
19	3	6	2	12	54	19.44	8.25	8.44	7.44	9.56
19	3	6	2	13	54	19.44	8.25	8.5	7.44	9.56
19	3	6	2	14	54	19.44	8.31	8.5	7.44	9.56
19	3	6	2	15	55	19.37	8.31	8.5	7.44	9.56
19	3	6	2	16	55	19.37	8.31	8.5	7.44	9.56
19	3	6	2	17	55	19.37	8.38	8.5	7.44	9.56
19	3	6	2	18	55	19.37	8.38	8.5	7.44	9.56
19	3	6	2	19	56	19.37	8.44	8.5	7.44	9.56
19	3	6	2	20	56	19.37	8.44	8.5	7.44	9.63
19	3	6	2	21	56	19.31	8.44	8.5	7.44	9.63
19	3	6	2	22	56	19.37	8.5	8.5	7.44	9.63
19	3	6	2	23	57	19.31	8.5	8.56	7.5	9.63
19	3	6	2	24	57	19.37	8.56	8.56	7.5	9.63
19	3	6	2	25	57	19.37	8.56	8.56	7.5	9.63
19	3	6	2	26	57	19.37	8.56	8.56	7.5	9.63
19	3	6	2	27	58	19.37	8.63	8.56	7.5	9.63
19	3	6	2	28	58	19.44	8.63	8.56	7.5	9.63
19	3	6	2	29	58	19.44	8.69	8.56	7.5	9.63
19	3	6	2	30	58	19.44	8.69	8.56	7.5	9.63
19	3	6	2	31	59	19.44	8.75	8.56	7.5	9.63
19	3	6	2	32	59	19.44	8.75	8.63	7.5	9.69
19	3	6	2	33	59	19.44	8.75	8.56	7.5	9.69
19	3	6	2	35	0	19.44	8.75	8.63	7.5	9.69
19	3	6	2	36	0	19.37	8.81	8.63	7.5	9.69
19	3	6	2	37	0	19.37	8.81	8.63	7.5	9.69
19	3	6	2	38	0	19.37	8.88	8.63	7.5	9.69
19	3	6	2	39	1	19.31	8.88	8.63	7.5	9.69
19	3	6	2	40	1	19.31	8.88	8.63	7.56	9.69
19	3	6	2	41	1	19.31	8.94	8.63	7.56	9.69
19	3	6	2	42	1	19.25	8.94	8.63	7.56	9.69
19	3	6	2	43	2	19.25	8.94	8.63	7.56	9.69
19	3	6	2	44	2	19.25	9	8.63	7.56	9.69
19	3	6	2	45	2	19.25	9	8.63	7.56	9.69
19	3	6	2	46	2	19.31	9	8.69	7.56	9.69
19	3	6	2	47	3	19.31	9.06	8.69	7.56	9.69
19	3	6	2	48	3	19.31	9.06	8.69	7.56	9.75
19	3	6	2	49	3	19.31	9.06	8.69	7.56	9.75
19	3	6	2	50	3	19.37	9.13	8.69	7.56	9.75
19	3	6	2	51	4	19.37	9.13	8.69	7.56	9.75
19	3	6	2	52	4	19.37	9.13	8.69	7.56	9.75
19	3	6	2	53	4	19.37	9.19	8.69	7.56	9.75
19	3	6	2	54	4	19.37	9.19	8.69	7.56	9.75
19	3	6	2	55	5	19.37	9.19	8.69	7.63	9.75
19	3	6	2	56	5	19.37	9.25	8.69	7.63	9.75
19	3	6	2	57	5	19.37	9.25	8.75	7.63	9.75
19	3	6	2	58	5	19.31	9.25	8.75	7.63	9.75
19	3	6	2	59	6	19.31	9.31	8.75	7.63	9.75
19	3	6	3	0	6	19.31	9.31	8.75	7.63	9.75
19	3	6	3	1	6	19.25	9.31	8.75	7.63	9.81
19	3	6	3	2	6	19.25	9.31	8.75	7.63	9.81
19	3	6	3	3	7	19.25	9.38	8.75	7.63	9.81
19	3	6	3	4	7	19.25	9.38	8.75	7.63	9.81
19	3	6	3	5	7	19.25	9.38	8.75	7.63	9.81
19	3	6	3	6	8	19.31	9.44	8.75	7.63	9.81
19	3	6	3	7	8	19.31	9.44	8.75	7.63	9.81
19	3	6	3	8	8	19.31	9.44	8.75	7.63	9.81
19	3	6	3	9	8	19.37	9.44	8.75	7.63	9.81
19	3	6	3	10	9	19.37	9.5	8.81	7.63	9.81
19	3	6	3	11	9	19.44	9.5	8.81	7.63	9.81
19	3	6	3	12	9	19.44	9.5	8.81	7.69	9.81
19	3	6	3	13	9	19.5	9.56	8.75	7.69	9.81
19	3	6	3	14	10	19.5	9.56	8.81	7.69	9.81
19	3	6	3	15	10	19.5	9.56	8.81	7.69	9.81
19	3	6	3	16	10	19.5	9.63	8.81	7.69	9.81
19	3	6	3	17	10	19.5	9.63	8.81	7.69	9.88
19	3	6	3	18	11	19.5	9.63	8.81	7.69	9.88
19	3	6	3	19	11	19.5	9.63	8.81	7.69	9.81
19	3	6	3	20	11	19.44	9.69	8.81	7.69	9.88
19	3	6	3	21	11	19.44	9.69	8.81	7.69	9.88
19	3	6	3	22	12	19.44	9.69	8.81	7.69	9.88
19	3	6	3	23	12	19.44	9.75	8.81	7.69	9.88
19	3	6	3	24	12	19.37	9.75	8.81	7.69	9.88
19	3	6	3	25	12	19.37	9.75	8.81	7.75	9.88
19	3	6	3	26	13	19.37	9.75	8.88	7.69	9.88
19	3	6	3	27	13	19.31	9.81	8.88	7.75	9.88
19	3	6	3	28	13	19.31	9.81	8.88	7.75	9.88
19	3	6	3	29	13	19.31	9.88	8.88	7.75	9.88
19	3	6	3	30	14	19.31	9.88	8.88	7.75	9.88
19	3	6	3	31	14	19.37	9.88	8.88	7.75	9.88
19	3	6	3	32	14	19.37	9.88	8.88	7.75	9.88
19	3	6	3	33	15	19.37	9.88	8.88	7.75	9.88
19	3	6	3	34	15	19.37	9.94	8.88	7.75	9.94
19	3	6	3	35	15	19.37	9.94	8.88	7.75	9.94
19	3	6	3	36	15	19.44	9.94	8.88	7.75	9.94
19	3	6	3	37	16	19.44	10	8.88	7.75	9.94
19	3	6	3	38	16	19.44	10	8.88	7.75	9.94
19	3	6	3	39	16	19.37	10	8.88	7.75	9.94
19	3	6	3	40	16	19.37	10.06	8.94	7.75	9.94
19	3	6	3	41	17	19.37	10.06	8.94	7.81	9.94
19	3	6	3	42	17	19.37	10.06	8.94	7.75	9.94
19	3	6	3	43	17	19.31	10.06	8.94	7.81	9.94
19	3	6	3	44	17	19.31	10.13	8.94	7.81	9.94
19	3	6	3	45	18	19.31	10.13	8.94	7.81	9.94
19	3	6	3	46	18	19.25	10.13	8.94	7.81	9.94
19	3	6	3	47	18	19.25	10.19	8.94	7.81	9.94
19	3	6	3	48	19	19.25	10.19	8.94	7.81	9.94
19	3	6	3	49	19	19.25	10.19	8.94	7.81	9.94
19	3	6	3	50	19	19.25	10.19	8.94	7.81	10
19	3	6	3	51	19	19.31	10.25	8.94	7.81	10
19	3	6	3	52	20	19.31	10.25	8.94	7.81	10
19	3	6	3	53	20	19.31	10.25	8.94	7.81	10
19	3	6	3	54	20	19.31	10.25	8.94	7.81	10
19	3	6	3	55	20	19.31	10.31	9	7.88	10
19	3	6	3	56	21	19.31	10.31	8.94	7.88	10
19	3	6	3	57	21	19.25	10.31	9	7.88	10
19	3	6	3	58	21	19.25	10.38	9	7.88	10
19	3	6	3	59	21	19.25	10.38	9	7.88	10
19	3	6	4	0	22	19.25	10.38	9	7.88	10
19	3	6	4	1	22	19.19	10.38	9	7.88	10
19	3	6	4	2	22	19.19	10.38	9	7.88	10
19	3	6	4	3	22	19.19	10.44	9	7.88	10
19	3	6	4	4	23	19.19	10.44	9	7.88	10
19	3	6	4	5	23	19.19	10.44	9	7.88	10.06
19	3	6	4	6	23	19.25	10.44	9	7.88	10
19	3	6	4	7	23	19.25	10.44	9	7.88	10
19	3	6	4	8	24	19.25	10.5	9	7.88	10
19	3	6	4	9	24	19.31	10.5	9.06	7.88	10.06
19	3	6	4	10	24	19.31	10.5	9.06	7.88	10
19	3	6	4	11	24	19.37	10.5	9.06	7.94	10.06
19	3	6	4	12	25	19.37	10.56	9.06	7.94	10.06
19	3	6	4	13	25	19.37	10.56	9.06	7.94	10.06
19	3	6	4	14	25	19.37	10.56	9.06	7.94	10.06
19	3	6	4	15	25	19.37	10.56	9.06	7.94	10.06
19	3	6	4	16	26	19.37	10.63	9.06	7.94	10.06
19	3	6	4	17	26	19.31	10.63	9.06	7.94	10.06
19	3	6	4	18	26	19.31	10.63	9.06	7.94	10.06
19	3	6	4	19	27	19.31	10.63	9.06	7.94	10.06
19	3	6	4	20	27	19.31	10.63	9.06	7.94	10.06
19	3	6	4	21	27	19.25	10.69	9.06	7.94	10.06
19	3	6	4	22	27	19.25	10.69	9.06	7.94	10.06
19	3	6	4	23	28	19.25	10.69	9.06	7.94	10.06
19	3	6	4	24	28	19.31	10.69	9.06	7.94	10.06
19	3	6	4	25	28	19.25	10.69	9.06	7.94	10.06
19	3	6	4	26	28	19.31	10.75	9.06	8	10.13
19	3	6	4	27	29	19.31	10.75	9.13	8	10.13
19	3	6	4	28	31	19.31	10.75	9.06	8	10.13
19	3	6	4	29	31	19.37	10.81	9.13	8	10.13
19	3	6	4	30	31	19.37	10.81	9.13	8	10.13
19	3	6	4	31	32	19.37	10.81	9.13	8	10.13
19	3	6	4	32	32	19.37	10.81	9.13	8	10.13
19	3	6	4	33	32	19.37	10.81	9.13	8	10.13
19	3	6	4	34	32	19.37	10.88	9.13	8	10.13
19	3	6	4	35	33	19.37	10.81	9.13	8	10.13
19	3	6	4	36	33	19.31	10.88	9.13	8	10.13
19	3	6	4	37	33	19.31	10.88	9.13	8	10.13
19	3	6	4	38	34	19.31	10.88	9.13	8	10.13
19	3	6	4	39	34	19.31	10.94	9.13	8	10.13
19	3	6	4	40	34	19.25	10.94	9.13	8	10.13
19	3	6	4	41	34	19.25	10.94	9.13	8	10.13
19	3	6	4	42	35	19.25	10.94	9.19	8.06	10.13
19	3	6	4	43	35	19.25	10.94	9.19	8.06	10.19
19	3	6	4	44	35	19.25	11	9.19	8.06	10.19
19	3	6	4	45	35	19.25	11	9.19	8.06	10.19
19	3	6	4	46	36	19.25	11	9.19	8.06	10.19
19	3	6	4	47	36	19.25	11	9.19	8.06	10.19
19	3	6	4	48	36	19.25	11	9.19	8.06	10.19
19	3	6	4	49	36	19.31	11	9.19	8.06	10.19
19	3	6	4	50	37	19.31	11.06	9.19	8.06	10.19
19	3	6	4	51	37	19.31	11.06	9.19	8.06	10.19
19	3	6	4	52	37	19.31	11.06	9.19	8.06	10.19
19	3	6	4	53	37	19.31	11.06	9.19	8.06	10.19
19	3	6	4	54	38	19.37	11.06	9.19	8.06	10.19
19	3	6	4	55	38	19.37	11.13	9.25	8.06	10.19
19	3	6	4	56	38	19.37	11.13	9.19	8.13	10.19
19	3	6	4	57	39	19.37	11.13	9.19	8.06	10.19
19	3	6	4	58	39	19.37	11.13	9.25	8.13	10.19
19	3	6	4	59	39	19.37	11.19	9.25	8.13	10.19
19	3	6	5	0	39	19.37	11.19	9.25	8.13	10.19
19	3	6	5	1	40	19.37	11.19	9.25	8.13	10.19
19	3	6	5	2	40	19.37	11.19	9.25	8.13	10.19
19	3	6	5	3	40	19.37	11.19	9.25	8.13	10.19
19	3	6	5	4	40	19.37	11.19	9.25	8.13	10.25
19	3	6	5	5	41	19.37	11.19	9.25	8.13	10.25
19	3	6	5	6	41	19.31	11.25	9.25	8.13	10.25
19	3	6	5	7	41	19.31	11.25	9.25	8.13	10.25
19	3	6	5	8	41	19.31	11.25	9.25	8.13	10.25
19	3	6	5	9	42	19.31	11.25	9.25	8.13	10.25
19	3	6	5	10	42	19.25	11.25	9.25	8.13	10.25
19	3	6	5	11	42	19.25	11.31	9.25	8.19	10.25
19	3	6	5	12	42	19.25	11.31	9.25	8.19	10.25
19	3	6	5	13	43	19.25	11.31	9.31	8.19	10.25
19	3	6	5	14	43	19.25	11.31	9.31	8.19	10.25
19	3	6	5	15	43	19.25	11.31	9.31	8.19	10.25
19	3	6	5	16	43	19.25	11.38	9.25	8.19	10.25
19	3	6	5	17	44	19.25	11.38	9.31	8.19	10.25
19	3	6	5	18	44	19.31	11.38	9.31	8.19	10.25
19	3	6	5	19	44	19.31	11.38	9.31	8.19	10.25
19	3	6	5	20	44	19.31	11.38	9.31	8.19	10.25
19	3	6	5	21	45	19.31	11.38	9.31	8.19	10.25
19	3	6	5	22	45	19.31	11.38	9.31	8.19	10.31
19	3	6	5	23	45	19.31	11.44	9.31	8.19	10.25
19	3	6	5	24	46	19.31	11.44	9.31	8.19	10.31
19	3	6	5	25	46	19.31	11.44	9.31	8.25	10.31
19	3	6	5	26	46	19.25	11.44	9.31	8.25	10.31
19	3	6	5	27	46	19.25	11.44	9.38	8.25	10.31
19	3	6	5	28	47	19.25	11.5	9.31	8.25	10.31
19	3	6	5	29	47	19.19	11.5	9.38	8.25	10.31
19	3	6	5	30	47	19.19	11.5	9.38	8.25	10.31
19	3	6	5	31	47	19.12	11.5	9.38	8.25	10.31
19	3	6	5	32	48	19.12	11.5	9.38	8.25	10.31
19	3	6	5	33	48	19.12	11.5	9.38	8.25	10.31
19	3	6	5	34	48	19.12	11.56	9.38	8.25	10.31
19	3	6	5	35	48	19.06	11.56	9.38	8.25	10.31
19	3	6	5	36	49	19.12	11.56	9.38	8.25	10.31
19	3	6	5	37	49	19.12	11.56	9.38	8.25	10.31
19	3	6	5	38	49	19.12	11.56	9.38	8.25	10.31
19	3	6	5	39	49	19.12	11.56	9.38	8.25	10.31
19	3	6	5	40	50	19.12	11.63	9.38	8.25	10.31
19	3	6	5	41	50	19.12	11.63	9.38	8.25	10.38
19	3	6	5	42	50	19.19	11.63	9.38	8.31	10.38
19	3	6	5	43	50	19.19	11.63	9.44	8.31	10.38
19	3	6	5	44	51	19.19	11.63	9.44	8.31	10.38
19	3	6	5	45	51	19.19	11.63	9.38	8.31	10.38
19	3	6	5	46	51	19.25	11.63	9.44	8.31	10.38
19	3	6	5	47	51	19.25	11.69	9.44	8.31	10.38
19	3	6	5	48	52	19.25	11.69	9.44	8.31	10.38
19	3	6	5	49	52	19.25	11.69	9.44	8.31	10.38
19	3	6	5	50	52	19.25	11.69	9.44	8.31	10.38
19	3	6	5	51	52	19.25	11.69	9.44	8.31	10.38
19	3	6	5	52	53	19.25	11.69	9.44	8.31	10.38
19	3	6	5	53	53	19.25	11.69	9.44	8.38	10.38
19	3	6	5	54	53	19.19	11.75	9.44	8.38	10.38
19	3	6	5	55	54	19.19	11.75	9.44	8.38	10.38
19	3	6	5	56	54	19.19	11.75	9.44	8.38	10.38
19	3	6	5	57	54	19.19	11.75	9.44	8.38	10.38
19	3	6	5	58	54	19.19	11.75	9.44	8.38	10.38
19	3	6	5	59	55	19.19	11.75	9.44	8.38	10.44
19	3	6	6	0	55	19.19	11.81	9.44	8.38	10.38
19	3	6	6	1	55	19.19	11.81	9.5	8.38	10.38
19	3	6	6	2	55	19.19	11.81	9.5	8.38	10.38
19	3	6	6	3	56	19.19	11.81	9.5	8.38	10.38
19	3	6	6	4	56	19.25	11.81	9.5	8.38	10.44
19	3	6	6	5	56	19.25	11.81	9.5	8.38	10.44
19	3	6	6	6	56	19.25	11.81	9.5	8.38	10.44
19	3	6	6	7	57	19.31	11.88	9.5	8.38	10.44
19	3	6	6	8	57	19.31	11.88	9.5	8.38	10.44
19	3	6	6	9	57	19.31	11.88	9.5	8.38	10.44
19	3	6	6	10	57	19.37	11.88	9.5	8.38	10.44
19	3	6	6	11	58	19.37	11.88	9.5	8.44	10.44
19	3	6	6	12	58	19.37	11.88	9.5	8.44	10.44
19	3	6	6	13	58	19.31	11.88	9.5	8.44	10.44
19	3	6	6	14	58	19.31	11.94	9.5	8.44	10.44
19	3	6	6	15	59	19.31	11.94	9.56	8.44	10.44
19	3	6	6	16	59	19.31	11.94	9.5	8.44	10.44
19	3	6	6	17	59	19.31	11.94	9.56	8.44	10.44
19	3	6	6	18	59	19.25	11.94	9.56	8.44	10.44
19	3	6	6	20	0	19.25	11.94	9.56	8.44	10.44
19	3	6	6	21	0	19.19	11.94	9.56	8.44	10.44
19	3	6	6	22	0	19.19	11.94	9.56	8.44	10.44
19	3	6	6	23	1	19.19	12	9.56	8.44	10.5
19	3	6	6	24	1	19.25	12	9.56	8.44	10.5
19	3	6	6	25	1	19.25	12	9.56	8.44	10.5
19	3	6	6	26	1	19.25	12	9.56	8.44	10.5
19	3	6	6	27	2	19.25	12	9.56	8.5	10.5
19	3	6	6	28	2	19.31	12	9.56	8.5	10.5
19	3	6	6	29	2	19.31	12.06	9.56	8.5	10.5
19	3	6	6	30	2	19.37	12.06	9.56	8.44	10.5
19	3	6	6	31	3	19.37	12.06	9.56	8.5	10.5
19	3	6	6	32	3	19.44	12.06	9.56	8.5	10.5
19	3	6	6	33	3	19.44	12.06	9.56	8.5	10.5
19	3	6	6	34	3	19.37	12.06	9.56	8.5	10.5
19	3	6	6	35	4	19.44	12.06	9.56	8.5	10.5
19	3	6	6	36	4	19.37	12.13	9.63	8.5	10.56
19	3	6	6	37	4	19.37	12.13	9.63	8.5	10.56
19	3	6	6	38	4	19.37	12.13	9.63	8.5	10.5
19	3	6	6	39	5	19.37	12.13	9.63	8.5	10.5
19	3	6	6	40	5	19.37	12.13	9.63	8.5	10.56
19	3	6	6	41	5	19.31	12.13	9.63	8.56	10.56
19	3	6	6	42	5	19.31	12.13	9.63	8.56	10.56
19	3	6	6	43	6	19.31	12.13	9.63	8.5	10.56
19	3	6	6	44	6	19.31	12.19	9.63	8.56	10.56
19	3	6	6	45	6	19.25	12.19	9.63	8.56	10.56
19	3	6	6	46	6	19.31	12.19	9.63	8.56	10.56